## Contents
SPI ETH ESP32'!A1
Pinout ETH DEEK'!A1
JTAG conn'!A1
SD usage'!A1
UDP packets'!A1
UDP datagram definition'!A1
Message Types'!A1
Meas_sample!A1
Sensor Type table'!A1
ADC autorange'!A1
Sensor configuration set'!A1
LabVIEW data log'!A1
Error Codes'!A1
ADC selection'!A1
Time budget'!A1
TO DO'!A1
System setup'!A1
Sensors Example datasheet val.'!A1
IssueTracking!A1
Button and LED UI table'!A1

## SPI ETH ESP32
Contents!A1
ESP32 pin | ESP32 | Huzzah | cable | datasheet | meas on | ETH DEEK | brown | 3v3 | IO
VSPI | IO | pin |  | ESP marking | cables | pin conn
29 | IO5 | SCK/5 | y-purple | CS | SCK | 13 | black | GND
37 | IO23 | SDA/23 | y-grey | MOSI | highZ?CS | 10 | red | ETH_RST | 4
31 | IO19 | MISO/19 | y-white | MISO | highZ?MISO | 12 | orange | SD_CS | 17
30 | IO18 | MOSI/18 | grey | SCK | MOSI | 11 | yellow | SCK | 5
|  | GND | white |  | 0 | GND | green | MOSI | 18
|  | USB | purple |  | 5V | 5V | white | MISO | 19
[7 empty] | grey | CS_ETH | 23

## Pinout ETH DEEK
Contents!A1
ETH* | name
pin
10 | CS
11 | MOSI
12 | MISO
13 | SCK
GND
5V
*Deek-robot
NANO Ethernet Shield V1.0

## JTAG conn
Contents!A1
JLINK | JLINK | cable | cable | ESP32 | ESP32 | huzzah | aligned with>
pin | name | color | color 2 | pin | signal |  | esp docs:
1 | Vcc | red | red | 2 |  | 3V |  | ESP32 Pin | JTAG Signal
4 | GND | black | brown | 1 | GND | GND | 1 | CHIP_PU | TRST_N
5 | TDI | white | black | 14 | IO12 | 12 | 2 | MTDO / GPIO15 | TDO
7 | TMS | grey | white | 13 | IO14 | 14 | 3 | MTDI / GPIO12 | TDI
9 | TCK | purple | grey | 16 | IO13 | 13 | 4 | MTCK / GPIO13 | TCK
13 | TDO | blue | blue | 23 | IO15 | 15 | 5 | MTMS / GPIO14 | TMS
15 | nRST | green | green | 3 | EN | RST | 6 | GND | GND
3 | nTRST |  | maroon |  |  |  | https://docs.espressif.com/projects/esp-idf/en/latest/api-guides/jtag-debugging/configure-other-jtag.html

## SD usage
Contents!A1
| qty | Bytes
Total memory |  | 2000000000 | 2GB SD card
One sensor |  | 2 | bytes
Sensors | 12 | 24 | bytes
period |  | 1 | s
Lasts for |  | 83333333.3333 | s
|  | 23148.1481 | h
|  | 925.9259 | d
|  | 2.5368 | y

## UDP packets
Contents!A1
Buffer ENC28J60 | 8000 | B
UDP message size max | 65507 | B
Ethernet max packet size | 1500 | B
Reasonable safe max UDP message size | 508 | B
Wireshark Filter
(udp) && ((udp.dstport == 1111)¦¦(udp.dstport == 1100)¦¦(udp.dstport == 1024)¦¦(udp.dstport == 65500)¦¦(udp.dstport == 65511))

## UDP datagram definition
Contents!A1
UDP datagram definition
Sensor unit to Master |  |  | Master to Sensor unit
Byte order | Saved data | Representation | Byte order | Saved data | Representation
0 | timestamp epoch time or sample order MSB | uint_32 | 0 | datagram order MSB | uint_32
1 | timestamp epoch time or sample order | uint_32 | 1 | datagram order | uint_32
2 | timestamp epoch time or sample order | uint_32 | 2 | datagram order | uint_32
3 | timestamp epoch time or sample order LSB | uint_32 | 3 | datagram order LSB | uint_32
4 | Message type | uint_8 | 4 | Message type | uint_8
5 | Unit/sender number - last byte of IP | uint8_t | 5 | Command / ACK | uint8_t
6 |  | uint_32 | 6 |  | uint_32
7 |  | uint_32 | 7 |  | uint_32
8 |  | uint_32 | 8 |  | uint_32
9 |  |  | 9
10 |  |  | 10
11 |  |  | 11

## Message Types
Contents!A1
Sensor unit |  |  |  | Master
Byte | Value | Message type
4 | 4 | ANNOUNCEMENT announcing sensor unit is on network
EOP | EOP | EOP |  | Byte | Value | Message type
|  |  |  | 4 | 1 | ACKNOWLEDGEMENT
|  |  |  | EOP | EOP | EOP
Sensor unit enters IDLE state
Byte | Value
4 | 3 | ERROR |  | Master logs error
5 | uint_8 | sender number
6 | uint_8 | ERROR CODE
|  |  |  | Byte | Value
|  |  |  | 4 | 5 | Sensor power control
|  |  |  | 5..6 | 0xFFFF | Power to all sensors on
[5 empty] | 0x0000 | Power to all sensors off
Note: all combinations of bits. Ordered as on IO expander. Log 1 will turn the output of IO exp LOW, thus voltage output on sensor port HIGH. Byte 5 controls port 1, byte 6 is for port 0.
Byte | Value
4 | 5 | Sensor power report
5 | uint_8 | sender number
6..7 | uint_16 | Inverted logic levels on IO expander pins
Master checks against requested setting
|  |  |  | Byte | Value | Message type
|  | Unit resets |  | 4 | 2 | RESET unit
|  |  |  | EOP | EOP | EOP
|  |  |  | Byte | Value | Message type
|  | Trigger measurement ON |  | 4 | 3 | Start measurement trigger
|  |  |  | 5 | uint8_t | Sample period in 0.1s units (1 is 100ms)
|  |  |  | 6 | uint8_t | Sensor configuration set
|  |  |  | EOP | EOP | EOP
|  |  |  | Byte | Value | Message type
|  | Stop measurement |  | 4 | 4 | Stop measurement trigger
|  |  |  | EOP | EOP | EOP
Byte | Value
4 | 2 | MEASUREMENT SAMPLE
5 | uint_8 | sender number |  |  |  |  | byte | 6
6 | uint_8 | range A1.0 / sensor type |  |  |  |  | ADC
7..8 | uint_16 | data raw A1.0 MSB:LSB |  |  |  |  | CH
9 | uint_8 | range A1.1 / type
10..11 | uint_16 | data raw A1.2
12 | uint_8 | range A1.2
13..14 | uint_16 | data raw A1.2 MSB:LSB
15 | uint_8 | range A1.3
16..17 | uint_16 | data raw A1.3
18 | uint_8 | range A2.0 |  | Byte | Value
19..20 | uint_16 | data raw A2.0 MSB:LSB
21 | uint_8 | range A2.1 |  |  | 6 | SENSORS DETECTED
22..23 | uint_16 | data raw A2.1 |  |  | uint_8 | SENSOR 1 TYPE
… | … | … |  |  | X | table
39 | uint_8 | range A3.3 |  |  | uint_8 | SENSOR 1 OK
40..41 | uint_16 | data raw A3.3 |  |  | 0xFF | OK
42 | uint_8 | sensor type |  |  | 0X00 | NOK
43..44 | uint_16 | Data MSB:LSB |  |  | uint_8 | SENSOR 2 TYPE
… | … | … |  |  | uint_8 | SENSOR 2 OK
end |  | end character |  | … | … | …
[5 empty] | uint_8 | SENSOR 12 TYPE
[5 empty] | uint_8 | SENSOR 12 OK
Byte | Value
4 | 5 | Sensor power control | done
Start trigger
Stop trigger
Pause trigger
Error
Built-in Test trigger
Sensors detected
Read One Shot
Retrieve One measurement
Sensor power control
Set period

## Meas_sample
Contents!A1
Measurement sample datagram definition
Unit is sending the data in following format | [6 empty] | Data from one sensor
Byte | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | … | X
Example 1 | 0 | 0 | 0 | 5 | 2 | 13 | 0 | 0x01 | 0xA1 | 0 | 0x01 | 0xA0 | … | 0xFF
Meaning | Sample number 5 |  |  |  | payload is measurement sample if ==2 | sender address (last byte of IP addr) | see table sensor type. | Sensor reading, MSB:LSB |  | see table sensor type. | Sensor reading, MSB:LSB |  | … | end of data
Example 1 meaning | Sample number 5 |  |  |  | sending measurement sample | sender address is 13 | Analog input A1.0 | Sensor reads int16 0x01A1 = decimal 417 |  | Analog input A1.1 | Sensor reads int16 0x01A0 = decimal 410 |  | … | No meaning

## Sensor Type table
Contents!A1
input may be determined also by index in datagram
Sensor type table |  | note: to get second digit of hex number (range) use remainder after division by 16 |  |  |  |  | not needed for Labview saved data or SD saved data (already converted do decimal) | Safer is to use the sensor type value
Sensor type value (HEX) | Sensor type value (DEC) | Sensor type | Analog input | Conversion calculation | Output units | Note | Data Representation in datagram | Index in datagram | AI port
0 | 0 | ADC analog input 6.144V range | A1.0 | 0.1875*X | mV |  | int16 | 6 | A1.0
1 | 1 | ADC analog input 4.096V range | A1.0 | 0.125*X | mV |  | int16 | 9 | A1.1
2 | 2 | ADC analog input 2.048V range | A1.0 | 0.0625*X | mV |  | int16 | 12 | A1.2
3 | 3 | ADC analog input 1.024V range | A1.0 | 0.03215*X | mV |  | int16 | 15 | A1.3
4 | 4 | ADC analog input 0.512V range | A1.0 | 0.015625*X | mV |  | int16 | 18 | A2.0
5 | 5 | ADC analog input 0.256V range | A1.0 | 0.007813*X | mV |  | int16 | 21 | A2.1
6 | 6 | ADC analog input 0.256V range | A1.0 | 0.007813*X | mV |  | int16 | 24 | A2.2
7 | 7 | ADC analog input 0.256V range | A1.0 | 0.007813*X | mV |  | int16 | 27 | A2.3
10 | 16 | ADC analog input 6.144V range | A1.1 | 0.1875*X | mV |  | int16 | 30 | A3.0
11 | 17 | ADC analog input 4.096V range | A1.1 | 0.125*X | mV |  | int16 | 33 | A3.1
12 | 18 | ADC analog input 2.048V range | A1.1 | 0.0625*X | mV |  | int16 | 36 | A3.2
13 | 19 | ADC analog input 1.024V range | A1.1 | 0.03215*X | mV |  | int16 | 39 | A3.3
14 | 20 | ADC analog input 0.512V range | A1.1 | 0.015625*X | mV |  | int16
15 | 21 | ADC analog input 0.256V range | A1.1 | 0.007813*X | mV |  | int16
16 | 22 | ADC analog input 0.256V range | A1.1 | 0.007813*X | mV |  | int16
17 | 23 | ADC analog input 0.256V range | A1.1 | 0.007813*X | mV |  | int16 | Rhcode | RH
20 | 32 | ADC analog input 6.144V range | A1.2 | 0.1875*X | mV |  | int16 | 28000 | 47.4058
21 | 33 | ADC analog input 4.096V range | A1.2 | 0.125*X | mV |  | int16 | Temp code
22 | 34 | ADC analog input 2.048V range | A1.2 | 0.0625*X | mV |  | int16 | 23000 | 14.8193
23 | 35 | ADC analog input 1.024V range | A1.2 | 0.03215*X | mV |  | int16
24 | 36 | ADC analog input 0.512V range | A1.2 | 0.015625*X | mV |  | int16
25 | 37 | ADC analog input 0.256V range | A1.2 | 0.007813*X | mV |  | int16
26 | 38 | ADC analog input 0.256V range | A1.2 | 0.007813*X | mV |  | int16
27 | 39 | ADC analog input 0.256V range | A1.2 | 0.007813*X | mV |  | int16
30 | 48 | ADC analog input 6.144V range | A1.3 | 0.1875*X | mV |  | int16
31 | 49 | ADC analog input 4.096V range | A1.3 | 0.125*X | mV |  | int16
32 | 50 | ADC analog input 2.048V range | A1.3 | 0.0625*X | mV |  | int16
33 | 51 | ADC analog input 1.024V range | A1.3 | 0.03215*X | mV |  | int16
34 | 52 | ADC analog input 0.512V range | A1.3 | 0.015625*X | mV |  | int16
35 | 53 | ADC analog input 0.256V range | A1.3 | 0.007813*X | mV |  | int16
36 | 54 | ADC analog input 0.256V range | A1.3 | 0.007813*X | mV |  | int16
37 | 55 | ADC analog input 0.256V range | A1.3 | 0.007813*X | mV |  | int16
40 | 64 | ADC analog input 6.144V range | A2.0 | 0.1875*X | mV |  | int16
41 | 65 | ADC analog input 4.096V range | A2.0 | 0.125*X | mV |  | int16
42 | 66 | ADC analog input 2.048V range | A2.0 | 0.0625*X | mV |  | int16
43 | 67 | ADC analog input 1.024V range | A2.0 | 0.03215*X | mV |  | int16
44 | 68 | ADC analog input 0.512V range | A2.0 | 0.015625*X | mV |  | int16
45 | 69 | ADC analog input 0.256V range | A2.0 | 0.007813*X | mV |  | int16
46 | 70 | ADC analog input 0.256V range | A2.0 | 0.007813*X | mV |  | int16
47 | 71 | ADC analog input 0.256V range | A2.0 | 0.007813*X | mV |  | int16
50 | 80 | ADC analog input 6.144V range | A2.1 | 0.1875*X | mV |  | int16
51 | 81 | ADC analog input 4.096V range | A2.1 | 0.125*X | mV |  | int16
52 | 82 | ADC analog input 2.048V range | A2.1 | 0.0625*X | mV |  | int16
53 | 83 | ADC analog input 1.024V range | A2.1 | 0.03215*X | mV |  | int16
54 | 84 | ADC analog input 0.512V range | A2.1 | 0.015625*X | mV |  | int16
55 | 85 | ADC analog input 0.256V range | A2.1 | 0.007813*X | mV |  | int16
56 | 86 | ADC analog input 0.256V range | A2.1 | 0.007813*X | mV |  | int16
57 | 87 | ADC analog input 0.256V range | A2.1 | 0.007813*X | mV |  | int16
60 | 96 | ADC analog input 6.144V range | A2.2 | 0.1875*X | mV |  | int16
61 | 97 | ADC analog input 4.096V range | A2.2 | 0.125*X | mV |  | int16
62 | 98 | ADC analog input 2.048V range | A2.2 | 0.0625*X | mV |  | int16
63 | 99 | ADC analog input 1.024V range | A2.2 | 0.03215*X | mV |  | int16
64 | 100 | ADC analog input 0.512V range | A2.2 | 0.015625*X | mV |  | int16
65 | 101 | ADC analog input 0.256V range | A2.2 | 0.007813*X | mV |  | int16
66 | 102 | ADC analog input 0.256V range | A2.2 | 0.007813*X | mV |  | int16
67 | 103 | ADC analog input 0.256V range | A2.2 | 0.007813*X | mV |  | int16
70 | 112 | ADC analog input 6.144V range | A2.3 | 0.1875*X | mV |  | int16
71 | 113 | ADC analog input 4.096V range | A2.3 | 0.125*X | mV |  | int16
72 | 114 | ADC analog input 2.048V range | A2.3 | 0.0625*X | mV |  | int16
73 | 115 | ADC analog input 1.024V range | A2.3 | 0.03215*X | mV |  | int16
74 | 116 | ADC analog input 0.512V range | A2.3 | 0.015625*X | mV |  | int16
75 | 117 | ADC analog input 0.256V range | A2.3 | 0.007813*X | mV |  | int16
76 | 118 | ADC analog input 0.256V range | A2.3 | 0.007813*X | mV |  | int16
77 | 119 | ADC analog input 0.256V range | A2.3 | 0.007813*X | mV |  | int16
80 | 128 | ADC analog input 6.144V range | A3.0 | 0.1875*X | mV |  | int16
81 | 129 | ADC analog input 4.096V range | A3.0 | 0.125*X | mV |  | int16
82 | 130 | ADC analog input 2.048V range | A3.0 | 0.0625*X | mV |  | int16
83 | 131 | ADC analog input 1.024V range | A3.0 | 0.03215*X | mV |  | int16
84 | 132 | ADC analog input 0.512V range | A3.0 | 0.015625*X | mV |  | int16
85 | 133 | ADC analog input 0.256V range | A3.0 | 0.007813*X | mV |  | int16
86 | 134 | ADC analog input 0.256V range | A3.0 | 0.007813*X | mV |  | int16
87 | 135 | ADC analog input 0.256V range | A3.0 | 0.007813*X | mV |  | int16
90 | 144 | ADC analog input 6.144V range | A3.1 | 0.1875*X | mV |  | int16
91 | 145 | ADC analog input 4.096V range | A3.1 | 0.125*X | mV |  | int16
92 | 146 | ADC analog input 2.048V range | A3.1 | 0.0625*X | mV |  | int16
93 | 147 | ADC analog input 1.024V range | A3.1 | 0.03215*X | mV |  | int16
94 | 148 | ADC analog input 0.512V range | A3.1 | 0.015625*X | mV |  | int16
95 | 149 | ADC analog input 0.256V range | A3.1 | 0.007813*X | mV |  | int16
96 | 150 | ADC analog input 0.256V range | A3.1 | 0.007813*X | mV |  | int16
97 | 151 | ADC analog input 0.256V range | A3.1 | 0.007813*X | mV |  | int16
A0 | 160 | ADC analog input 6.144V range | A3.2 | 0.1875*X | mV |  | int16
A1 | 161 | ADC analog input 4.096V range | A3.2 | 0.125*X | mV |  | int16
A2 | 162 | ADC analog input 2.048V range | A3.2 | 0.0625*X | mV |  | int16
A3 | 163 | ADC analog input 1.024V range | A3.2 | 0.03215*X | mV |  | int16
A4 | 164 | ADC analog input 0.512V range | A3.2 | 0.015625*X | mV |  | int16
A5 | 165 | ADC analog input 0.256V range | A3.2 | 0.007813*X | mV |  | int16
A6 | 166 | ADC analog input 0.256V range | A3.2 | 0.007813*X | mV |  | int16
A7 | 167 | ADC analog input 0.256V range | A3.2 | 0.007813*X | mV |  | int16
B0 | 176 | ADC analog input 6.144V range | A3.3 | 0.1875*X | mV |  | int16
B1 | 177 | ADC analog input 4.096V range | A3.3 | 0.125*X | mV |  | int16
B2 | 178 | ADC analog input 2.048V range | A3.3 | 0.0625*X | mV |  | int16
B3 | 179 | ADC analog input 1.024V range | A3.3 | 0.03215*X | mV |  | int16
B4 | 180 | ADC analog input 0.512V range | A3.3 | 0.015625*X | mV |  | int16
B5 | 181 | ADC analog input 0.256V range | A3.3 | 0.007813*X | mV |  | int16
B6 | 182 | ADC analog input 0.256V range | A3.3 | 0.007813*X | mV |  | int16
B7 | 183 | ADC analog input 0.256V range | A3.3 | 0.007813*X | mV |  | int16
… | … | … | … | … | … | Free for additional sensors
F0 | 240 | Light reading gain 1 |  | X*0.0576 | Lux | see VEML7700 datasheet | uint16
F1 | 241 | Light reading gain 2 |  | X/2*0.0576 | Lux | see VEML7700 datasheet | uint16
F2 | 242 | Light reading gain 1/8 |  | X*8*0.0576 | Lux | see VEML7700 datasheet | uint16
F3 | 243 | Light reading gain 1/4 |  | X*4*0.0576 | Lux | see VEML7700 datasheet | uint16
FD | 253 | Temperature reading |  | ((175.72*X)/65536)-46.85 | °C | see Si7021 datasheet | uint16
FE | 254 | Relative Humidity reading |  | ((125*X)/65536)-6 | %RH | see Si7021 datasheet | uint16
FF | 255 | No sensor, indicates end of data
Sample data |  | not complete formula
Sensor type | Raw value | Converted | Units
64 | 6374 | 1195.125 | mV
64 | 6372 | 1194.75 | mV
64 | 6372 | 1194.75 | mV
64 | 6370 | 1194.375 | mV
64 | 6367 | 1193.8125 | mV
64 | 6368 | 1194 | mV
64 | 6368 | 1194 | mV
64 | 6366 | 1193.625 | mV
65 | 9543 | 894.6562 | mV
66 | 19062 | 893.5312 | mV
66 | 19060 | 893.4375 | mV
66 | 19052 | 893.0625 | mV
66 | 19051 | 893.0156 | mV
66 | 19049 | 892.9219 | mV
66 | 19045 | 892.7344 | mV
66 | 19038 | 892.4062 | mV
66 | 19039 | 892.4531 | mV
66 | 19039 | 892.4531 | mV
66 | 19033 | 892.1719 | mV
66 | 19028 | 891.9375 | mV
66 | 19032 | 892.125 | mV
66 | 19027 | 891.8906 | mV
66 | 19023 | 891.7031 | mV
66 | 19024 | 891.75 | mV
66 | 19023 | 891.7031 | mV
66 | 19019 | 891.5156 | mV
66 | 19022 | 891.6562 | mV
66 | 19019 | 891.5156 | mV
66 | 19011 | 891.1406 | mV
66 | 19014 | 891.2812 | mV
66 | 19011 | 891.1406 | mV
66 | 19003 | 890.7656 | mV
66 | 19009 | 891.0469 | mV
66 | 19002 | 890.7188 | mV
66 | 19001 | 890.6719 | mV
66 | 19003 | 890.7656 | mV
66 | 18994 | 890.3438 | mV
66 | 18998 | 890.5312 | mV
66 | 18995 | 890.3906 | mV
66 | 18991 | 890.2031 | mV
66 | 18997 | 890.4844 | mV
66 | 18987 | 890.0156 | mV
66 | 18991 | 890.2031 | mV
66 | 18991 | 890.2031 | mV
66 | 18984 | 889.875 | mV
66 | 18990 | 890.1562 | mV
66 | 18980 | 889.6875 | mV
66 | 18985 | 889.9219 | mV
66 | 18977 | 889.5469 | mV
66 | 18983 | 889.8281 | mV
66 | 18973 | 889.3594 | mV
66 | 18978 | 889.5938 | mV
66 | 18974 | 889.4062 | mV
66 | 18973 | 889.3594 | mV
66 | 18974 | 889.4062 | mV
66 | 18967 | 889.0781 | mV
66 | 18973 | 889.3594 | mV
66 | 18966 | 889.0312 | mV
66 | 18971 | 889.2656 | mV
66 | 18963 | 888.8906 | mV
66 | 18969 | 889.1719 | mV
66 | 18966 | 889.0312 | mV
66 | 18963 | 888.8906 | mV
66 | 18965 | 888.9844 | mV
66 | 18959 | 888.7031 | mV
66 | 18963 | 888.8906 | mV
66 | 18955 | 888.5156 | mV
66 | 18960 | 888.75
66 | 18952 | 888.375
66 | 18959 | 888.7031
66 | 18951 | 888.3281
66 | 18958 | 888.6562
66 | 18947 | 888.1406
66 | 18954 | 888.4688
66 | 18946 | 888.0938
66 | 18952 | 888.375
66 | 18944 | 888
66 | 18950 | 888.2812
66 | 18943 | 887.9531
66 | 18950 | 888.2812
66 | 18941 | 887.8594
66 | 18947 | 888.1406
66 | 18941 | 887.8594
66 | 18947 | 888.1406
66 | 18938 | 887.7188
66 | 18943 | 887.9531
66 | 18935 | 887.5781
66 | 18940 | 887.8125
66 | 18932 | 887.4375
66 | 18938 | 887.7188
66 | 18933 | 887.4844
66 | 18938 | 887.7188
66 | 18931 | 887.3906
66 | 18936 | 887.625
66 | 18932 | 887.4375
66 | 18935 | 887.5781
66 | 18930 | 887.3438
66 | 18934 | 887.5312
66 | 18931 | 887.3906
66 | 18932 | 887.4375
66 | 18930 | 887.3438
66 | 18930 | 887.3438
66 | 18926 | 887.1562
66 | 18928 | 887.25
66 | 18925 | 887.1094
66 | 18927 | 887.2031
66 | 18923 | 887.0156
66 | 18924 | 887.0625
66 | 18922 | 886.9688
66 | 18922 | 886.9688
66 | 18922 | 886.9688
66 | 18922 | 886.9688
66 | 18918 | 886.7812
66 | 18921 | 886.9219
66 | 18920 | 886.875
66 | 18922 | 886.9688
66 | 18918 | 886.7812
66 | 18921 | 886.9219
66 | 18918 | 886.7812
66 | 18917 | 886.7344
66 | 18917 | 886.7344
66 | 18916 | 886.6875
66 | 18912 | 886.5
66 | 18915 | 886.6406
66 | 18910 | 886.4062
66 | 18913 | 886.5469
66 | 18910 | 886.4062
66 | 18910 | 886.4062
66 | 18911 | 886.4531
66 | 18910 | 886.4062
66 | 18910 | 886.4062
66 | 18910 | 886.4062
66 | 18908 | 886.3125
66 | 18911 | 886.4531
66 | 18907 | 886.2656
66 | 18909 | 886.3594
66 | 18908 | 886.3125
66 | 18906 | 886.2188
66 | 18906 | 886.2188
66 | 18904 | 886.125
66 | 18906 | 886.2188
66 | 18902 | 886.0312
66 | 18904 | 886.125
66 | 18901 | 885.9844
66 | 18902 | 886.0312
66 | 18899 | 885.8906
66 | 18903 | 886.0781
66 | 18901 | 885.9844
66 | 18900 | 885.9375
66 | 18903 | 886.0781
66 | 18899 | 885.8906
66 | 18901 | 885.9844
66 | 18901 | 885.9844
66 | 18896 | 885.75
66 | 18903 | 886.0781
66 | 18898 | 885.8438
66 | 18896 | 885.75
66 | 18900 | 885.9375
66 | 18896 | 885.75
66 | 18894 | 885.6562
66 | 18897 | 885.7969
66 | 18895 | 885.7031
66 | 18893 | 885.6094
66 | 18895 | 885.7031
66 | 18895 | 885.7031
66 | 18894 | 885.6562
66 | 18894 | 885.6562
66 | 18892 | 885.5625
66 | 18893 | 885.6094
66 | 18896 | 885.75
66 | 18895 | 885.7031
66 | 18891 | 885.5156
66 | 18895 | 885.7031
66 | 18892 | 885.5625
66 | 18886 | 885.2812
66 | 18892 | 885.5625
66 | 18889 | 885.4219
66 | 18885 | 885.2344
66 | 18890 | 885.4688
66 | 18889 | 885.4219
66 | 18884 | 885.1875
66 | 18888 | 885.375
66 | 18889 | 885.4219
66 | 18884 | 885.1875
66 | 18886 | 885.2812
66 | 18890 | 885.4688
66 | 18889 | 885.4219
66 | 18884 | 885.1875
66 | 18889 | 885.4219
66 | 18886 | 885.2812
66 | 18882 | 885.0938
66 | 18884 | 885.1875
66 | 18886 | 885.2812
66 | 18884 | 885.1875
66 | 18879 | 884.9531
66 | 18884 | 885.1875
66 | 18883 | 885.1406
66 | 18881 | 885.0469
66 | 18878 | 884.9062
66 | 18883 | 885.1406
66 | 18882 | 885.0938
66 | 18881 | 885.0469
66 | 18880 | 885
66 | 18883 | 885.1406
66 | 18882 | 885.0938
66 | 18878 | 884.9062
66 | 18880 | 885
66 | 18883 | 885.1406
66 | 18880 | 885
66 | 18876 | 884.8125
66 | 18881 | 885.0469
66 | 18880 | 885
66 | 18879 | 884.9531
66 | 18874 | 884.7188
66 | 18879 | 884.9531
66 | 18877 | 884.8594
66 | 18877 | 884.8594
66 | 18872 | 884.625
66 | 18879 | 884.9531
66 | 18878 | 884.9062
66 | 18878 | 884.9062
66 | 18874 | 884.7188
66 | 18879 | 884.9531
66 | 18880 | 885
66 | 18878 | 884.9062
66 | 18872 | 884.625
66 | 18875 | 884.7656
66 | 18877 | 884.8594
66 | 18877 | 884.8594
66 | 18876 | 884.8125
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18876 | 884.8125
66 | 18874 | 884.7188
66 | 18874 | 884.7188
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18873 | 884.6719
66 | 18877 | 884.8594
66 | 18875 | 884.7656
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18867 | 884.3906
66 | 18869 | 884.4844
66 | 18874 | 884.7188
66 | 18875 | 884.7656
66 | 18873 | 884.6719
66 | 18873 | 884.6719
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18869 | 884.4844
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18866 | 884.3438
66 | 18871 | 884.5781
66 | 18874 | 884.7188
66 | 18873 | 884.6719
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18871 | 884.5781
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18863 | 884.2031
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18863 | 884.2031
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18861 | 884.1094
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18842 | 883.2188
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18843 | 883.2656
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18848 | 883.5
66 | 18842 | 883.2188
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18847 | 883.4531
66 | 18841 | 883.1719
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18844 | 883.3125
66 | 18847 | 883.4531
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18847 | 883.4531
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18842 | 883.2188
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18842 | 883.2188
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18840 | 883.125
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18841 | 883.1719
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18843 | 883.2656
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18843 | 883.2656
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18842 | 883.2188
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18845 | 883.3594
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18848 | 883.5
66 | 18843 | 883.2656
66 | 18847 | 883.4531
66 | 18842 | 883.2188
66 | 18848 | 883.5
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18851 | 883.6406
66 | 18845 | 883.3594
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18849 | 883.5469
66 | 18841 | 883.1719
66 | 18848 | 883.5
66 | 18842 | 883.2188
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18851 | 883.6406
66 | 18844 | 883.3125
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18844 | 883.3125
66 | 18851 | 883.6406
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18851 | 883.6406
66 | 18842 | 883.2188
66 | 18849 | 883.5469
66 | 18841 | 883.1719
66 | 18848 | 883.5
66 | 18842 | 883.2188
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18852 | 883.6875
66 | 18845 | 883.3594
66 | 18852 | 883.6875
66 | 18842 | 883.2188
66 | 18852 | 883.6875
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18851 | 883.6406
66 | 18842 | 883.2188
66 | 18851 | 883.6406
66 | 18843 | 883.2656
66 | 18852 | 883.6875
66 | 18842 | 883.2188
66 | 18852 | 883.6875
66 | 18843 | 883.2656
66 | 18851 | 883.6406
66 | 18844 | 883.3125
66 | 18852 | 883.6875
66 | 18843 | 883.2656
66 | 18852 | 883.6875
66 | 18844 | 883.3125
66 | 18852 | 883.6875
66 | 18842 | 883.2188
66 | 18852 | 883.6875
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18846 | 883.4062
66 | 18851 | 883.6406
66 | 18844 | 883.3125
66 | 18851 | 883.6406
66 | 18847 | 883.4531
66 | 18851 | 883.6406
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18847 | 883.4531
66 | 18846 | 883.4062
66 | 18851 | 883.6406
66 | 18843 | 883.2656
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18849 | 883.5469
66 | 18849 | 883.5469
66 | 18851 | 883.6406
66 | 18845 | 883.3594
66 | 18852 | 883.6875
66 | 18847 | 883.4531
66 | 18851 | 883.6406
66 | 18847 | 883.4531
66 | 18843 | 883.2656
66 | 18851 | 883.6406
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18843 | 883.2656
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18847 | 883.4531
66 | 18852 | 883.6875
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18852 | 883.6875
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18843 | 883.2656
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18842 | 883.2188
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18847 | 883.4531
66 | 18850 | 883.5938
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18844 | 883.3125
66 | 18850 | 883.5938
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18850 | 883.5938
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18847 | 883.4531
66 | 18852 | 883.6875
66 | 18855 | 883.8281
66 | 18852 | 883.6875
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18855 | 883.8281
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18847 | 883.4531
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18860 | 884.0625
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18856 | 883.875
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18859 | 884.0156
66 | 18863 | 884.2031
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18861 | 884.1094
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18855 | 883.8281
66 | 18860 | 884.0625
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18856 | 883.875
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18858 | 883.9688
66 | 18855 | 883.8281
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18863 | 884.2031
66 | 18860 | 884.0625
66 | 18856 | 883.875
66 | 18862 | 884.1562
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18861 | 884.1094
66 | 18855 | 883.8281
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18855 | 883.8281
66 | 18860 | 884.0625
66 | 18863 | 884.2031
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18855 | 883.8281
66 | 18860 | 884.0625
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18862 | 884.1562
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18854 | 883.7812
66 | 18860 | 884.0625
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18855 | 883.8281
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18857 | 883.9219
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18863 | 884.2031
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18860 | 884.0625
66 | 18856 | 883.875
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18855 | 883.8281
66 | 18860 | 884.0625
66 | 18863 | 884.2031
66 | 18862 | 884.1562
66 | 18857 | 883.9219
66 | 18862 | 884.1562
66 | 18864 | 884.25
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18856 | 883.875
66 | 18861 | 884.1094
66 | 18863 | 884.2031
66 | 18860 | 884.0625
66 | 18854 | 883.7812
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18854 | 883.7812
66 | 18861 | 884.1094
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18866 | 884.3438
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18856 | 883.875
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18856 | 883.875
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18863 | 884.2031
66 | 18860 | 884.0625
66 | 18866 | 884.3438
66 | 18860 | 884.0625
66 | 18863 | 884.2031
66 | 18866 | 884.3438
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18858 | 883.9688
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18858 | 883.9688
66 | 18865 | 884.2969
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18866 | 884.3438
66 | 18861 | 884.1094
66 | 18864 | 884.25
66 | 18868 | 884.4375
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18857 | 883.9219
66 | 18865 | 884.2969
66 | 18863 | 884.2031
66 | 18859 | 884.0156
66 | 18866 | 884.3438
66 | 18861 | 884.1094
66 | 18863 | 884.2031
66 | 18867 | 884.3906
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18861 | 884.1094
66 | 18867 | 884.3906
66 | 18863 | 884.2031
66 | 18863 | 884.2031
66 | 18866 | 884.3438
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18859 | 884.0156
66 | 18868 | 884.4375
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18862 | 884.1562
66 | 18869 | 884.4844
66 | 18862 | 884.1562
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18869 | 884.4844
66 | 18861 | 884.1094
66 | 18867 | 884.3906
66 | 18860 | 884.0625
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18862 | 884.1562
66 | 18867 | 884.3906
66 | 18860 | 884.0625
66 | 18867 | 884.3906
66 | 18861 | 884.1094
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18870 | 884.5312
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18863 | 884.2031
66 | 18868 | 884.4375
66 | 18861 | 884.1094
66 | 18868 | 884.4375
66 | 18862 | 884.1562
66 | 18869 | 884.4844
66 | 18860 | 884.0625
66 | 18867 | 884.3906
66 | 18861 | 884.1094
66 | 18871 | 884.5781
66 | 18862 | 884.1562
66 | 18871 | 884.5781
66 | 18864 | 884.25
66 | 18871 | 884.5781
66 | 18865 | 884.2969
66 | 18869 | 884.4844
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18868 | 884.4375
66 | 18866 | 884.3438
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18871 | 884.5781
66 | 18869 | 884.4844
66 | 18871 | 884.5781
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18866 | 884.3438
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18862 | 884.1562
66 | 18871 | 884.5781
66 | 18863 | 884.2031
66 | 18869 | 884.4844
66 | 18871 | 884.5781
66 | 18870 | 884.5312
66 | 18873 | 884.6719
66 | 18868 | 884.4375
66 | 18873 | 884.6719
66 | 18866 | 884.3438
66 | 18873 | 884.6719
66 | 18871 | 884.5781
66 | 18868 | 884.4375
66 | 18872 | 884.625
66 | 18865 | 884.2969
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18865 | 884.2969
66 | 18872 | 884.625
66 | 18866 | 884.3438
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18865 | 884.2969
66 | 18874 | 884.7188
66 | 18871 | 884.5781
66 | 18868 | 884.4375
66 | 18874 | 884.7188
66 | 18867 | 884.3906
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18865 | 884.2969
66 | 18872 | 884.625
66 | 18872 | 884.625
66 | 18865 | 884.2969
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18874 | 884.7188
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18874 | 884.7188
66 | 18866 | 884.3438
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18865 | 884.2969
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18863 | 884.2031
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18864 | 884.25
66 | 18871 | 884.5781
66 | 18870 | 884.5312
66 | 18867 | 884.3906
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18866 | 884.3438
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18866 | 884.3438
66 | 18872 | 884.625
66 | 18872 | 884.625
66 | 18865 | 884.2969
66 | 18870 | 884.5312
66 | 18872 | 884.625
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18871 | 884.5781
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18873 | 884.6719
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18873 | 884.6719
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18873 | 884.6719
66 | 18866 | 884.3438
66 | 18869 | 884.4844
66 | 18873 | 884.6719
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18866 | 884.3438
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18864 | 884.25
66 | 18871 | 884.5781
66 | 18873 | 884.6719
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18875 | 884.7656
66 | 18871 | 884.5781
66 | 18867 | 884.3906
66 | 18875 | 884.7656
66 | 18872 | 884.625
66 | 18864 | 884.25
66 | 18868 | 884.4375
66 | 18873 | 884.6719
66 | 18870 | 884.5312
66 | 18865 | 884.2969
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18865 | 884.2969
66 | 18872 | 884.625
66 | 18876 | 884.8125
66 | 18872 | 884.625
66 | 18865 | 884.2969
66 | 18871 | 884.5781
66 | 18873 | 884.6719
66 | 18871 | 884.5781
66 | 18866 | 884.3438
66 | 18870 | 884.5312
66 | 18872 | 884.625
66 | 18871 | 884.5781
66 | 18863 | 884.2031
66 | 18869 | 884.4844
66 | 18872 | 884.625
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18872 | 884.625
66 | 18875 | 884.7656
66 | 18871 | 884.5781
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18873 | 884.6719
66 | 18873 | 884.6719
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18863 | 884.2031
66 | 18866 | 884.3438
66 | 18870 | 884.5312
66 | 18873 | 884.6719
66 | 18870 | 884.5312
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18872 | 884.625
66 | 18875 | 884.7656
66 | 18876 | 884.8125
66 | 18871 | 884.5781
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18870 | 884.5312
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18874 | 884.7188
66 | 18872 | 884.625
66 | 18871 | 884.5781
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18872 | 884.625
66 | 18875 | 884.7656
66 | 18875 | 884.7656
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18872 | 884.625
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18868 | 884.4375
66 | 18872 | 884.625
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18865 | 884.2969
66 | 18863 | 884.2031
66 | 18867 | 884.3906
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18872 | 884.625
66 | 18868 | 884.4375
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18873 | 884.6719
66 | 18872 | 884.625
66 | 18872 | 884.625
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18871 | 884.5781
66 | 18870 | 884.5312
66 | 18871 | 884.5781
66 | 18872 | 884.625
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18869 | 884.4844
66 | 18870 | 884.5312
66 | 18872 | 884.625
66 | 18874 | 884.7188
66 | 18872 | 884.625
66 | 18871 | 884.5781
66 | 18869 | 884.4844
66 | 18864 | 884.25
66 | 18862 | 884.1562
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18869 | 884.4844
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18865 | 884.2969
66 | 18861 | 884.1094
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18872 | 884.625
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18866 | 884.3438
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18870 | 884.5312
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18863 | 884.2031
66 | 18862 | 884.1562
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18870 | 884.5312
66 | 18866 | 884.3438
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18869 | 884.4844
66 | 18868 | 884.4375
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18864 | 884.25
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18866 | 884.3438
66 | 18868 | 884.4375
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18857 | 883.9219
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18856 | 883.875
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18855 | 883.8281
66 | 18861 | 884.1094
66 | 18864 | 884.25
66 | 18860 | 884.0625
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18860 | 884.0625
66 | 18868 | 884.4375
66 | 18863 | 884.2031
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18865 | 884.2969
66 | 18855 | 883.8281
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18866 | 884.3438
66 | 18863 | 884.2031
66 | 18861 | 884.1094
66 | 18866 | 884.3438
66 | 18859 | 884.0156
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18856 | 883.875
66 | 18864 | 884.25
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18859 | 884.0156
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18857 | 883.9219
66 | 18867 | 884.3906
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18858 | 883.9688
66 | 18863 | 884.2031
66 | 18856 | 883.875
66 | 18864 | 884.25
66 | 18856 | 883.875
66 | 18864 | 884.25
66 | 18862 | 884.1562
66 | 18860 | 884.0625
66 | 18865 | 884.2969
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18859 | 884.0156
66 | 18863 | 884.2031
66 | 18857 | 883.9219
66 | 18861 | 884.1094
66 | 18856 | 883.875
66 | 18862 | 884.1562
66 | 18854 | 883.7812
66 | 18863 | 884.2031
66 | 18858 | 883.9688
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18861 | 884.1094
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18860 | 884.0625
66 | 18858 | 883.9688
66 | 18860 | 884.0625
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18854 | 883.7812
66 | 18862 | 884.1562
66 | 18853 | 883.7344
66 | 18860 | 884.0625
66 | 18853 | 883.7344
66 | 18862 | 884.1562
66 | 18855 | 883.8281
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18854 | 883.7812
66 | 18858 | 883.9688
66 | 18851 | 883.6406
66 | 18859 | 884.0156
66 | 18850 | 883.5938
66 | 18858 | 883.9688
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18859 | 884.0156
66 | 18854 | 883.7812
66 | 18859 | 884.0156
66 | 18850 | 883.5938
66 | 18861 | 884.1094
66 | 18853 | 883.7344
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18852 | 883.6875
66 | 18858 | 883.9688
66 | 18849 | 883.5469
66 | 18857 | 883.9219
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18857 | 883.9219
66 | 18849 | 883.5469
66 | 18857 | 883.9219
66 | 18850 | 883.5938
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18852 | 883.6875
66 | 18858 | 883.9688
66 | 18851 | 883.6406
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18851 | 883.6406
66 | 18858 | 883.9688
66 | 18852 | 883.6875
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18848 | 883.5
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18848 | 883.5
66 | 18857 | 883.9219
66 | 18854 | 883.7812
66 | 18850 | 883.5938
66 | 18858 | 883.9688
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18859 | 884.0156
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18857 | 883.9219
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18848 | 883.5
66 | 18853 | 883.7344
66 | 18856 | 883.875
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18857 | 883.9219
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18857 | 883.9219
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18857 | 883.9219
66 | 18854 | 883.7812
66 | 18848 | 883.5
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18851 | 883.6406
66 | 18847 | 883.4531
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18856 | 883.875
66 | 18851 | 883.6406
66 | 18848 | 883.5
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18856 | 883.875
66 | 18852 | 883.6875
66 | 18846 | 883.4062
66 | 18848 | 883.5
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18844 | 883.3125
66 | 18849 | 883.5469
66 | 18854 | 883.7812
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18852 | 883.6875
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18851 | 883.6406
66 | 18845 | 883.3594
66 | 18850 | 883.5938
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18851 | 883.6406
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18854 | 883.7812
66 | 18855 | 883.8281
66 | 18851 | 883.6406
66 | 18845 | 883.3594
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18847 | 883.4531
66 | 18843 | 883.2656
66 | 18847 | 883.4531
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18846 | 883.4062
66 | 18844 | 883.3125
66 | 18849 | 883.5469
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18844 | 883.3125
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18843 | 883.2656
66 | 18841 | 883.1719
66 | 18843 | 883.2656
66 | 18848 | 883.5
66 | 18849 | 883.5469
66 | 18850 | 883.5938
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18845 | 883.3594
66 | 18844 | 883.3125
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18844 | 883.3125
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18840 | 883.125
66 | 18840 | 883.125
66 | 18843 | 883.2656
66 | 18844 | 883.3125
66 | 18847 | 883.4531
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18844 | 883.3125
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18844 | 883.3125
66 | 18844 | 883.3125
66 | 18844 | 883.3125
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18841 | 883.1719
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18840 | 883.125
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18845 | 883.3594
66 | 18844 | 883.3125
66 | 18845 | 883.3594
66 | 18843 | 883.2656
66 | 18844 | 883.3125
66 | 18844 | 883.3125
66 | 18842 | 883.2188
66 | 18844 | 883.3125
66 | 18844 | 883.3125
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18844 | 883.3125
66 | 18841 | 883.1719
66 | 18841 | 883.1719
66 | 18839 | 883.0781
66 | 18837 | 882.9844
66 | 18835 | 882.8906
66 | 18834 | 882.8438
66 | 18837 | 882.9844
66 | 18840 | 883.125
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18840 | 883.125
66 | 18836 | 882.9375
66 | 18834 | 882.8438
66 | 18832 | 882.75
66 | 18837 | 882.9844
66 | 18841 | 883.1719
66 | 18841 | 883.1719
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18840 | 883.125
66 | 18835 | 882.8906
66 | 18835 | 882.8906
66 | 18840 | 883.125
66 | 18841 | 883.1719
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18840 | 883.125
66 | 18837 | 882.9844
66 | 18832 | 882.75
66 | 18834 | 882.8438
66 | 18839 | 883.0781
66 | 18841 | 883.1719
66 | 18842 | 883.2188
66 | 18839 | 883.0781
66 | 18835 | 882.8906
66 | 18834 | 882.8438
66 | 18836 | 882.9375
66 | 18840 | 883.125
66 | 18841 | 883.1719
66 | 18843 | 883.2656
66 | 18839 | 883.0781
66 | 18837 | 882.9844
66 | 18833 | 882.7969
66 | 18838 | 883.0312
66 | 18840 | 883.125
66 | 18842 | 883.2188
66 | 18838 | 883.0312
66 | 18834 | 882.8438
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18839 | 883.0781
66 | 18840 | 883.125
66 | 18835 | 882.8906
66 | 18832 | 882.75
66 | 18832 | 882.75
66 | 18839 | 883.0781
66 | 18841 | 883.1719
66 | 18839 | 883.0781
66 | 18836 | 882.9375
66 | 18832 | 882.75
66 | 18838 | 883.0312
66 | 18839 | 883.0781
66 | 18840 | 883.125
66 | 18835 | 882.8906
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18839 | 883.0781
66 | 18839 | 883.0781
66 | 18836 | 882.9375
66 | 18830 | 882.6562
66 | 18837 | 882.9844
66 | 18838 | 883.0312
66 | 18840 | 883.125
66 | 18836 | 882.9375
66 | 18831 | 882.7031
66 | 18838 | 883.0312
66 | 18840 | 883.125
66 | 18839 | 883.0781
66 | 18835 | 882.8906
66 | 18833 | 882.7969
66 | 18838 | 883.0312
66 | 18838 | 883.0312
66 | 18835 | 882.8906
66 | 18830 | 882.6562
66 | 18835 | 882.8906
66 | 18838 | 883.0312
66 | 18834 | 882.8438
66 | 18830 | 882.6562
66 | 18835 | 882.8906
66 | 18839 | 883.0781
66 | 18838 | 883.0312
66 | 18832 | 882.75
66 | 18835 | 882.8906
66 | 18841 | 883.1719
66 | 18838 | 883.0312
66 | 18831 | 882.7031
66 | 18837 | 882.9844
66 | 18839 | 883.0781
66 | 18834 | 882.8438
66 | 18830 | 882.6562
66 | 18837 | 882.9844
66 | 18838 | 883.0312
66 | 18833 | 882.7969
66 | 18831 | 882.7031
66 | 18836 | 882.9375
66 | 18833 | 882.7969
66 | 18831 | 882.7031
66 | 18837 | 882.9844
66 | 18835 | 882.8906
66 | 18831 | 882.7031
66 | 18838 | 883.0312
66 | 18835 | 882.8906
66 | 18833 | 882.7969
66 | 18840 | 883.125
66 | 18834 | 882.8438
66 | 18833 | 882.7969
66 | 18839 | 883.0781
66 | 18833 | 882.7969
66 | 18832 | 882.75
66 | 18837 | 882.9844
66 | 18833 | 882.7969
66 | 18832 | 882.75
66 | 18837 | 882.9844
66 | 18830 | 882.6562
66 | 18834 | 882.8438
66 | 18837 | 882.9844
66 | 18831 | 882.7031
66 | 18833 | 882.7969
66 | 18838 | 883.0312
66 | 18831 | 882.7031
66 | 18836 | 882.9375
66 | 18837 | 882.9844
66 | 18830 | 882.6562
66 | 18835 | 882.8906
66 | 18836 | 882.9375
66 | 18827 | 882.5156
66 | 18836 | 882.9375
66 | 18833 | 882.7969
66 | 18831 | 882.7031
66 | 18834 | 882.8438
66 | 18826 | 882.4688
66 | 18834 | 882.8438
66 | 18834 | 882.8438
66 | 18831 | 882.7031
66 | 18835 | 882.8906
66 | 18827 | 882.5156
66 | 18835 | 882.8906
66 | 18835 | 882.8906
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18828 | 882.5625
66 | 18834 | 882.8438
66 | 18836 | 882.9375
66 | 18831 | 882.7031
66 | 18836 | 882.9375
66 | 18825 | 882.4219
66 | 18834 | 882.8438
66 | 18830 | 882.6562
66 | 18832 | 882.75
66 | 18833 | 882.7969
66 | 18827 | 882.5156
66 | 18835 | 882.8906
66 | 18827 | 882.5156
66 | 18836 | 882.9375
66 | 18829 | 882.6094
66 | 18837 | 882.9844
66 | 18833 | 882.7969
66 | 18833 | 882.7969
66 | 18834 | 882.8438
66 | 18832 | 882.75
66 | 18834 | 882.8438
66 | 18828 | 882.5625
66 | 18834 | 882.8438
66 | 18827 | 882.5156
66 | 18835 | 882.8906
66 | 18825 | 882.4219
66 | 18834 | 882.8438
66 | 18826 | 882.4688
66 | 18834 | 882.8438
66 | 18826 | 882.4688
66 | 18834 | 882.8438
66 | 18827 | 882.5156
66 | 18834 | 882.8438
66 | 18827 | 882.5156
66 | 18834 | 882.8438
66 | 18833 | 882.7969
66 | 18833 | 882.7969
66 | 18832 | 882.75
66 | 18832 | 882.75
66 | 18834 | 882.8438
66 | 18830 | 882.6562
66 | 18832 | 882.75
66 | 18830 | 882.6562
66 | 18833 | 882.7969
66 | 18826 | 882.4688
66 | 18834 | 882.8438
66 | 18826 | 882.4688
66 | 18834 | 882.8438
66 | 18825 | 882.4219
66 | 18836 | 882.9375
66 | 18828 | 882.5625
66 | 18833 | 882.7969
66 | 18834 | 882.8438
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18827 | 882.5156
66 | 18836 | 882.9375
66 | 18829 | 882.6094
66 | 18832 | 882.75
66 | 18833 | 882.7969
66 | 18826 | 882.4688
66 | 18835 | 882.8906
66 | 18827 | 882.5156
66 | 18831 | 882.7031
66 | 18833 | 882.7969
66 | 18824 | 882.375
66 | 18836 | 882.9375
66 | 18830 | 882.6562
66 | 18832 | 882.75
66 | 18836 | 882.9375
66 | 18830 | 882.6562
66 | 18831 | 882.7031
66 | 18834 | 882.8438
66 | 18825 | 882.4219
66 | 18835 | 882.8906
66 | 18831 | 882.7031
66 | 18830 | 882.6562
66 | 18834 | 882.8438
66 | 18827 | 882.5156
66 | 18831 | 882.7031
66 | 18835 | 882.8906
66 | 18827 | 882.5156
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18829 | 882.6094
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18831 | 882.7031
66 | 18830 | 882.6562
66 | 18837 | 882.9844
66 | 18831 | 882.7031
66 | 18829 | 882.6094
66 | 18837 | 882.9844
66 | 18831 | 882.7031
66 | 18828 | 882.5625
66 | 18835 | 882.8906
66 | 18831 | 882.7031
66 | 18826 | 882.4688
66 | 18834 | 882.8438
66 | 18833 | 882.7969
66 | 18827 | 882.5156
66 | 18830 | 882.6562
66 | 18836 | 882.9375
66 | 18833 | 882.7969
66 | 18827 | 882.5156
66 | 18834 | 882.8438
66 | 18835 | 882.8906
66 | 18832 | 882.75
66 | 18830 | 882.6562
66 | 18835 | 882.8906
66 | 18834 | 882.8438
66 | 18828 | 882.5625
66 | 18829 | 882.6094
66 | 18834 | 882.8438
66 | 18834 | 882.8438
66 | 18828 | 882.5625
66 | 18828 | 882.5625
66 | 18835 | 882.8906
66 | 18835 | 882.8906
66 | 18832 | 882.75
66 | 18827 | 882.5156
66 | 18834 | 882.8438
66 | 18835 | 882.8906
66 | 18836 | 882.9375
66 | 18835 | 882.8906
66 | 18831 | 882.7031
66 | 18827 | 882.5156
66 | 18828 | 882.5625
66 | 18833 | 882.7969
66 | 18836 | 882.9375
66 | 18833 | 882.7969
66 | 18832 | 882.75
66 | 18830 | 882.6562
66 | 18826 | 882.4688
66 | 18826 | 882.4688
66 | 18831 | 882.7031
66 | 18832 | 882.75
66 | 18835 | 882.8906
66 | 18836 | 882.9375
66 | 18835 | 882.8906
66 | 18835 | 882.8906
66 | 18830 | 882.6562
66 | 18830 | 882.6562
66 | 18827 | 882.5156
66 | 18827 | 882.5156
66 | 18831 | 882.7031
66 | 18831 | 882.7031
66 | 18834 | 882.8438
66 | 18833 | 882.7969
66 | 18835 | 882.8906
66 | 18835 | 882.8906
66 | 18836 | 882.9375
66 | 18837 | 882.9844
66 | 18835 | 882.8906
66 | 18836 | 882.9375
66 | 18835 | 882.8906
66 | 18833 | 882.7969
66 | 18835 | 882.8906
66 | 18835 | 882.8906
66 | 18833 | 882.7969
66 | 18832 | 882.75
66 | 18832 | 882.75
66 | 18832 | 882.75
66 | 18833 | 882.7969
66 | 18833 | 882.7969
66 | 18834 | 882.8438
66 | 18833 | 882.7969
66 | 18833 | 882.7969
66 | 18833 | 882.7969
66 | 18834 | 882.8438
66 | 18834 | 882.8438
66 | 18835 | 882.8906
66 | 18834 | 882.8438
66 | 18836 | 882.9375
66 | 18837 | 882.9844
66 | 18838 | 883.0312
66 | 18836 | 882.9375
66 | 18838 | 883.0312
66 | 18837 | 882.9844
66 | 18837 | 882.9844
66 | 18835 | 882.8906
66 | 18837 | 882.9844
66 | 18837 | 882.9844
66 | 18838 | 883.0312
66 | 18837 | 882.9844
66 | 18837 | 882.9844
66 | 18837 | 882.9844
66 | 18837 | 882.9844
66 | 18838 | 883.0312
66 | 18836 | 882.9375
66 | 18835 | 882.8906
66 | 18834 | 882.8438
66 | 18833 | 882.7969
66 | 18832 | 882.75
66 | 18830 | 882.6562
66 | 18831 | 882.7031
66 | 18833 | 882.7969
66 | 18837 | 882.9844
66 | 18838 | 883.0312
66 | 18836 | 882.9375
66 | 18837 | 882.9844
66 | 18833 | 882.7969
66 | 18830 | 882.6562
66 | 18829 | 882.6094
66 | 18832 | 882.75
66 | 18837 | 882.9844
66 | 18840 | 883.125
66 | 18839 | 883.0781
66 | 18835 | 882.8906
66 | 18832 | 882.75
66 | 18833 | 882.7969
66 | 18840 | 883.125
66 | 18839 | 883.0781
66 | 18837 | 882.9844
66 | 18831 | 882.7031
66 | 18831 | 882.7031
66 | 18837 | 882.9844
66 | 18839 | 883.0781
66 | 18840 | 883.125
66 | 18833 | 882.7969
66 | 18829 | 882.6094
66 | 18834 | 882.8438
66 | 18842 | 883.2188
66 | 18839 | 883.0781
66 | 18834 | 882.8438
66 | 18831 | 882.7031
66 | 18840 | 883.125
66 | 18841 | 883.1719
66 | 18837 | 882.9844
66 | 18835 | 882.8906
66 | 18841 | 883.1719
66 | 18841 | 883.1719
66 | 18837 | 882.9844
66 | 18831 | 882.7031
66 | 18840 | 883.125
66 | 18840 | 883.125
66 | 18832 | 882.75
66 | 18833 | 882.7969
66 | 18840 | 883.125
66 | 18838 | 883.0312
66 | 18831 | 882.7031
66 | 18840 | 883.125
66 | 18841 | 883.1719
66 | 18835 | 882.8906
66 | 18839 | 883.0781
66 | 18842 | 883.2188
66 | 18834 | 882.8438
66 | 18838 | 883.0312
66 | 18843 | 883.2656
66 | 18833 | 882.7969
66 | 18837 | 882.9844
66 | 18841 | 883.1719
66 | 18832 | 882.75
66 | 18840 | 883.125
66 | 18840 | 883.125
66 | 18833 | 882.7969
66 | 18841 | 883.1719
66 | 18834 | 882.8438
66 | 18839 | 883.0781
66 | 18841 | 883.1719
66 | 18836 | 882.9375
66 | 18845 | 883.3594
66 | 18835 | 882.8906
66 | 18844 | 883.3125
66 | 18841 | 883.1719
66 | 18839 | 883.0781
66 | 18843 | 883.2656
66 | 18836 | 882.9375
66 | 18843 | 883.2656
66 | 18835 | 882.8906
66 | 18842 | 883.2188
66 | 18835 | 882.8906
66 | 18840 | 883.125
66 | 18840 | 883.125
66 | 18837 | 882.9844
66 | 18841 | 883.1719
66 | 18837 | 882.9844
66 | 18842 | 883.2188
66 | 18839 | 883.0781
66 | 18843 | 883.2656
66 | 18840 | 883.125
66 | 18845 | 883.3594
66 | 18840 | 883.125
66 | 18844 | 883.3125
66 | 18841 | 883.1719
66 | 18844 | 883.3125
66 | 18839 | 883.0781
66 | 18842 | 883.2188
66 | 18841 | 883.1719
66 | 18841 | 883.1719
66 | 18841 | 883.1719
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18842 | 883.2188
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18843 | 883.2656
66 | 18845 | 883.3594
66 | 18842 | 883.2188
66 | 18845 | 883.3594
66 | 18840 | 883.125
66 | 18846 | 883.4062
66 | 18840 | 883.125
66 | 18844 | 883.3125
66 | 18840 | 883.125
66 | 18845 | 883.3594
66 | 18838 | 883.0312
66 | 18845 | 883.3594
66 | 18836 | 882.9375
66 | 18843 | 883.2656
66 | 18836 | 882.9375
66 | 18840 | 883.125
66 | 18844 | 883.3125
66 | 18839 | 883.0781
66 | 18846 | 883.4062
66 | 18839 | 883.0781
66 | 18845 | 883.3594
66 | 18843 | 883.2656
66 | 18842 | 883.2188
66 | 18846 | 883.4062
66 | 18837 | 882.9844
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18838 | 883.0312
66 | 18843 | 883.2656
66 | 18840 | 883.125
66 | 18838 | 883.0312
66 | 18845 | 883.3594
66 | 18838 | 883.0312
66 | 18839 | 883.0781
66 | 18847 | 883.4531
66 | 18843 | 883.2656
66 | 18838 | 883.0312
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18838 | 883.0312
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18840 | 883.125
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18837 | 882.9844
66 | 18847 | 883.4531
66 | 18843 | 883.2656
66 | 18837 | 882.9844
66 | 18840 | 883.125
66 | 18846 | 883.4062
66 | 18844 | 883.3125
66 | 18839 | 883.0781
66 | 18842 | 883.2188
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18840 | 883.125
66 | 18842 | 883.2188
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18841 | 883.1719
66 | 18842 | 883.2188
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18843 | 883.2656
66 | 18839 | 883.0781
66 | 18842 | 883.2188
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18846 | 883.4062
66 | 18839 | 883.0781
66 | 18840 | 883.125
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18849 | 883.5469
66 | 18843 | 883.2656
66 | 18839 | 883.0781
66 | 18840 | 883.125
66 | 18844 | 883.3125
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18844 | 883.3125
66 | 18839 | 883.0781
66 | 18837 | 882.9844
66 | 18839 | 883.0781
66 | 18843 | 883.2656
66 | 18846 | 883.4062
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18843 | 883.2656
66 | 18839 | 883.0781
66 | 18840 | 883.125
66 | 18843 | 883.2656
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18846 | 883.4062
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18845 | 883.3594
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18847 | 883.4531
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18846 | 883.4062
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18845 | 883.3594
66 | 18844 | 883.3125
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18848 | 883.5
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18845 | 883.3594
66 | 18846 | 883.4062
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18845 | 883.3594
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18845 | 883.3594
66 | 18843 | 883.2656
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18844 | 883.3125
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18848 | 883.5
66 | 18844 | 883.3125
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18847 | 883.4531
66 | 18845 | 883.3594
66 | 18854 | 883.7812
66 | 18855 | 883.8281
66 | 18847 | 883.4531
66 | 18851 | 883.6406
66 | 18855 | 883.8281
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18855 | 883.8281
66 | 18850 | 883.5938
66 | 18844 | 883.3125
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18843 | 883.2656
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18845 | 883.3594
66 | 18853 | 883.7344
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18846 | 883.4062
66 | 18855 | 883.8281
66 | 18851 | 883.6406
66 | 18846 | 883.4062
66 | 18855 | 883.8281
66 | 18846 | 883.4062
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18844 | 883.3125
66 | 18853 | 883.7344
66 | 18846 | 883.4062
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18843 | 883.2656
66 | 18854 | 883.7812
66 | 18846 | 883.4062
66 | 18852 | 883.6875
66 | 18853 | 883.7344
66 | 18844 | 883.3125
66 | 18854 | 883.7812
66 | 18847 | 883.4531
66 | 18854 | 883.7812
66 | 18847 | 883.4531
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18845 | 883.3594
66 | 18853 | 883.7344
66 | 18844 | 883.3125
66 | 18853 | 883.7344
66 | 18845 | 883.3594
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18848 | 883.5
66 | 18853 | 883.7344
66 | 18845 | 883.3594
66 | 18854 | 883.7812
66 | 18847 | 883.4531
66 | 18856 | 883.875
66 | 18845 | 883.3594
66 | 18855 | 883.8281
66 | 18846 | 883.4062
66 | 18854 | 883.7812
66 | 18844 | 883.3125
66 | 18853 | 883.7344
66 | 18846 | 883.4062
66 | 18852 | 883.6875
66 | 18846 | 883.4062
66 | 18853 | 883.7344
66 | 18846 | 883.4062
66 | 18853 | 883.7344
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18849 | 883.5469
66 | 18855 | 883.8281
66 | 18849 | 883.5469
66 | 18854 | 883.7812
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18849 | 883.5469
66 | 18851 | 883.6406
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18849 | 883.5469
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18848 | 883.5
66 | 18856 | 883.875
66 | 18849 | 883.5469
66 | 18853 | 883.7344
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18845 | 883.3594
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18854 | 883.7812
66 | 18844 | 883.3125
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18849 | 883.5469
66 | 18855 | 883.8281
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18856 | 883.875
66 | 18846 | 883.4062
66 | 18853 | 883.7344
66 | 18854 | 883.7812
66 | 18847 | 883.4531
66 | 18854 | 883.7812
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18849 | 883.5469
66 | 18855 | 883.8281
66 | 18851 | 883.6406
66 | 18849 | 883.5469
66 | 18854 | 883.7812
66 | 18852 | 883.6875
66 | 18850 | 883.5938
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18846 | 883.4062
66 | 18853 | 883.7344
66 | 18855 | 883.8281
66 | 18850 | 883.5938
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18855 | 883.8281
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18858 | 883.9688
66 | 18855 | 883.8281
66 | 18849 | 883.5469
66 | 18854 | 883.7812
66 | 18857 | 883.9219
66 | 18853 | 883.7344
66 | 18847 | 883.4531
66 | 18849 | 883.5469
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18850 | 883.5938
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18854 | 883.7812
66 | 18849 | 883.5469
66 | 18852 | 883.6875
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18846 | 883.4062
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18851 | 883.6406
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18851 | 883.6406
66 | 18854 | 883.7812
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18852 | 883.6875
66 | 18848 | 883.5
66 | 18848 | 883.5
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18855 | 883.8281
66 | 18858 | 883.9688
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18853 | 883.7344
66 | 18851 | 883.6406
66 | 18849 | 883.5469
66 | 18848 | 883.5
66 | 18850 | 883.5938
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18852 | 883.6875
66 | 18851 | 883.6406
66 | 18849 | 883.5469
66 | 18853 | 883.7344
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18856 | 883.875
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18846 | 883.4062
66 | 18850 | 883.5938
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18855 | 883.8281
66 | 18851 | 883.6406
66 | 18850 | 883.5938
66 | 18855 | 883.8281
66 | 18858 | 883.9688
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18856 | 883.875
66 | 18850 | 883.5938
66 | 18852 | 883.6875
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18857 | 883.9219
66 | 18853 | 883.7344
66 | 18849 | 883.5469
66 | 18857 | 883.9219
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18860 | 884.0625
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18852 | 883.6875
66 | 18852 | 883.6875
66 | 18857 | 883.9219
66 | 18860 | 884.0625
66 | 18853 | 883.7344
66 | 18850 | 883.5938
66 | 18856 | 883.875
66 | 18860 | 884.0625
66 | 18856 | 883.875
66 | 18851 | 883.6406
66 | 18856 | 883.875
66 | 18860 | 884.0625
66 | 18859 | 884.0156
66 | 18850 | 883.5938
66 | 18858 | 883.9688
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18850 | 883.5938
66 | 18856 | 883.875
66 | 18859 | 884.0156
66 | 18852 | 883.6875
66 | 18855 | 883.8281
66 | 18859 | 884.0156
66 | 18855 | 883.8281
66 | 18853 | 883.7344
66 | 18861 | 884.1094
66 | 18859 | 884.0156
66 | 18853 | 883.7344
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18853 | 883.7344
66 | 18860 | 884.0625
66 | 18860 | 884.0625
66 | 18852 | 883.6875
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18851 | 883.6406
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18852 | 883.6875
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18853 | 883.7344
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18855 | 883.8281
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18855 | 883.8281
66 | 18863 | 884.2031
66 | 18855 | 883.8281
66 | 18862 | 884.1562
66 | 18858 | 883.9688
66 | 18856 | 883.875
66 | 18862 | 884.1562
66 | 18855 | 883.8281
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18856 | 883.875
66 | 18863 | 884.2031
66 | 18853 | 883.7344
66 | 18864 | 884.25
66 | 18860 | 884.0625
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18856 | 883.875
66 | 18863 | 884.2031
66 | 18857 | 883.9219
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18855 | 883.8281
66 | 18859 | 884.0156
66 | 18855 | 883.8281
66 | 18863 | 884.2031
66 | 18854 | 883.7812
66 | 18864 | 884.25
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18862 | 884.1562
66 | 18858 | 883.9688
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18860 | 884.0625
66 | 18855 | 883.8281
66 | 18861 | 884.1094
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18858 | 883.9688
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18856 | 883.875
66 | 18863 | 884.2031
66 | 18853 | 883.7344
66 | 18863 | 884.2031
66 | 18858 | 883.9688
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18857 | 883.9219
66 | 18867 | 884.3906
66 | 18861 | 884.1094
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18856 | 883.875
66 | 18865 | 884.2969
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18868 | 884.4375
66 | 18866 | 884.3438
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18863 | 884.2031
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18862 | 884.1562
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18856 | 883.875
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18866 | 884.3438
66 | 18867 | 884.3906
66 | 18858 | 883.9688
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18857 | 883.9219
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18855 | 883.8281
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18857 | 883.9219
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18856 | 883.875
66 | 18861 | 884.1094
66 | 18865 | 884.2969
66 | 18863 | 884.2031
66 | 18857 | 883.9219
66 | 18862 | 884.1562
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18857 | 883.9219
66 | 18863 | 884.2031
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18864 | 884.25
66 | 18866 | 884.3438
66 | 18859 | 884.0156
66 | 18856 | 883.875
66 | 18859 | 884.0156
66 | 18864 | 884.25
66 | 18863 | 884.2031
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18863 | 884.2031
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18867 | 884.3906
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18860 | 884.0625
66 | 18867 | 884.3906
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18862 | 884.1562
66 | 18856 | 883.875
66 | 18858 | 883.9688
66 | 18861 | 884.1094
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18859 | 884.0156
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18863 | 884.2031
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18861 | 884.1094
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18861 | 884.1094
66 | 18860 | 884.0625
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18858 | 883.9688
66 | 18859 | 884.0156
66 | 18861 | 884.1094
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18857 | 883.9219
66 | 18859 | 884.0156
66 | 18857 | 883.9219
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18856 | 883.875
66 | 18856 | 883.875
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18857 | 883.9219
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18861 | 884.1094
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18865 | 884.2969
66 | 18867 | 884.3906
66 | 18868 | 884.4375
66 | 18868 | 884.4375
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18863 | 884.2031
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18858 | 883.9688
66 | 18870 | 884.5312
66 | 18864 | 884.25
66 | 18864 | 884.25
66 | 18865 | 884.2969
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18858 | 883.9688
66 | 18869 | 884.4844
66 | 18859 | 884.0156
66 | 18860 | 884.0625
66 | 18863 | 884.2031
66 | 18862 | 884.1562
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18857 | 883.9219
66 | 18854 | 883.7812
66 | 18856 | 883.875
66 | 18863 | 884.2031
66 | 18864 | 884.25
66 | 18867 | 884.3906
66 | 18866 | 884.3438
66 | 18863 | 884.2031
66 | 18855 | 883.8281
66 | 18859 | 884.0156
66 | 18866 | 884.3438
66 | 18864 | 884.25
66 | 18862 | 884.1562
66 | 18859 | 884.0156
66 | 18855 | 883.8281
66 | 18856 | 883.875
66 | 18862 | 884.1562
66 | 18863 | 884.2031
66 | 18863 | 884.2031
66 | 18860 | 884.0625
66 | 18858 | 883.9688
66 | 18857 | 883.9219
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18864 | 884.25
66 | 18861 | 884.1094
66 | 18858 | 883.9688
66 | 18863 | 884.2031
66 | 18866 | 884.3438
66 | 18863 | 884.2031
66 | 18859 | 884.0156
66 | 18855 | 883.8281
66 | 18852 | 883.6875
66 | 18854 | 883.7812
66 | 18858 | 883.9688
66 | 18862 | 884.1562
66 | 18866 | 884.3438
66 | 18866 | 884.3438
66 | 18869 | 884.4844
66 | 18867 | 884.3906
66 | 18865 | 884.2969
66 | 18859 | 884.0156
66 | 18859 | 884.0156
66 | 18865 | 884.2969
66 | 18865 | 884.2969
66 | 18860 | 884.0625
66 | 18855 | 883.8281
66 | 18848 | 883.5
66 | 18851 | 883.6406
66 | 18855 | 883.8281
66 | 18854 | 883.7812
66 | 18853 | 883.7344
66 | 18848 | 883.5
66 | 18846 | 883.4062
66 | 18851 | 883.6406
66 | 18853 | 883.7344
66 | 18853 | 883.7344
66 | 18847 | 883.4531
66 | 18840 | 883.125
66 | 18836 | 882.9375
66 | 18833 | 882.7969
66 | 18834 | 882.8438
66 | 18832 | 882.75
66 | 18822 | 882.2812
66 | 18792 | 880.875
66 | 18746 | 878.7188
66 | 18677 | 875.4844
66 | 18582 | 871.0312
66 | 18546 | 869.3438
66 | 18512 | 867.75
66 | 18416 | 863.25
66 | 18342 | 859.7812
66 | 18296 | 857.625
66 | 18262 | 856.0312
66 | 18232 | 854.625
66 | 18189 | 852.6094
66 | 18141 | 850.3594
66 | 18096 | 848.25
66 | 18053 | 846.2344
66 | 18009 | 844.1719
66 | 17967 | 842.2031
66 | 17943 | 841.0781
66 | 17927 | 840.3281
66 | 17915 | 839.7656
66 | 17887 | 838.4531
66 | 17866 | 837.4688
66 | 17845 | 836.4844
66 | 17828 | 835.6875
66 | 17797 | 834.2344
66 | 17770 | 832.9688
66 | 17744 | 831.75
66 | 17719 | 830.5781
66 | 17678 | 828.6562
66 | 17643 | 827.0156
66 | 17605 | 825.2344
66 | 17570 | 823.5938
66 | 17536 | 822
66 | 17509 | 820.7344
66 | 17488 | 819.75
66 | 17474 | 819.0938
66 | 17463 | 818.5781
66 | 17455 | 818.2031
66 | 17449 | 817.9219
66 | 17435 | 817.2656
66 | 17422 | 816.6562
66 | 17411 | 816.1406
66 | 17394 | 815.3438
66 | 17383 | 814.8281
66 | 17363 | 813.8906
66 | 17342 | 812.9062
66 | 17326 | 812.1562
66 | 17312 | 811.5
66 | 17293 | 810.6094
66 | 17273 | 809.6719
66 | 17251 | 808.6406
66 | 17235 | 807.8906
66 | 17223 | 807.3281
66 | 17211 | 806.7656
66 | 17199 | 806.2031
66 | 17195 | 806.0156
66 | 17184 | 805.5
66 | 17173 | 804.9844
66 | 17162 | 804.4688
66 | 17152 | 804
66 | 17129 | 802.9219
66 | 17120 | 802.5
66 | 17103 | 801.7031
66 | 17087 | 800.9531
66 | 17074 | 800.3438
66 | 17059 | 799.6406
66 | 17052 | 799.3125
66 | 17056 | 799.5
66 | 17045 | 798.9844
66 | 17048 | 799.125
66 | 17044 | 798.9375
66 | 17043 | 798.8906
66 | 17038 | 798.6562
66 | 17029 | 798.2344
66 | 17020 | 797.8125
66 | 17011 | 797.3906
66 | 16998 | 796.7812
66 | 16982 | 796.0312
66 | 16966 | 795.2812
66 | 16951 | 794.5781
66 | 16937 | 793.9219
66 | 16924 | 793.3125
66 | 16916 | 792.9375
66 | 16916 | 792.9375
66 | 16905 | 792.4219
66 | 16911 | 792.7031
66 | 16905 | 792.4219
66 | 16909 | 792.6094
66 | 16906 | 792.4688
66 | 16905 | 792.4219
66 | 16898 | 792.0938
66 | 16891 | 791.7656
66 | 16880 | 791.25
66 | 16873 | 790.9219
66 | 16860 | 790.3125
66 | 16854 | 790.0312
66 | 16841 | 789.4219
66 | 16844 | 789.5625
66 | 16831 | 788.9531
66 | 16842 | 789.4688
66 | 16832 | 789
66 | 16841 | 789.4219
66 | 16844 | 789.5625
66 | 16849 | 789.7969
66 | 16850 | 789.8438
66 | 16849 | 789.7969
66 | 16846 | 789.6562
66 | 16839 | 789.3281
66 | 16830 | 788.9062
66 | 16819 | 788.3906
66 | 16808 | 787.875
66 | 16797 | 787.3594
66 | 16787 | 786.8906
66 | 16780 | 786.5625
66 | 16775 | 786.3281
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16774 | 786.2812
66 | 16766 | 785.9062
66 | 16758 | 785.5312
66 | 16748 | 785.0625
66 | 16739 | 784.6406
66 | 16729 | 784.1719
66 | 16720 | 783.75
66 | 16715 | 783.5156
66 | 16710 | 783.2812
66 | 16711 | 783.3281
66 | 16712 | 783.375
66 | 16715 | 783.5156
66 | 16720 | 783.75
66 | 16725 | 783.9844
66 | 16728 | 784.125
66 | 16728 | 784.125
66 | 16727 | 784.0781
66 | 16723 | 783.8906
66 | 16713 | 783.4219
66 | 16706 | 783.0938
66 | 16694 | 782.5312
66 | 16686 | 782.1562
66 | 16675 | 781.6406
66 | 16670 | 781.4062
66 | 16666 | 781.2188
66 | 16669 | 781.3594
66 | 16669 | 781.3594
66 | 16676 | 781.6875
66 | 16674 | 781.5938
66 | 16683 | 782.0156
66 | 16679 | 781.8281
66 | 16687 | 782.2031
66 | 16677 | 781.7344
66 | 16680 | 781.875
66 | 16665 | 781.1719
66 | 16665 | 781.1719
66 | 16648 | 780.375
66 | 16637 | 779.8594
66 | 16632 | 779.625
66 | 16627 | 779.3906
66 | 16624 | 779.25
66 | 16623 | 779.2031
66 | 16622 | 779.1562
66 | 16623 | 779.2031
66 | 16626 | 779.3438
66 | 16630 | 779.5312
66 | 16634 | 779.7188
66 | 16636 | 779.8125
66 | 16638 | 779.9062
66 | 16636 | 779.8125
66 | 16632 | 779.625
66 | 16624 | 779.25
66 | 16618 | 778.9688
66 | 16611 | 778.6406
66 | 16596 | 777.9375
66 | 16591 | 777.7031
66 | 16593 | 777.7969
66 | 16591 | 777.7031
66 | 16597 | 777.9844
66 | 16598 | 778.0312
66 | 16604 | 778.3125
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16616 | 778.875
66 | 16615 | 778.8281
66 | 16611 | 778.6406
66 | 16607 | 778.4531
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16574 | 776.9062
66 | 16566 | 776.5312
66 | 16562 | 776.3438
66 | 16561 | 776.2969
66 | 16572 | 776.8125
66 | 16569 | 776.6719
66 | 16573 | 776.8594
66 | 16587 | 777.5156
66 | 16582 | 777.2812
66 | 16586 | 777.4688
66 | 16580 | 777.1875
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16569 | 776.6719
66 | 16554 | 775.9688
66 | 16549 | 775.7344
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16541 | 775.3594
66 | 16546 | 775.5938
66 | 16553 | 775.9219
66 | 16556 | 776.0625
66 | 16563 | 776.3906
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16553 | 775.9219
66 | 16548 | 775.6875
66 | 16536 | 775.125
66 | 16533 | 774.9844
66 | 16520 | 774.375
66 | 16521 | 774.4219
66 | 16518 | 774.2812
66 | 16516 | 774.1875
66 | 16520 | 774.375
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16536 | 775.125
66 | 16533 | 774.9844
66 | 16537 | 775.1719
66 | 16536 | 775.125
66 | 16535 | 775.0781
66 | 16531 | 774.8906
66 | 16527 | 774.7031
66 | 16518 | 774.2812
66 | 16510 | 773.9062
66 | 16505 | 773.6719
66 | 16496 | 773.25
66 | 16495 | 773.2031
66 | 16490 | 772.9688
66 | 16498 | 773.3438
66 | 16494 | 773.1562
66 | 16509 | 773.8594
66 | 16504 | 773.625
66 | 16515 | 774.1406
66 | 16519 | 774.3281
66 | 16519 | 774.3281
66 | 16520 | 774.375
66 | 16518 | 774.2812
66 | 16512 | 774
66 | 16504 | 773.625
66 | 16496 | 773.25
66 | 16487 | 772.8281
66 | 16481 | 772.5469
66 | 16475 | 772.2656
66 | 16474 | 772.2188
66 | 16474 | 772.2188
66 | 16476 | 772.3125
66 | 16483 | 772.6406
66 | 16495 | 773.2031
66 | 16500 | 773.4375
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16503 | 773.5781
66 | 16499 | 773.3906
66 | 16496 | 773.25
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16471 | 772.0781
66 | 16468 | 771.9375
66 | 16456 | 771.375
66 | 16459 | 771.5156
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16458 | 771.4688
66 | 16470 | 772.0312
66 | 16467 | 771.8906
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16477 | 772.3594
66 | 16480 | 772.5
66 | 16476 | 772.3125
66 | 16471 | 772.0781
66 | 16463 | 771.7031
66 | 16455 | 771.3281
66 | 16447 | 770.9531
66 | 16439 | 770.5781
66 | 16435 | 770.3906
66 | 16432 | 770.25
66 | 16431 | 770.2031
66 | 16435 | 770.3906
66 | 16439 | 770.5781
66 | 16442 | 770.7188
66 | 16447 | 770.9531
66 | 16454 | 771.2812
66 | 16457 | 771.4219
66 | 16462 | 771.6562
66 | 16459 | 771.5156
66 | 16456 | 771.375
66 | 16450 | 771.0938
66 | 16443 | 770.7656
66 | 16436 | 770.4375
66 | 16425 | 769.9219
66 | 16424 | 769.875
66 | 16421 | 769.7344
66 | 16419 | 769.6406
66 | 16420 | 769.6875
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16436 | 770.4375
66 | 16442 | 770.7188
66 | 16449 | 771.0469
66 | 16451 | 771.1406
66 | 16453 | 771.2344
66 | 16447 | 770.9531
66 | 16442 | 770.7188
66 | 16437 | 770.4844
66 | 16427 | 770.0156
66 | 16424 | 769.875
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16403 | 768.8906
66 | 16403 | 768.8906
66 | 16406 | 769.0312
66 | 16409 | 769.1719
66 | 16415 | 769.4531
66 | 16421 | 769.7344
66 | 16426 | 769.9688
66 | 16431 | 770.2031
66 | 16432 | 770.25
66 | 16434 | 770.3438
66 | 16430 | 770.1562
66 | 16426 | 769.9688
66 | 16426 | 769.9688
66 | 16412 | 769.3125
66 | 16407 | 769.0781
66 | 16397 | 768.6094
66 | 16397 | 768.6094
66 | 16390 | 768.2812
66 | 16396 | 768.5625
66 | 16393 | 768.4219
66 | 16405 | 768.9844
66 | 16408 | 769.125
66 | 16412 | 769.3125
66 | 16416 | 769.5
66 | 16427 | 770.0156
66 | 16428 | 770.0625
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16408 | 769.125
66 | 16400 | 768.75
66 | 16390 | 768.2812
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16385 | 768.0469
66 | 16391 | 768.3281
66 | 16392 | 768.375
66 | 16404 | 768.9375
66 | 16402 | 768.8438
66 | 16414 | 769.4062
66 | 16411 | 769.2656
66 | 16421 | 769.7344
66 | 16411 | 769.2656
66 | 16415 | 769.4531
66 | 16403 | 768.8906
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16385 | 768.0469
66 | 16382 | 767.9062
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16394 | 768.4688
66 | 16401 | 768.7969
66 | 16406 | 769.0312
66 | 16410 | 769.2188
66 | 16408 | 769.125
66 | 16408 | 769.125
66 | 16400 | 768.75
66 | 16393 | 768.4219
66 | 16388 | 768.1875
66 | 16379 | 767.7656
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16362 | 766.9688
66 | 16367 | 767.2031
66 | 16362 | 766.9688
66 | 16373 | 767.4844
66 | 16371 | 767.3906
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16396 | 768.5625
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16396 | 768.5625
66 | 16392 | 768.375
66 | 16384 | 768
66 | 16377 | 767.6719
66 | 16370 | 767.3438
66 | 16363 | 767.0156
66 | 16357 | 766.7344
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16354 | 766.5938
66 | 16361 | 766.9219
66 | 16374 | 767.5312
66 | 16380 | 767.8125
66 | 16377 | 767.6719
66 | 16380 | 767.8125
66 | 16391 | 768.3281
66 | 16380 | 767.8125
66 | 16385 | 768.0469
66 | 16371 | 767.3906
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16352 | 766.5
66 | 16346 | 766.2188
66 | 16342 | 766.0312
66 | 16342 | 766.0312
66 | 16345 | 766.1719
66 | 16349 | 766.3594
66 | 16354 | 766.5938
66 | 16361 | 766.9219
66 | 16367 | 767.2031
66 | 16371 | 767.3906
66 | 16374 | 767.5312
66 | 16377 | 767.6719
66 | 16384 | 768
66 | 16372 | 767.4375
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16350 | 766.4062
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16352 | 766.5
66 | 16359 | 766.8281
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16371 | 767.3906
66 | 16365 | 767.1094
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16340 | 765.9375
66 | 16336 | 765.75
66 | 16333 | 765.6094
66 | 16332 | 765.5625
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16356 | 766.6875
66 | 16358 | 766.7812
66 | 16353 | 766.5469
66 | 16356 | 766.6875
66 | 16341 | 765.9844
66 | 16343 | 766.0781
66 | 16326 | 765.2812
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16315 | 764.7656
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16332 | 765.5625
66 | 16343 | 766.0781
66 | 16348 | 766.3125
66 | 16354 | 766.5938
66 | 16354 | 766.5938
66 | 16355 | 766.6406
66 | 16354 | 766.5938
66 | 16348 | 766.3125
66 | 16341 | 765.9844
66 | 16334 | 765.6562
66 | 16325 | 765.2344
66 | 16317 | 764.8594
66 | 16310 | 764.5312
66 | 16312 | 764.625
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16317 | 764.8594
66 | 16331 | 765.5156
66 | 16329 | 765.4219
66 | 16345 | 766.1719
66 | 16339 | 765.8906
66 | 16350 | 766.4062
66 | 16341 | 765.9844
66 | 16348 | 766.3125
66 | 16334 | 765.6562
66 | 16335 | 765.7031
66 | 16322 | 765.0938
66 | 16317 | 764.8594
66 | 16313 | 764.6719
66 | 16306 | 764.3438
66 | 16303 | 764.2031
66 | 16302 | 764.1562
66 | 16306 | 764.3438
66 | 16310 | 764.5312
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16331 | 765.5156
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16310 | 764.5312
66 | 16304 | 764.25
66 | 16298 | 763.9688
66 | 16295 | 763.8281
66 | 16293 | 763.7344
66 | 16295 | 763.8281
66 | 16299 | 764.0156
66 | 16311 | 764.5781
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16323 | 765.1406
66 | 16332 | 765.5625
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16312 | 764.625
66 | 16313 | 764.6719
66 | 16296 | 763.875
66 | 16301 | 764.1094
66 | 16288 | 763.5
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16295 | 763.8281
66 | 16299 | 764.0156
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16330 | 765.4688
66 | 16325 | 765.2344
66 | 16312 | 764.625
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16295 | 763.8281
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16305 | 764.2969
66 | 16303 | 764.2031
66 | 16319 | 764.9531
66 | 16312 | 764.625
66 | 16325 | 765.2344
66 | 16318 | 764.9062
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16295 | 763.8281
66 | 16295 | 763.8281
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16280 | 763.125
66 | 16285 | 763.3594
66 | 16291 | 763.6406
66 | 16297 | 763.9219
66 | 16303 | 764.2031
66 | 16308 | 764.4375
66 | 16312 | 764.625
66 | 16312 | 764.625
66 | 16311 | 764.5781
66 | 16308 | 764.4375
66 | 16304 | 764.25
66 | 16299 | 764.0156
66 | 16290 | 763.5938
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16288 | 763.5
66 | 16288 | 763.5
66 | 16295 | 763.8281
66 | 16307 | 764.3906
66 | 16306 | 764.3438
66 | 16315 | 764.7656
66 | 16307 | 764.3906
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16278 | 763.0312
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16278 | 763.0312
66 | 16284 | 763.3125
66 | 16291 | 763.6406
66 | 16296 | 763.875
66 | 16301 | 764.1094
66 | 16303 | 764.2031
66 | 16303 | 764.2031
66 | 16299 | 764.0156
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16276 | 762.9375
66 | 16272 | 762.75
66 | 16267 | 762.5156
66 | 16270 | 762.6562
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16278 | 763.0312
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16297 | 763.9219
66 | 16305 | 764.2969
66 | 16300 | 764.0625
66 | 16305 | 764.2969
66 | 16292 | 763.6875
66 | 16294 | 763.7812
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16269 | 762.6094
66 | 16279 | 763.0781
66 | 16281 | 763.1719
66 | 16287 | 763.4531
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16299 | 764.0156
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16277 | 762.9844
66 | 16270 | 762.6562
66 | 16266 | 762.4688
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16268 | 762.5625
66 | 16273 | 762.7969
66 | 16280 | 763.125
66 | 16285 | 763.3594
66 | 16289 | 763.5469
66 | 16293 | 763.7344
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16280 | 763.125
66 | 16268 | 762.5625
66 | 16263 | 762.3281
66 | 16258 | 762.0938
66 | 16257 | 762.0469
66 | 16250 | 761.7188
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16287 | 763.4531
66 | 16293 | 763.7344
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16249 | 761.6719
66 | 16254 | 761.9062
66 | 16260 | 762.1875
66 | 16262 | 762.2812
66 | 16273 | 762.7969
66 | 16279 | 763.0781
66 | 16285 | 763.3594
66 | 16287 | 763.4531
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16275 | 762.8906
66 | 16269 | 762.6094
66 | 16260 | 762.1875
66 | 16253 | 761.8594
66 | 16248 | 761.625
66 | 16245 | 761.4844
66 | 16245 | 761.4844
66 | 16248 | 761.625
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16266 | 762.4688
66 | 16273 | 762.7969
66 | 16278 | 763.0312
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16270 | 762.6562
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16253 | 761.8594
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16269 | 762.6094
66 | 16275 | 762.8906
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16272 | 762.75
66 | 16267 | 762.5156
66 | 16263 | 762.3281
66 | 16252 | 761.8125
66 | 16251 | 761.7656
66 | 16239 | 761.2031
66 | 16245 | 761.4844
66 | 16236 | 761.0625
66 | 16247 | 761.5781
66 | 16253 | 761.8594
66 | 16252 | 761.8125
66 | 16261 | 762.2344
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16271 | 762.7031
66 | 16265 | 762.4219
66 | 16263 | 762.3281
66 | 16255 | 761.9531
66 | 16250 | 761.7188
66 | 16244 | 761.4375
66 | 16240 | 761.25
66 | 16238 | 761.1562
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16268 | 762.5625
66 | 16273 | 762.7969
66 | 16277 | 762.9844
66 | 16269 | 762.6094
66 | 16261 | 762.2344
66 | 16256 | 762
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16238 | 761.1562
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16243 | 761.3906
66 | 16248 | 761.625
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16270 | 762.6562
66 | 16267 | 762.5156
66 | 16274 | 762.8438
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16253 | 761.8594
66 | 16248 | 761.625
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16234 | 760.9688
66 | 16229 | 760.7344
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16242 | 761.3438
66 | 16247 | 761.5781
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16265 | 762.4219
66 | 16260 | 762.1875
66 | 16257 | 762.0469
66 | 16252 | 761.8125
66 | 16244 | 761.4375
66 | 16240 | 761.25
66 | 16238 | 761.1562
66 | 16231 | 760.8281
66 | 16228 | 760.6875
66 | 16232 | 760.875
66 | 16236 | 761.0625
66 | 16242 | 761.3438
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16236 | 761.0625
66 | 16231 | 760.8281
66 | 16230 | 760.7812
66 | 16225 | 760.5469
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16235 | 761.0156
66 | 16240 | 761.25
66 | 16252 | 761.8125
66 | 16249 | 761.6719
66 | 16262 | 762.2812
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16241 | 761.2969
66 | 16240 | 761.25
66 | 16227 | 760.6406
66 | 16225 | 760.5469
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16225 | 760.5469
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16255 | 761.9531
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16256 | 762
66 | 16247 | 761.5781
66 | 16241 | 761.2969
66 | 16235 | 761.0156
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16217 | 760.1719
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16224 | 760.5
66 | 16235 | 761.0156
66 | 16240 | 761.25
66 | 16244 | 761.4375
66 | 16250 | 761.7188
66 | 16253 | 761.8594
66 | 16253 | 761.8594
66 | 16250 | 761.7188
66 | 16247 | 761.5781
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16226 | 760.5938
66 | 16220 | 760.3125
66 | 16215 | 760.0781
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16240 | 761.25
66 | 16245 | 761.4844
66 | 16249 | 761.6719
66 | 16249 | 761.6719
66 | 16254 | 761.9062
66 | 16246 | 761.5312
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16231 | 760.8281
66 | 16220 | 760.3125
66 | 16215 | 760.0781
66 | 16212 | 759.9375
66 | 16214 | 760.0312
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16231 | 760.8281
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16246 | 761.5312
66 | 16249 | 761.6719
66 | 16244 | 761.4375
66 | 16242 | 761.3438
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16232 | 760.875
66 | 16237 | 761.1094
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16244 | 761.4375
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16210 | 759.8438
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16232 | 760.875
66 | 16238 | 761.1562
66 | 16241 | 761.2969
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16225 | 760.5469
66 | 16219 | 760.2656
66 | 16214 | 760.0312
66 | 16208 | 759.75
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16224 | 760.5
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16241 | 761.2969
66 | 16240 | 761.25
66 | 16240 | 761.25
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16213 | 759.9844
66 | 16207 | 759.7031
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16238 | 761.1562
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16201 | 759.4219
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16225 | 760.5469
66 | 16230 | 760.7812
66 | 16237 | 761.1094
66 | 16237 | 761.1094
66 | 16237 | 761.1094
66 | 16235 | 761.0156
66 | 16227 | 760.6406
66 | 16220 | 760.3125
66 | 16214 | 760.0312
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16219 | 760.2656
66 | 16230 | 760.7812
66 | 16233 | 760.9219
66 | 16235 | 761.0156
66 | 16235 | 761.0156
66 | 16234 | 760.9688
66 | 16230 | 760.7812
66 | 16223 | 760.4531
66 | 16218 | 760.2188
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16204 | 759.5625
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16226 | 760.5938
66 | 16229 | 760.7344
66 | 16233 | 760.9219
66 | 16233 | 760.9219
66 | 16233 | 760.9219
66 | 16228 | 760.6875
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16202 | 759.4688
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16230 | 760.7812
66 | 16232 | 760.875
66 | 16235 | 761.0156
66 | 16230 | 760.7812
66 | 16223 | 760.4531
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16192 | 759
66 | 16197 | 759.2344
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16225 | 760.5469
66 | 16230 | 760.7812
66 | 16232 | 760.875
66 | 16228 | 760.6875
66 | 16223 | 760.4531
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16196 | 759.1875
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16225 | 760.5469
66 | 16229 | 760.7344
66 | 16227 | 760.6406
66 | 16226 | 760.5938
66 | 16220 | 760.3125
66 | 16215 | 760.0781
66 | 16209 | 759.7969
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16226 | 760.5938
66 | 16228 | 760.6875
66 | 16224 | 760.5
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16194 | 759.0938
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16224 | 760.5
66 | 16227 | 760.6406
66 | 16225 | 760.5469
66 | 16220 | 760.3125
66 | 16215 | 760.0781
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16219 | 760.2656
66 | 16225 | 760.5469
66 | 16225 | 760.5469
66 | 16219 | 760.2656
66 | 16215 | 760.0781
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16220 | 760.3125
66 | 16223 | 760.4531
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16191 | 758.9531
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16192 | 759
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16219 | 760.2656
66 | 16219 | 760.2656
66 | 16214 | 760.0312
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16196 | 759.1875
66 | 16184 | 758.625
66 | 16202 | 759.4688
66 | 16182 | 758.5312
66 | 16214 | 760.0312
66 | 16194 | 759.0938
66 | 16187 | 758.7656
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16194 | 759.0938
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16181 | 758.4844
66 | 16180 | 758.4375
66 | 16181 | 758.4844
66 | 16182 | 758.5312
66 | 16190 | 758.9062
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16215 | 760.0781
66 | 16209 | 759.7969
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16188 | 758.8125
66 | 16197 | 759.2344
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16211 | 759.8906
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16212 | 759.9375
66 | 16214 | 760.0312
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16197 | 759.2344
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16188 | 758.8125
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16188 | 758.8125
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16181 | 758.4844
66 | 16184 | 758.625
66 | 16194 | 759.0938
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16193 | 759.0469
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16188 | 758.8125
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16174 | 758.1562
66 | 16180 | 758.4375
66 | 16190 | 758.9062
66 | 16196 | 759.1875
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16188 | 758.8125
66 | 16183 | 758.5781
66 | 16175 | 758.2031
66 | 16172 | 758.0625
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16174 | 758.1562
66 | 16179 | 758.3906
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16190 | 758.9062
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16173 | 758.1094
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16188 | 758.8125
66 | 16194 | 759.0938
66 | 16201 | 759.4219
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16191 | 758.9531
66 | 16187 | 758.7656
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16182 | 758.5312
66 | 16191 | 758.9531
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16195 | 759.1406
66 | 16188 | 758.8125
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16197 | 759.2344
66 | 16201 | 759.4219
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16172 | 758.0625
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16171 | 758.0156
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16187 | 758.7656
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16170 | 757.9688
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16170 | 757.9688
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16188 | 758.8125
66 | 16183 | 758.5781
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16168 | 757.875
66 | 16172 | 758.0625
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16193 | 759.0469
66 | 16197 | 759.2344
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16189 | 758.8594
66 | 16183 | 758.5781
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16187 | 758.7656
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16171 | 758.0156
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16181 | 758.4844
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16186 | 758.7188
66 | 16181 | 758.4844
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16160 | 757.5
66 | 16155 | 757.2656
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16172 | 758.0625
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16160 | 757.5
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16189 | 758.8594
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16169 | 757.9219
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16170 | 757.9688
66 | 16167 | 757.8281
66 | 16160 | 757.5
66 | 16153 | 757.1719
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16161 | 757.5469
66 | 16169 | 757.9219
66 | 16175 | 758.2031
66 | 16178 | 758.3438
66 | 16181 | 758.4844
66 | 16190 | 758.9062
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16183 | 758.5781
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16167 | 757.8281
66 | 16176 | 758.25
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16153 | 757.1719
66 | 16152 | 757.125
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16170 | 757.9688
66 | 16177 | 758.2969
66 | 16183 | 758.5781
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16147 | 756.8906
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16149 | 756.9844
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16151 | 757.0781
66 | 16151 | 757.0781
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16167 | 757.8281
66 | 16174 | 758.1562
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16152 | 757.125
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16170 | 757.9688
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16159 | 757.4531
66 | 16155 | 757.2656
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16171 | 758.0156
66 | 16176 | 758.25
66 | 16179 | 758.3906
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16175 | 758.2031
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16161 | 757.5469
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16161 | 757.5469
66 | 16157 | 757.3594
66 | 16153 | 757.1719
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16171 | 758.0156
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16141 | 756.6094
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16178 | 758.3438
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16157 | 757.3594
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16136 | 756.375
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16173 | 758.1094
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16177 | 758.2969
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16168 | 757.875
66 | 16174 | 758.1562
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16155 | 757.2656
66 | 16154 | 757.2188
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16171 | 758.0156
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16166 | 757.7812
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16175 | 758.2031
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16144 | 756.75
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16170 | 757.9688
66 | 16177 | 758.2969
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16173 | 758.1094
66 | 16162 | 757.5938
66 | 16157 | 757.3594
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16172 | 758.0625
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16148 | 756.9375
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16145 | 756.7969
66 | 16151 | 757.0781
66 | 16158 | 757.4062
66 | 16165 | 757.7344
66 | 16170 | 757.9688
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16149 | 756.9844
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16169 | 757.9219
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16155 | 757.2656
66 | 16146 | 756.8438
66 | 16136 | 756.375
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16144 | 756.75
66 | 16135 | 756.3281
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16153 | 757.1719
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16153 | 757.1719
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16136 | 756.375
66 | 16133 | 756.2344
66 | 16146 | 756.8438
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16149 | 756.9844
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16138 | 756.4688
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16159 | 757.4531
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16170 | 757.9688
66 | 16172 | 758.0625
66 | 16160 | 757.5
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16141 | 756.6094
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16139 | 756.5156
66 | 16150 | 757.0312
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16155 | 757.2656
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16149 | 756.9844
66 | 16156 | 757.3125
66 | 16158 | 757.4062
66 | 16171 | 758.0156
66 | 16162 | 757.5938
66 | 16170 | 757.9688
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16155 | 757.2656
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16172 | 758.0625
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16138 | 756.4688
66 | 16142 | 756.6562
66 | 16150 | 757.0312
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16138 | 756.4688
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16156 | 757.3125
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16170 | 757.9688
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16153 | 757.1719
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16128 | 756
66 | 16134 | 756.2812
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16160 | 757.5
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16139 | 756.5156
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16168 | 757.875
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16145 | 756.7969
66 | 16133 | 756.2344
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16133 | 756.2344
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16140 | 756.5625
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16160 | 757.5
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16135 | 756.3281
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16134 | 756.2812
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16150 | 757.0312
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16155 | 757.2656
66 | 16150 | 757.0312
66 | 16141 | 756.6094
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16157 | 757.3594
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16138 | 756.4688
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16150 | 757.0312
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16141 | 756.6094
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16136 | 756.375
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16137 | 756.4219
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16147 | 756.8906
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16135 | 756.3281
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16149 | 756.9844
66 | 16144 | 756.75
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16139 | 756.5156
66 | 16147 | 756.8906
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16150 | 757.0312
66 | 16155 | 757.2656
66 | 16154 | 757.2188
66 | 16147 | 756.8906
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16149 | 756.9844
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16151 | 757.0781
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16141 | 756.6094
66 | 16135 | 756.3281
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16137 | 756.4219
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16151 | 757.0781
66 | 16148 | 756.9375
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16134 | 756.2812
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16150 | 757.0312
66 | 16143 | 756.7031
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16136 | 756.375
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16136 | 756.375
66 | 16141 | 756.6094
66 | 16128 | 756
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16112 | 755.25
66 | 16122 | 755.7188
66 | 16111 | 755.2031
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16111 | 755.2031
66 | 16122 | 755.7188
66 | 16116 | 755.4375
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16148 | 756.9375
66 | 16141 | 756.6094
66 | 16149 | 756.9844
66 | 16138 | 756.4688
66 | 16142 | 756.6562
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16147 | 756.8906
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16147 | 756.8906
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16132 | 756.1875
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16137 | 756.4219
66 | 16132 | 756.1875
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16138 | 756.4688
66 | 16138 | 756.4688
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16134 | 756.2812
66 | 16139 | 756.5156
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16134 | 756.2812
66 | 16139 | 756.5156
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16111 | 755.2031
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16135 | 756.3281
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16128 | 756
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16134 | 756.2812
66 | 16143 | 756.7031
66 | 16139 | 756.5156
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16132 | 756.1875
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16135 | 756.3281
66 | 16128 | 756
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16134 | 756.2812
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16133 | 756.2344
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16092 | 754.3125
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16093 | 754.3594
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16091 | 754.2656
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16093 | 754.3594
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16092 | 754.3125
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16092 | 754.3125
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16091 | 754.2656
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16087 | 754.0781
66 | 16093 | 754.3594
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16092 | 754.3125
66 | 16103 | 754.8281
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16094 | 754.4062
66 | 16102 | 754.7812
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16077 | 753.6094
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16076 | 753.5625
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16093 | 754.3594
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16076 | 753.5625
66 | 16076 | 753.5625
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16075 | 753.5156
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16076 | 753.5625
66 | 16074 | 753.4688
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16075 | 753.5156
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16074 | 753.4688
66 | 16074 | 753.4688
66 | 16075 | 753.5156
66 | 16082 | 753.8438
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16074 | 753.4688
66 | 16073 | 753.4219
66 | 16076 | 753.5625
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16104 | 754.875
66 | 16092 | 754.3125
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16072 | 753.375
66 | 16073 | 753.4219
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16075 | 753.5156
66 | 16072 | 753.375
66 | 16075 | 753.5156
66 | 16078 | 753.6562
66 | 16086 | 754.0312
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16073 | 753.4219
66 | 16072 | 753.375
66 | 16074 | 753.4688
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16091 | 754.2656
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16072 | 753.375
66 | 16072 | 753.375
66 | 16076 | 753.5625
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16075 | 753.5156
66 | 16072 | 753.375
66 | 16071 | 753.3281
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16092 | 754.3125
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16072 | 753.375
66 | 16071 | 753.3281
66 | 16073 | 753.4219
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16075 | 753.5156
66 | 16071 | 753.3281
66 | 16070 | 753.2812
66 | 16075 | 753.5156
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16071 | 753.3281
66 | 16069 | 753.2344
66 | 16072 | 753.375
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16073 | 753.4219
66 | 16070 | 753.2812
66 | 16070 | 753.2812
66 | 16074 | 753.4688
66 | 16082 | 753.8438
66 | 16092 | 754.3125
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16091 | 754.2656
66 | 16081 | 753.7969
66 | 16073 | 753.4219
66 | 16071 | 753.3281
66 | 16070 | 753.2812
66 | 16072 | 753.375
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16070 | 753.2812
66 | 16067 | 753.1406
66 | 16069 | 753.2344
66 | 16075 | 753.5156
66 | 16084 | 753.9375
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16071 | 753.3281
66 | 16067 | 753.1406
66 | 16069 | 753.2344
66 | 16074 | 753.4688
66 | 16082 | 753.8438
66 | 16092 | 754.3125
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16071 | 753.3281
66 | 16068 | 753.1875
66 | 16069 | 753.2344
66 | 16073 | 753.4219
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16071 | 753.3281
66 | 16068 | 753.1875
66 | 16067 | 753.1406
66 | 16072 | 753.375
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16071 | 753.3281
66 | 16068 | 753.1875
66 | 16068 | 753.1875
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16070 | 753.2812
66 | 16066 | 753.0938
66 | 16067 | 753.1406
66 | 16072 | 753.375
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16069 | 753.2344
66 | 16067 | 753.1406
66 | 16067 | 753.1406
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16070 | 753.2812
66 | 16067 | 753.1406
66 | 16067 | 753.1406
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16069 | 753.2344
66 | 16065 | 753.0469
66 | 16066 | 753.0938
66 | 16073 | 753.4219
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16092 | 754.3125
66 | 16082 | 753.8438
66 | 16073 | 753.4219
66 | 16068 | 753.1875
66 | 16064 | 753
66 | 16067 | 753.1406
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16071 | 753.3281
66 | 16066 | 753.0938
66 | 16064 | 753
66 | 16067 | 753.1406
66 | 16073 | 753.4219
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16070 | 753.2812
66 | 16065 | 753.0469
66 | 16064 | 753
66 | 16065 | 753.0469
66 | 16074 | 753.4688
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16065 | 753.0469
66 | 16062 | 752.9062
66 | 16063 | 752.9531
66 | 16071 | 753.3281
66 | 16077 | 753.6094
66 | 16090 | 754.2188
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16091 | 754.2656
66 | 16079 | 753.7031
66 | 16069 | 753.2344
66 | 16062 | 752.9062
66 | 16060 | 752.8125
66 | 16054 | 752.5312
66 | 16060 | 752.8125
66 | 16069 | 753.2344
66 | 16079 | 753.7031
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16089 | 754.1719
66 | 16077 | 753.6094
66 | 16061 | 752.8594
66 | 16056 | 752.625
66 | 16052 | 752.4375
66 | 16056 | 752.625
66 | 16063 | 752.9531
66 | 16072 | 753.375
66 | 16091 | 754.2656
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16097 | 754.5469
66 | 16088 | 754.125
66 | 16076 | 753.5625
66 | 16066 | 753.0938
66 | 16061 | 752.8594
66 | 16063 | 752.9531
66 | 16062 | 752.9062
66 | 16067 | 753.1406
66 | 16073 | 753.4219
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16074 | 753.4688
66 | 16065 | 753.0469
66 | 16060 | 752.8125
66 | 16060 | 752.8125
66 | 16065 | 753.0469
66 | 16074 | 753.4688
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16094 | 754.4062
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16073 | 753.4219
66 | 16065 | 753.0469
66 | 16059 | 752.7656
66 | 16060 | 752.8125
66 | 16065 | 753.0469
66 | 16072 | 753.375
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16069 | 753.2344
66 | 16063 | 752.9531
66 | 16060 | 752.8125
66 | 16062 | 752.9062
66 | 16066 | 753.0938
66 | 16075 | 753.5156
66 | 16078 | 753.6562
66 | 16089 | 754.1719
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16078 | 753.6562
66 | 16067 | 753.1406
66 | 16059 | 752.7656
66 | 16058 | 752.7188
66 | 16053 | 752.4844
66 | 16059 | 752.7656
66 | 16068 | 753.1875
66 | 16080 | 753.75
66 | 16091 | 754.2656
66 | 16106 | 754.9688
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16069 | 753.2344
66 | 16062 | 752.9062
66 | 16062 | 752.9062
66 | 16059 | 752.7656
66 | 16060 | 752.8125
66 | 16068 | 753.1875
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16064 | 753
66 | 16059 | 752.7656
66 | 16055 | 752.5781
66 | 16058 | 752.7188
66 | 16065 | 753.0469
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16074 | 753.4688
66 | 16076 | 753.5625
66 | 16069 | 753.2344
66 | 16063 | 752.9531
66 | 16063 | 752.9531
66 | 16061 | 752.8594
66 | 16067 | 753.1406
66 | 16075 | 753.5156
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16091 | 754.2656
66 | 16080 | 753.75
66 | 16071 | 753.3281
66 | 16063 | 752.9531
66 | 16062 | 752.9062
66 | 16058 | 752.7188
66 | 16057 | 752.6719
66 | 16063 | 752.9531
66 | 16065 | 753.0469
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16074 | 753.4688
66 | 16071 | 753.3281
66 | 16061 | 752.8594
66 | 16056 | 752.625
66 | 16050 | 752.3438
66 | 16055 | 752.5781
66 | 16064 | 753
66 | 16075 | 753.5156
66 | 16095 | 754.4531
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16093 | 754.3594
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16069 | 753.2344
66 | 16060 | 752.8125
66 | 16048 | 752.25
66 | 16050 | 752.3438
66 | 16063 | 752.9531
66 | 16068 | 753.1875
66 | 16076 | 753.5625
66 | 16079 | 753.7031
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16065 | 753.0469
66 | 16064 | 753
66 | 16058 | 752.7188
66 | 16052 | 752.4375
66 | 16051 | 752.3906
66 | 16058 | 752.7188
66 | 16068 | 753.1875
66 | 16084 | 753.9375
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16065 | 753.0469
66 | 16055 | 752.5781
66 | 16046 | 752.1562
66 | 16051 | 752.3906
66 | 16065 | 753.0469
66 | 16072 | 753.375
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16092 | 754.3125
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16089 | 754.1719
66 | 16077 | 753.6094
66 | 16061 | 752.8594
66 | 16054 | 752.5312
66 | 16054 | 752.5312
66 | 16057 | 752.6719
66 | 16058 | 752.7188
66 | 16059 | 752.7656
66 | 16069 | 753.2344
66 | 16081 | 753.7969
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16072 | 753.375
66 | 16062 | 752.9062
66 | 16054 | 752.5312
66 | 16046 | 752.1562
66 | 16048 | 752.25
66 | 16056 | 752.625
66 | 16072 | 753.375
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16061 | 752.8594
66 | 16050 | 752.3438
66 | 16046 | 752.1562
66 | 16048 | 752.25
66 | 16060 | 752.8125
66 | 16067 | 753.1406
66 | 16073 | 753.4219
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16091 | 754.2656
66 | 16076 | 753.5625
66 | 16063 | 752.9531
66 | 16056 | 752.625
66 | 16052 | 752.4375
66 | 16056 | 752.625
66 | 16054 | 752.5312
66 | 16059 | 752.7656
66 | 16060 | 752.8125
66 | 16070 | 753.2812
66 | 16083 | 753.8906
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16076 | 753.5625
66 | 16068 | 753.1875
66 | 16058 | 752.7188
66 | 16045 | 752.1094
66 | 16046 | 752.1562
66 | 16052 | 752.4375
66 | 16065 | 753.0469
66 | 16070 | 753.2812
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16092 | 754.3125
66 | 16075 | 753.5156
66 | 16065 | 753.0469
66 | 16061 | 752.8594
66 | 16059 | 752.7656
66 | 16053 | 752.4844
66 | 16052 | 752.4375
66 | 16048 | 752.25
66 | 16058 | 752.7188
66 | 16070 | 753.2812
66 | 16087 | 754.0781
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16062 | 752.9062
66 | 16055 | 752.5781
66 | 16052 | 752.4375
66 | 16054 | 752.5312
66 | 16051 | 752.3906
66 | 16051 | 752.3906
66 | 16058 | 752.7188
66 | 16071 | 753.3281
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16101 | 754.7344
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16079 | 753.7031
66 | 16063 | 752.9531
66 | 16052 | 752.4375
66 | 16047 | 752.2031
66 | 16046 | 752.1562
66 | 16055 | 752.5781
66 | 16060 | 752.8125
66 | 16067 | 753.1406
66 | 16074 | 753.4688
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16091 | 754.2656
66 | 16081 | 753.7969
66 | 16068 | 753.1875
66 | 16053 | 752.4844
66 | 16049 | 752.2969
66 | 16046 | 752.1562
66 | 16057 | 752.6719
66 | 16059 | 752.7656
66 | 16066 | 753.0938
66 | 16073 | 753.4219
66 | 16075 | 753.5156
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16062 | 752.9062
66 | 16052 | 752.4375
66 | 16046 | 752.1562
66 | 16045 | 752.1094
66 | 16052 | 752.4375
66 | 16058 | 752.7188
66 | 16065 | 753.0469
66 | 16074 | 753.4688
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16093 | 754.3594
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16062 | 752.9062
66 | 16053 | 752.4844
66 | 16047 | 752.2031
66 | 16043 | 752.0156
66 | 16046 | 752.1562
66 | 16058 | 752.7188
66 | 16065 | 753.0469
66 | 16073 | 753.4219
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16082 | 753.8438
66 | 16069 | 753.2344
66 | 16060 | 752.8125
66 | 16047 | 752.2031
66 | 16048 | 752.25
66 | 16049 | 752.2969
66 | 16052 | 752.4375
66 | 16057 | 752.6719
66 | 16070 | 753.2812
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16075 | 753.5156
66 | 16068 | 753.1875
66 | 16062 | 752.9062
66 | 16063 | 752.9531
66 | 16058 | 752.7188
66 | 16056 | 752.625
66 | 16058 | 752.7188
66 | 16058 | 752.7188
66 | 16061 | 752.8594
66 | 16069 | 753.2344
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16092 | 754.3125
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16071 | 753.3281
66 | 16063 | 752.9531
66 | 16058 | 752.7188
66 | 16056 | 752.625
66 | 16057 | 752.6719
66 | 16058 | 752.7188
66 | 16059 | 752.7656
66 | 16067 | 753.1406
66 | 16075 | 753.5156
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16069 | 753.2344
66 | 16063 | 752.9531
66 | 16058 | 752.7188
66 | 16053 | 752.4844
66 | 16052 | 752.4375
66 | 16051 | 752.3906
66 | 16058 | 752.7188
66 | 16066 | 753.0938
66 | 16073 | 753.4219
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16067 | 753.1406
66 | 16058 | 752.7188
66 | 16048 | 752.25
66 | 16047 | 752.2031
66 | 16048 | 752.25
66 | 16052 | 752.4375
66 | 16057 | 752.6719
66 | 16067 | 753.1406
66 | 16075 | 753.5156
66 | 16089 | 754.1719
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16067 | 753.1406
66 | 16061 | 752.8594
66 | 16054 | 752.5312
66 | 16050 | 752.3438
66 | 16055 | 752.5781
66 | 16056 | 752.625
66 | 16061 | 752.8594
66 | 16065 | 753.0469
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16075 | 753.5156
66 | 16070 | 753.2812
66 | 16063 | 752.9531
66 | 16064 | 753
66 | 16057 | 752.6719
66 | 16054 | 752.5312
66 | 16056 | 752.625
66 | 16057 | 752.6719
66 | 16062 | 752.9062
66 | 16062 | 752.9062
66 | 16069 | 753.2344
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16061 | 752.8594
66 | 16053 | 752.4844
66 | 16048 | 752.25
66 | 16046 | 752.1562
66 | 16050 | 752.3438
66 | 16059 | 752.7656
66 | 16063 | 752.9531
66 | 16071 | 753.3281
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16075 | 753.5156
66 | 16069 | 753.2344
66 | 16060 | 752.8125
66 | 16056 | 752.625
66 | 16056 | 752.625
66 | 16054 | 752.5312
66 | 16056 | 752.625
66 | 16059 | 752.7656
66 | 16066 | 753.0938
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16068 | 753.1875
66 | 16061 | 752.8594
66 | 16054 | 752.5312
66 | 16053 | 752.4844
66 | 16053 | 752.4844
66 | 16055 | 752.5781
66 | 16056 | 752.625
66 | 16061 | 752.8594
66 | 16070 | 753.2812
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16070 | 753.2812
66 | 16063 | 752.9531
66 | 16058 | 752.7188
66 | 16055 | 752.5781
66 | 16052 | 752.4375
66 | 16055 | 752.5781
66 | 16053 | 752.4844
66 | 16060 | 752.8125
66 | 16067 | 753.1406
66 | 16076 | 753.5625
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16069 | 753.2344
66 | 16063 | 752.9531
66 | 16055 | 752.5781
66 | 16054 | 752.5312
66 | 16057 | 752.6719
66 | 16055 | 752.5781
66 | 16058 | 752.7188
66 | 16062 | 752.9062
66 | 16068 | 753.1875
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16071 | 753.3281
66 | 16065 | 753.0469
66 | 16058 | 752.7188
66 | 16055 | 752.5781
66 | 16053 | 752.4844
66 | 16056 | 752.625
66 | 16060 | 752.8125
66 | 16066 | 753.0938
66 | 16072 | 753.375
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16076 | 753.5625
66 | 16069 | 753.2344
66 | 16062 | 752.9062
66 | 16055 | 752.5781
66 | 16048 | 752.25
66 | 16044 | 752.0625
66 | 16043 | 752.0156
66 | 16045 | 752.1094
66 | 16049 | 752.2969
66 | 16054 | 752.5312
66 | 16062 | 752.9062
66 | 16069 | 753.2344
66 | 16075 | 753.5156
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16072 | 753.375
66 | 16065 | 753.0469
66 | 16058 | 752.7188
66 | 16049 | 752.2969
66 | 16046 | 752.1562
66 | 16042 | 751.9688
66 | 16042 | 751.9688
66 | 16047 | 752.2031
66 | 16053 | 752.4844
66 | 16058 | 752.7188
66 | 16066 | 753.0938
66 | 16073 | 753.4219
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16073 | 753.4219
66 | 16066 | 753.0938
66 | 16058 | 752.7188
66 | 16052 | 752.4375
66 | 16046 | 752.1562
66 | 16042 | 751.9688
66 | 16043 | 752.0156
66 | 16044 | 752.0625
66 | 16049 | 752.2969
66 | 16057 | 752.6719
66 | 16064 | 753
66 | 16071 | 753.3281
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16075 | 753.5156
66 | 16071 | 753.3281
66 | 16066 | 753.0938
66 | 16066 | 753.0938
66 | 16051 | 752.3906
66 | 16048 | 752.25
66 | 16045 | 752.1094
66 | 16048 | 752.25
66 | 16052 | 752.4375
66 | 16062 | 752.9062
66 | 16070 | 753.2812
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16092 | 754.3125
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16064 | 753
66 | 16057 | 752.6719
66 | 16049 | 752.2969
66 | 16045 | 752.1094
66 | 16046 | 752.1562
66 | 16048 | 752.25
66 | 16052 | 752.4375
66 | 16059 | 752.7656
66 | 16064 | 753
66 | 16072 | 753.375
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16067 | 753.1406
66 | 16058 | 752.7188
66 | 16051 | 752.3906
66 | 16047 | 752.2031
66 | 16044 | 752.0625
66 | 16046 | 752.1562
66 | 16053 | 752.4844
66 | 16057 | 752.6719
66 | 16063 | 752.9531
66 | 16070 | 753.2812
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16074 | 753.4688
66 | 16065 | 753.0469
66 | 16055 | 752.5781
66 | 16045 | 752.1094
66 | 16042 | 751.9688
66 | 16040 | 751.875
66 | 16043 | 752.0156
66 | 16047 | 752.2031
66 | 16053 | 752.4844
66 | 16068 | 753.1875
66 | 16075 | 753.5156
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16069 | 753.2344
66 | 16062 | 752.9062
66 | 16054 | 752.5312
66 | 16052 | 752.4375
66 | 16048 | 752.25
66 | 16042 | 751.9688
66 | 16042 | 751.9688
66 | 16048 | 752.25
66 | 16055 | 752.5781
66 | 16065 | 753.0469
66 | 16075 | 753.5156
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16068 | 753.1875
66 | 16058 | 752.7188
66 | 16050 | 752.3438
66 | 16036 | 751.6875
66 | 16032 | 751.5
66 | 16032 | 751.5
66 | 16038 | 751.7812
66 | 16046 | 752.1562
66 | 16055 | 752.5781
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16072 | 753.375
66 | 16061 | 752.8594
66 | 16052 | 752.4375
66 | 16038 | 751.7812
66 | 16035 | 751.6406
66 | 16032 | 751.5
66 | 16036 | 751.6875
66 | 16040 | 751.875
66 | 16049 | 752.2969
66 | 16057 | 752.6719
66 | 16070 | 753.2812
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16070 | 753.2812
66 | 16060 | 752.8125
66 | 16052 | 752.4375
66 | 16046 | 752.1562
66 | 16043 | 752.0156
66 | 16050 | 752.3438
66 | 16051 | 752.3906
66 | 16054 | 752.5312
66 | 16060 | 752.8125
66 | 16066 | 753.0938
66 | 16072 | 753.375
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16069 | 753.2344
66 | 16063 | 752.9531
66 | 16057 | 752.6719
66 | 16052 | 752.4375
66 | 16048 | 752.25
66 | 16046 | 752.1562
66 | 16050 | 752.3438
66 | 16056 | 752.625
66 | 16061 | 752.8594
66 | 16068 | 753.1875
66 | 16074 | 753.4688
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16074 | 753.4688
66 | 16063 | 752.9531
66 | 16053 | 752.4844
66 | 16048 | 752.25
66 | 16042 | 751.9688
66 | 16040 | 751.875
66 | 16043 | 752.0156
66 | 16047 | 752.2031
66 | 16056 | 752.625
66 | 16066 | 753.0938
66 | 16073 | 753.4219
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16065 | 753.0469
66 | 16054 | 752.5312
66 | 16046 | 752.1562
66 | 16034 | 751.5938
66 | 16030 | 751.4062
66 | 16029 | 751.3594
66 | 16034 | 751.5938
66 | 16041 | 751.9219
66 | 16051 | 752.3906
66 | 16068 | 753.1875
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16061 | 752.8594
66 | 16051 | 752.3906
66 | 16042 | 751.9688
66 | 16037 | 751.7344
66 | 16032 | 751.5
66 | 16031 | 751.4531
66 | 16036 | 751.6875
66 | 16048 | 752.25
66 | 16056 | 752.625
66 | 16063 | 752.9531
66 | 16073 | 753.4219
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16069 | 753.2344
66 | 16062 | 752.9062
66 | 16054 | 752.5312
66 | 16046 | 752.1562
66 | 16041 | 751.9219
66 | 16042 | 751.9688
66 | 16039 | 751.8281
66 | 16039 | 751.8281
66 | 16041 | 751.9219
66 | 16042 | 751.9688
66 | 16050 | 752.3438
66 | 16059 | 752.7656
66 | 16068 | 753.1875
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16070 | 753.2812
66 | 16056 | 752.625
66 | 16046 | 752.1562
66 | 16039 | 751.8281
66 | 16034 | 751.5938
66 | 16029 | 751.3594
66 | 16034 | 751.5938
66 | 16043 | 752.0156
66 | 16047 | 752.2031
66 | 16054 | 752.5312
66 | 16063 | 752.9531
66 | 16065 | 753.0469
66 | 16073 | 753.4219
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16068 | 753.1875
66 | 16058 | 752.7188
66 | 16048 | 752.25
66 | 16034 | 751.5938
66 | 16029 | 751.3594
66 | 16030 | 751.4062
66 | 16034 | 751.5938
66 | 16047 | 752.2031
66 | 16055 | 752.5781
66 | 16061 | 752.8594
66 | 16071 | 753.3281
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16076 | 753.5625
66 | 16070 | 753.2812
66 | 16063 | 752.9531
66 | 16061 | 752.8594
66 | 16051 | 752.3906
66 | 16041 | 751.9219
66 | 16035 | 751.6406
66 | 16027 | 751.2656
66 | 16030 | 751.4062
66 | 16036 | 751.6875
66 | 16044 | 752.0625
66 | 16053 | 752.4844
66 | 16069 | 753.2344
66 | 16079 | 753.7031
66 | 16086 | 754.0312
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16070 | 753.2812
66 | 16062 | 752.9062
66 | 16061 | 752.8594
66 | 16050 | 752.3438
66 | 16041 | 751.9219
66 | 16030 | 751.4062
66 | 16029 | 751.3594
66 | 16032 | 751.5
66 | 16037 | 751.7344
66 | 16046 | 752.1562
66 | 16063 | 752.9531
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16068 | 753.1875
66 | 16058 | 752.7188
66 | 16050 | 752.3438
66 | 16039 | 751.8281
66 | 16035 | 751.6406
66 | 16033 | 751.5469
66 | 16038 | 751.7812
66 | 16041 | 751.9219
66 | 16045 | 752.1094
66 | 16053 | 752.4844
66 | 16053 | 752.4844
66 | 16063 | 752.9531
66 | 16072 | 753.375
66 | 16051 | 752.3906
66 | 16037 | 751.7344
66 | 16068 | 753.1875
66 | 16063 | 752.9531
66 | 16049 | 752.2969
66 | 16039 | 751.8281
66 | 16036 | 751.6875
66 | 16046 | 752.1562
66 | 16053 | 752.4844
66 | 16062 | 752.9062
66 | 16070 | 753.2812
66 | 16078 | 753.6562
66 | 16077 | 753.6094
66 | 16078 | 753.6562
66 | 16078 | 753.6562
66 | 16075 | 753.5156
66 | 16067 | 753.1406
66 | 16062 | 752.9062
66 | 16057 | 752.6719
66 | 16045 | 752.1094
66 | 16038 | 751.7812
66 | 16033 | 751.5469
66 | 16027 | 751.2656
66 | 16030 | 751.4062
66 | 16038 | 751.7812
66 | 16047 | 752.2031
66 | 16057 | 752.6719
66 | 16066 | 753.0938
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16066 | 753.0938
66 | 16058 | 752.7188
66 | 16039 | 751.8281
66 | 16032 | 751.5
66 | 16028 | 751.3125
66 | 16028 | 751.3125
66 | 16031 | 751.4531
66 | 16045 | 752.1094
66 | 16052 | 752.4375
66 | 16061 | 752.8594
66 | 16068 | 753.1875
66 | 16070 | 753.2812
66 | 16076 | 753.5625
66 | 16078 | 753.6562
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16071 | 753.3281
66 | 16059 | 752.7656
66 | 16044 | 752.0625
66 | 16040 | 751.875
66 | 16036 | 751.6875
66 | 16032 | 751.5
66 | 16037 | 751.7344
66 | 16046 | 752.1562
66 | 16049 | 752.2969
66 | 16055 | 752.5781
66 | 16056 | 752.625
66 | 16063 | 752.9531
66 | 16068 | 753.1875
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16069 | 753.2344
66 | 16054 | 752.5312
66 | 16049 | 752.2969
66 | 16043 | 752.0156
66 | 16046 | 752.1562
66 | 16040 | 751.875
66 | 16038 | 751.7812
66 | 16035 | 751.6406
66 | 16038 | 751.7812
66 | 16047 | 752.2031
66 | 16055 | 752.5781
66 | 16072 | 753.375
66 | 16076 | 753.5625
66 | 16080 | 753.75
66 | 16080 | 753.75
66 | 16077 | 753.6094
66 | 16071 | 753.3281
66 | 16063 | 752.9531
66 | 16050 | 752.3438
66 | 16044 | 752.0625
66 | 16036 | 751.6875
66 | 16033 | 751.5469
66 | 16032 | 751.5
66 | 16034 | 751.5938
66 | 16041 | 751.9219
66 | 16054 | 752.5312
66 | 16061 | 752.8594
66 | 16066 | 753.0938
66 | 16071 | 753.3281
66 | 16068 | 753.1875
66 | 16070 | 753.2812
66 | 16072 | 753.375
66 | 16075 | 753.5156
66 | 16069 | 753.2344
66 | 16058 | 752.7188
66 | 16050 | 752.3438
66 | 16037 | 751.7344
66 | 16033 | 751.5469
66 | 16034 | 751.5938
66 | 16043 | 752.0156
66 | 16046 | 752.1562
66 | 16051 | 752.3906
66 | 16055 | 752.5781
66 | 16056 | 752.625
66 | 16064 | 753
66 | 16068 | 753.1875
66 | 16072 | 753.375
66 | 16080 | 753.75
66 | 16075 | 753.5156
66 | 16068 | 753.1875
66 | 16060 | 752.8125
66 | 16050 | 752.3438
66 | 16036 | 751.6875
66 | 16033 | 751.5469
66 | 16032 | 751.5
66 | 16035 | 751.6406
66 | 16046 | 752.1562
66 | 16050 | 752.3438
66 | 16054 | 752.5312
66 | 16059 | 752.7656
66 | 16060 | 752.8125
66 | 16068 | 753.1875
66 | 16071 | 753.3281
66 | 16080 | 753.75
66 | 16075 | 753.5156
66 | 16068 | 753.1875
66 | 16055 | 752.5781
66 | 16047 | 752.2031
66 | 16041 | 751.9219
66 | 16035 | 751.6406
66 | 16035 | 751.6406
66 | 16042 | 751.9688
66 | 16041 | 751.9219
66 | 16046 | 752.1562
66 | 16045 | 752.1094
66 | 16052 | 752.4375
66 | 16061 | 752.8594
66 | 16068 | 753.1875
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16074 | 753.4688
66 | 16062 | 752.9062
66 | 16055 | 752.5781
66 | 16049 | 752.2969
66 | 16043 | 752.0156
66 | 16038 | 751.7812
66 | 16040 | 751.875
66 | 16041 | 751.9219
66 | 16041 | 751.9219
66 | 16043 | 752.0156
66 | 16044 | 752.0625
66 | 16049 | 752.2969
66 | 16057 | 752.6719
66 | 16065 | 753.0469
66 | 16071 | 753.3281
66 | 16079 | 753.7031
66 | 16077 | 753.6094
66 | 16071 | 753.3281
66 | 16065 | 753.0469
66 | 16048 | 752.25
66 | 16042 | 751.9688
66 | 16038 | 751.7812
66 | 16034 | 751.5938
66 | 16040 | 751.875
66 | 16040 | 751.875
66 | 16042 | 751.9688
66 | 16041 | 751.9219
66 | 16048 | 752.25
66 | 16056 | 752.625
66 | 16064 | 753
66 | 16077 | 753.6094
66 | 16078 | 753.6562
66 | 16074 | 753.4688
66 | 16064 | 753
66 | 16058 | 752.7188
66 | 16052 | 752.4375
66 | 16045 | 752.1094
66 | 16048 | 752.25
66 | 16041 | 751.9219
66 | 16038 | 751.7812
66 | 16031 | 751.4531
66 | 16034 | 751.5938
66 | 16042 | 751.9688
66 | 16052 | 752.4375
66 | 16064 | 753
66 | 16068 | 753.1875
66 | 16071 | 753.3281
66 | 16069 | 753.2344
66 | 16069 | 753.2344
66 | 16063 | 752.9531
66 | 16060 | 752.8125
66 | 16056 | 752.625
66 | 16054 | 752.5312
66 | 16047 | 752.2031
66 | 16040 | 751.875
66 | 16037 | 751.7344
66 | 16029 | 751.3594
66 | 16031 | 751.4531
66 | 16039 | 751.8281
66 | 16055 | 752.5781
66 | 16061 | 752.8594
66 | 16066 | 753.0938
66 | 16066 | 753.0938
66 | 16066 | 753.0938
66 | 16068 | 753.1875
66 | 16068 | 753.1875
66 | 16070 | 753.2812
66 | 16062 | 752.9062
66 | 16053 | 752.4844
66 | 16042 | 751.9688
66 | 16034 | 751.5938
66 | 16028 | 751.3125
66 | 16032 | 751.5
66 | 16039 | 751.8281
66 | 16043 | 752.0156
66 | 16049 | 752.2969
66 | 16048 | 752.25
66 | 16055 | 752.5781
66 | 16063 | 752.9531
66 | 16070 | 753.2812
66 | 16079 | 753.7031
66 | 16075 | 753.5156
66 | 16070 | 753.2812
66 | 16056 | 752.625
66 | 16048 | 752.25
66 | 16042 | 751.9688
66 | 16040 | 751.875
66 | 16042 | 751.9688
66 | 16037 | 751.7344
66 | 16037 | 751.7344
66 | 16036 | 751.6875
66 | 16040 | 751.875
66 | 16047 | 752.2031
66 | 16055 | 752.5781
66 | 16063 | 752.9531
66 | 16070 | 753.2812
66 | 16078 | 753.6562
66 | 16076 | 753.5625
66 | 16070 | 753.2812
66 | 16066 | 753.0938
66 | 16051 | 752.3906
66 | 16043 | 752.0156
66 | 16039 | 751.8281
66 | 16034 | 751.5938
66 | 16039 | 751.8281
66 | 16039 | 751.8281
66 | 16040 | 751.875
66 | 16046 | 752.1562
66 | 16044 | 752.0625
66 | 16052 | 752.4375
66 | 16060 | 752.8125
66 | 16065 | 753.0469
66 | 16070 | 753.2812
66 | 16077 | 753.6094
66 | 16075 | 753.5156
66 | 16070 | 753.2812
66 | 16061 | 752.8594
66 | 16052 | 752.4375
66 | 16040 | 751.875
66 | 16034 | 751.5938
66 | 16029 | 751.3594
66 | 16030 | 751.4062
66 | 16032 | 751.5
66 | 16046 | 752.1562
66 | 16049 | 752.2969
66 | 16056 | 752.625
66 | 16062 | 752.9062
66 | 16062 | 752.9062
66 | 16065 | 753.0469
66 | 16068 | 753.1875
66 | 16070 | 753.2812
66 | 16072 | 753.375
66 | 16063 | 752.9531
66 | 16054 | 752.5312
66 | 16041 | 751.9219
66 | 16034 | 751.5938
66 | 16030 | 751.4062
66 | 16030 | 751.4062
66 | 16038 | 751.7812
66 | 16043 | 752.0156
66 | 16047 | 752.2031
66 | 16053 | 752.4844
66 | 16059 | 752.7656
66 | 16058 | 752.7188
66 | 16064 | 753
66 | 16067 | 753.1406
66 | 16067 | 753.1406
66 | 16074 | 753.4688
66 | 16065 | 753.0469
66 | 16056 | 752.625
66 | 16042 | 751.9688
66 | 16036 | 751.6875
66 | 16031 | 751.4531
66 | 16029 | 751.3594
66 | 16039 | 751.8281
66 | 16040 | 751.875
66 | 16041 | 751.9219
66 | 16047 | 752.2031
66 | 16047 | 752.2031
66 | 16056 | 752.625
66 | 16063 | 752.9531
66 | 16074 | 753.4688
66 | 16075 | 753.5156
66 | 16071 | 753.3281
66 | 16067 | 753.1406
66 | 16054 | 752.5312
66 | 16047 | 752.2031
66 | 16041 | 751.9219
66 | 16042 | 751.9688
66 | 16038 | 751.7812
66 | 16033 | 751.5469
66 | 16028 | 751.3125
66 | 16032 | 751.5
66 | 16038 | 751.7812
66 | 16046 | 752.1562
66 | 16062 | 752.9062
66 | 16066 | 753.0938
66 | 16069 | 753.2344
66 | 16065 | 753.0469
66 | 16065 | 753.0469
66 | 16065 | 753.0469
66 | 16067 | 753.1406
66 | 16059 | 752.7656
66 | 16050 | 752.3438
66 | 16035 | 751.6406
66 | 16030 | 751.4062
66 | 16029 | 751.3594
66 | 16037 | 751.7344
66 | 16039 | 751.8281
66 | 16042 | 751.9688
66 | 16046 | 752.1562
66 | 16046 | 752.1562
66 | 16052 | 752.4375
66 | 16063 | 752.9531
66 | 16073 | 753.4219
66 | 16072 | 753.375
66 | 16065 | 753.0469
66 | 16062 | 752.9062
66 | 16057 | 752.6719
66 | 16059 | 752.7656
66 | 16049 | 752.2969
66 | 16041 | 751.9219
66 | 16029 | 751.3594
66 | 16028 | 751.3125
66 | 16028 | 751.3125
66 | 16039 | 751.8281
66 | 16043 | 752.0156
66 | 16046 | 752.1562
66 | 16046 | 752.1562
66 | 16054 | 752.5312
66 | 16060 | 752.8125
66 | 16070 | 753.2812
66 | 16074 | 753.4688
66 | 16072 | 753.375
66 | 16068 | 753.1875
66 | 16054 | 752.5312
66 | 16048 | 752.25
66 | 16043 | 752.0156
66 | 16044 | 752.0625
66 | 16037 | 751.7344
66 | 16034 | 751.5938
66 | 16025 | 751.1719
66 | 16030 | 751.4062
66 | 16038 | 751.7812
66 | 16049 | 752.2969
66 | 16054 | 752.5312
66 | 16061 | 752.8594
66 | 16059 | 752.7656
66 | 16064 | 753
66 | 16066 | 753.0938
66 | 16072 | 753.375
66 | 16069 | 753.2344
66 | 16061 | 752.8594
66 | 16048 | 752.25
66 | 16040 | 751.875
66 | 16037 | 751.7344
66 | 16040 | 751.875
66 | 16035 | 751.6406
66 | 16033 | 751.5469
66 | 16028 | 751.3125
66 | 16034 | 751.5938
66 | 16042 | 751.9688
66 | 16057 | 752.6719
66 | 16062 | 752.9062
66 | 16065 | 753.0469
66 | 16061 | 752.8594
66 | 16065 | 753.0469
66 | 16066 | 753.0938
66 | 16066 | 753.0938
66 | 16065 | 753.0469
66 | 16056 | 752.625
66 | 16049 | 752.2969
66 | 16034 | 751.5938
66 | 16029 | 751.3594
66 | 16029 | 751.3594
66 | 16036 | 751.6875
66 | 16034 | 751.5938
66 | 16031 | 751.4531
66 | 16037 | 751.7344
66 | 16047 | 752.2031
66 | 16063 | 752.9531
66 | 16065 | 753.0469
66 | 16070 | 753.2812
66 | 16064 | 753
66 | 16064 | 753
66 | 16062 | 752.9062
66 | 16066 | 753.0938
66 | 16058 | 752.7188
66 | 16050 | 752.3438
66 | 16036 | 751.6875
66 | 16032 | 751.5
66 | 16032 | 751.5
66 | 16035 | 751.6406
66 | 16035 | 751.6406
66 | 16031 | 751.4531
66 | 16035 | 751.6406
66 | 16045 | 752.1094
66 | 16060 | 752.8125
66 | 16065 | 753.0469
66 | 16068 | 753.1875
66 | 16064 | 753
66 | 16062 | 752.9062
66 | 16062 | 752.9062
66 | 16067 | 753.1406
66 | 16058 | 752.7188
66 | 16045 | 752.1094
66 | 16038 | 751.7812
66 | 16033 | 751.5469
66 | 16036 | 751.6875
66 | 16034 | 751.5938
66 | 16032 | 751.5
66 | 16028 | 751.3125
66 | 16036 | 751.6875
66 | 16050 | 752.3438
66 | 16052 | 752.4375
66 | 16055 | 752.5781
66 | 16058 | 752.7188
66 | 16065 | 753.0469
66 | 16073 | 753.4219
66 | 16069 | 753.2344
66 | 16059 | 752.7656
66 | 16052 | 752.4375
66 | 16048 | 752.25
66 | 16048 | 752.25
66 | 16038 | 751.7812
66 | 16035 | 751.6406
66 | 16024 | 751.125
66 | 16025 | 751.1719
66 | 16037 | 751.7344
66 | 16040 | 751.875
66 | 16044 | 752.0625
66 | 16043 | 752.0156
66 | 16052 | 752.4375
66 | 16066 | 753.0938
66 | 16069 | 753.2344
66 | 16070 | 753.2812
66 | 16062 | 752.9062
66 | 16059 | 752.7656
66 | 16057 | 752.6719
66 | 16058 | 752.7188
66 | 16046 | 752.1562
66 | 16034 | 751.5938
66 | 16027 | 751.2656
66 | 16027 | 751.2656
66 | 16032 | 751.5
66 | 16032 | 751.5
66 | 16028 | 751.3125
66 | 16034 | 751.5938
66 | 16049 | 752.2969
66 | 16057 | 752.6719
66 | 16061 | 752.8594
66 | 16057 | 752.6719
66 | 16063 | 752.9531
66 | 16068 | 753.1875
66 | 16068 | 753.1875
66 | 16060 | 752.8125
66 | 16049 | 752.2969
66 | 16043 | 752.0156
66 | 16039 | 751.8281
66 | 16040 | 751.875
66 | 16030 | 751.4062
66 | 16021 | 750.9844
66 | 16024 | 751.125
66 | 16028 | 751.3125
66 | 16042 | 751.9688
66 | 16044 | 752.0625
66 | 16046 | 752.1562
66 | 16050 | 752.3438
66 | 16059 | 752.7656
66 | 16072 | 753.375
66 | 16071 | 753.3281
66 | 16061 | 752.8594
66 | 16056 | 752.625
66 | 16052 | 752.4375
66 | 16056 | 752.625
66 | 16045 | 752.1094
66 | 16033 | 751.5469
66 | 16026 | 751.2188
66 | 16025 | 751.1719
66 | 16033 | 751.5469
66 | 16034 | 751.5938
66 | 16034 | 751.5938
66 | 16034 | 751.5938
66 | 16043 | 752.0156
66 | 16058 | 752.7188
66 | 16062 | 752.9062
66 | 16066 | 753.0938
66 | 16060 | 752.8125
66 | 16059 | 752.7656
66 | 16061 | 752.8594
66 | 16065 | 753.0469
66 | 16055 | 752.5781
66 | 16049 | 752.2969
66 | 16033 | 751.5469
66 | 16031 | 751.4531
66 | 16026 | 751.2188
66 | 16032 | 751.5
66 | 16031 | 751.4531
66 | 16031 | 751.4531
66 | 16030 | 751.4062
66 | 16036 | 751.6875
66 | 16049 | 752.2969
66 | 16062 | 752.9062
66 | 16063 | 752.9531
66 | 16060 | 752.8125
66 | 16063 | 752.9531
66 | 16071 | 753.3281
66 | 16066 | 753.0938
66 | 16058 | 752.7188
66 | 16046 | 752.1562
66 | 16040 | 751.875
66 | 16036 | 751.6875
66 | 16038 | 751.7812
66 | 16032 | 751.5
66 | 16030 | 751.4062
66 | 16025 | 751.1719
66 | 16029 | 751.3594
66 | 16041 | 751.9219
66 | 16050 | 752.3438
66 | 16052 | 752.4375
66 | 16051 | 752.3906
66 | 16057 | 752.6719
66 | 16062 | 752.9062
66 | 16072 | 753.375
66 | 16066 | 753.0938
66 | 16061 | 752.8594
66 | 16050 | 752.3438
66 | 16044 | 752.0625
66 | 16046 | 752.1562
66 | 16039 | 751.8281
66 | 16032 | 751.5
66 | 16022 | 751.0312
66 | 16024 | 751.125
66 | 16028 | 751.3125
66 | 16040 | 751.875
66 | 16044 | 752.0625
66 | 16048 | 752.25
66 | 16049 | 752.2969
66 | 16056 | 752.625
66 | 16059 | 752.7656
66 | 16065 | 753.0469
66 | 16072 | 753.375
66 | 16066 | 753.0938
66 | 16058 | 752.7188
66 | 16045 | 752.1094
66 | 16038 | 751.7812
66 | 16032 | 751.5
66 | 16029 | 751.3594
66 | 16035 | 751.6406
66 | 16031 | 751.4531
66 | 16031 | 751.4531
66 | 16028 | 751.3125
66 | 16033 | 751.5469
66 | 16042 | 751.9688
66 | 16051 | 752.3906
66 | 16065 | 753.0469
66 | 16066 | 753.0938
66 | 16067 | 753.1406
66 | 16060 | 752.8125
66 | 16059 | 752.7656
66 | 16058 | 752.7188
66 | 16057 | 752.6719
66 | 16046 | 752.1562
66 | 16039 | 751.8281
66 | 16027 | 751.2656
66 | 16023 | 751.0781
66 | 16032 | 751.5
66 | 16031 | 751.4531
66 | 16035 | 751.6406
66 | 16032 | 751.5
66 | 16040 | 751.875
66 | 16057 | 752.6719
66 | 16060 | 752.8125
66 | 16063 | 752.9531
66 | 16061 | 752.8594
66 | 16061 | 752.8594
66 | 16068 | 753.1875
66 | 16063 | 752.9531
66 | 16053 | 752.4844
66 | 16040 | 751.875
66 | 16032 | 751.5
66 | 16029 | 751.3594
66 | 16035 | 751.6406
66 | 16031 | 751.4531
66 | 16030 | 751.4062
66 | 16031 | 751.4531
66 | 16032 | 751.5
66 | 16040 | 751.875
66 | 16049 | 752.2969
66 | 16064 | 753
66 | 16065 | 753.0469
66 | 16068 | 753.1875
66 | 16059 | 752.7656
66 | 16059 | 752.7656
66 | 16054 | 752.5312
66 | 16057 | 752.6719
66 | 16046 | 752.1562
66 | 16039 | 751.8281
66 | 16026 | 751.2188
66 | 16022 | 751.0312
66 | 16024 | 751.125
66 | 16033 | 751.5469
66 | 16035 | 751.6406
66 | 16038 | 751.7812
66 | 16038 | 751.7812
66 | 16047 | 752.2031
66 | 16061 | 752.8594
66 | 16066 | 753.0938
66 | 16066 | 753.0938
66 | 16064 | 753
66 | 16061 | 752.8594
66 | 16053 | 752.4844
66 | 16054 | 752.5312
66 | 16043 | 752.0156
66 | 16031 | 751.4531
66 | 16018 | 750.8438
66 | 16017 | 750.7969
66 | 16023 | 751.0781
66 | 16025 | 751.1719
66 | 16026 | 751.2188
66 | 16033 | 751.5469
66 | 16045 | 752.1094
66 | 16062 | 752.9062
66 | 16066 | 753.0938
66 | 16064 | 753
66 | 16064 | 753
66 | 16065 | 753.0469
66 | 16067 | 753.1406
66 | 16055 | 752.5781
66 | 16040 | 751.875
66 | 16030 | 751.4062
66 | 16023 | 751.0781
66 | 16026 | 751.2188
66 | 16024 | 751.125
66 | 16023 | 751.0781
66 | 16020 | 750.9375
66 | 16029 | 751.3594
66 | 16047 | 752.2031
66 | 16052 | 752.4375
66 | 16056 | 752.625
66 | 16061 | 752.8594
66 | 16065 | 753.0469
66 | 16075 | 753.5156
66 | 16068 | 753.1875
66 | 16055 | 752.5781
66 | 16045 | 752.1094
66 | 16038 | 751.7812
66 | 16037 | 751.7344
66 | 16025 | 751.1719
66 | 16014 | 750.6562
66 | 16016 | 750.75
66 | 16027 | 751.2656
66 | 16033 | 751.5469
66 | 16050 | 752.3438
66 | 16049 | 752.2969
66 | 16053 | 752.4844
66 | 16055 | 752.5781
66 | 16161 | 757.5469
66 | 16140 | 756.5625
66 | 16070 | 753.2812
66 | 16036 | 751.6875
66 | 16032 | 751.5
66 | 16025 | 751.1719
66 | 16019 | 750.8906
66 | 16083 | 753.8906
66 | 16123 | 755.7656
66 | 16104 | 754.875
66 | 16072 | 753.375
66 | 16071 | 753.3281
66 | 16084 | 753.9375
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16139 | 756.5156
66 | 16128 | 756
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16155 | 757.2656
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16174 | 758.1562
66 | 16178 | 758.3438
66 | 16167 | 757.8281
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16181 | 758.4844
66 | 16190 | 758.9062
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16196 | 759.1875
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16202 | 759.4688
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16218 | 760.2188
66 | 16212 | 759.9375
66 | 16222 | 760.4062
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16225 | 760.5469
66 | 16216 | 760.125
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16227 | 760.6406
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16226 | 760.5938
66 | 16233 | 760.9219
66 | 16235 | 761.0156
66 | 16231 | 760.8281
66 | 16227 | 760.6406
66 | 16224 | 760.5
66 | 16230 | 760.7812
66 | 16232 | 760.875
66 | 16236 | 761.0625
66 | 16238 | 761.1562
66 | 16240 | 761.25
66 | 16242 | 761.3438
66 | 16240 | 761.25
66 | 16240 | 761.25
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16248 | 761.625
66 | 16245 | 761.4844
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16249 | 761.6719
66 | 16242 | 761.3438
66 | 16248 | 761.625
66 | 16254 | 761.9062
66 | 16245 | 761.4844
66 | 16257 | 762.0469
66 | 16252 | 761.8125
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16249 | 761.6719
66 | 16259 | 762.1406
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16251 | 761.7656
66 | 16263 | 762.3281
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16263 | 762.3281
66 | 16269 | 762.6094
66 | 16266 | 762.4688
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16266 | 762.4688
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16276 | 762.9375
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16284 | 763.3125
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16278 | 763.0312
66 | 16288 | 763.5
66 | 16288 | 763.5
66 | 16278 | 763.0312
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16278 | 763.0312
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16288 | 763.5
66 | 16288 | 763.5
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16293 | 763.7344
66 | 16290 | 763.5938
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16295 | 763.8281
66 | 16291 | 763.6406
66 | 16298 | 763.9688
66 | 16288 | 763.5
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16288 | 763.5
66 | 16300 | 764.0625
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16295 | 763.8281
66 | 16307 | 764.3906
66 | 16295 | 763.8281
66 | 16306 | 764.3438
66 | 16300 | 764.0625
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16305 | 764.2969
66 | 16311 | 764.5781
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16308 | 764.4375
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16306 | 764.3438
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16310 | 764.5312
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16309 | 764.4844
66 | 16309 | 764.4844
66 | 16319 | 764.9531
66 | 16309 | 764.4844
66 | 16311 | 764.5781
66 | 16321 | 765.0469
66 | 16313 | 764.6719
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16315 | 764.7656
66 | 16324 | 765.1875
66 | 16316 | 764.8125
66 | 16315 | 764.7656
66 | 16324 | 765.1875
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16317 | 764.8594
66 | 16326 | 765.2812
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16328 | 765.375
66 | 16321 | 765.0469
66 | 16321 | 765.0469
66 | 16330 | 765.4688
66 | 16319 | 764.9531
66 | 16323 | 765.1406
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16322 | 765.0938
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16325 | 765.2344
66 | 16325 | 765.2344
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16338 | 765.8438
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16324 | 765.1875
66 | 16336 | 765.75
66 | 16329 | 765.4219
66 | 16329 | 765.4219
66 | 16338 | 765.8438
66 | 16326 | 765.2812
66 | 16337 | 765.7969
66 | 16331 | 765.5156
66 | 16332 | 765.5625
66 | 16342 | 766.0312
66 | 16330 | 765.4688
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16330 | 765.4688
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16333 | 765.6094
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16346 | 766.2188
66 | 16334 | 765.6562
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16346 | 766.2188
66 | 16334 | 765.6562
66 | 16345 | 766.1719
66 | 16337 | 765.7969
66 | 16346 | 766.2188
66 | 16337 | 765.7969
66 | 16347 | 766.2656
66 | 16340 | 765.9375
66 | 16346 | 766.2188
66 | 16341 | 765.9844
66 | 16347 | 766.2656
66 | 16342 | 766.0312
66 | 16349 | 766.3594
66 | 16345 | 766.1719
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16349 | 766.3594
66 | 16344 | 766.125
66 | 16351 | 766.4531
66 | 16346 | 766.2188
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16343 | 766.0781
66 | 16349 | 766.3594
66 | 16342 | 766.0312
66 | 16353 | 766.5469
66 | 16343 | 766.0781
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16345 | 766.1719
66 | 16356 | 766.6875
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16354 | 766.5938
66 | 16344 | 766.125
66 | 16352 | 766.5
66 | 16354 | 766.5938
66 | 16344 | 766.125
66 | 16353 | 766.5469
66 | 16356 | 766.6875
66 | 16347 | 766.2656
66 | 16354 | 766.5938
66 | 16358 | 766.7812
66 | 16349 | 766.3594
66 | 16353 | 766.5469
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16351 | 766.4531
66 | 16353 | 766.5469
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16362 | 766.9688
66 | 16358 | 766.7812
66 | 16352 | 766.5
66 | 16355 | 766.6406
66 | 16363 | 767.0156
66 | 16363 | 767.0156
66 | 16358 | 766.7812
66 | 16354 | 766.5938
66 | 16359 | 766.8281
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16355 | 766.6406
66 | 16355 | 766.6406
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16357 | 766.7344
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16363 | 767.0156
66 | 16359 | 766.8281
66 | 16358 | 766.7812
66 | 16362 | 766.9688
66 | 16371 | 767.3906
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16368 | 767.25
66 | 16365 | 767.1094
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16368 | 767.25
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16371 | 767.3906
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16370 | 767.3438
66 | 16373 | 767.4844
66 | 16376 | 767.625
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16377 | 767.6719
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16375 | 767.5781
66 | 16379 | 767.7656
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16369 | 767.2969
66 | 16373 | 767.4844
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16376 | 767.625
66 | 16374 | 767.5312
66 | 16382 | 767.9062
66 | 16386 | 768.0938
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16386 | 768.0938
66 | 16386 | 768.0938
66 | 16376 | 767.625
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16387 | 768.1406
66 | 16382 | 767.9062
66 | 16378 | 767.7188
66 | 16389 | 768.2344
66 | 16386 | 768.0938
66 | 16382 | 767.9062
66 | 16390 | 768.2812
66 | 16382 | 767.9062
66 | 16386 | 768.0938
66 | 16394 | 768.4688
66 | 16382 | 767.9062
66 | 16393 | 768.4219
66 | 16387 | 768.1406
66 | 16389 | 768.2344
66 | 16392 | 768.375
66 | 16387 | 768.1406
66 | 16394 | 768.4688
66 | 16385 | 768.0469
66 | 16394 | 768.4688
66 | 16383 | 767.9531
66 | 16395 | 768.5156
66 | 16385 | 768.0469
66 | 16394 | 768.4688
66 | 16387 | 768.1406
66 | 16397 | 768.6094
66 | 16389 | 768.2344
66 | 16397 | 768.6094
66 | 16394 | 768.4688
66 | 16395 | 768.5156
66 | 16398 | 768.6562
66 | 16389 | 768.2344
66 | 16399 | 768.7031
66 | 16387 | 768.1406
66 | 16399 | 768.7031
66 | 16392 | 768.375
66 | 16392 | 768.375
66 | 16400 | 768.75
66 | 16389 | 768.2344
66 | 16399 | 768.7031
66 | 16400 | 768.75
66 | 16389 | 768.2344
66 | 16402 | 768.8438
66 | 16401 | 768.7969
66 | 16393 | 768.4219
66 | 16404 | 768.9375
66 | 16403 | 768.8906
66 | 16394 | 768.4688
66 | 16400 | 768.75
66 | 16406 | 769.0312
66 | 16398 | 768.6562
66 | 16397 | 768.6094
66 | 16407 | 769.0781
66 | 16405 | 768.9844
66 | 16396 | 768.5625
66 | 16396 | 768.5625
66 | 16407 | 769.0781
66 | 16406 | 769.0312
66 | 16400 | 768.75
66 | 16398 | 768.6562
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16410 | 769.2188
66 | 16406 | 769.0312
66 | 16402 | 768.8438
66 | 16398 | 768.6562
66 | 16404 | 768.9375
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16411 | 769.2656
66 | 16407 | 769.0781
66 | 16403 | 768.8906
66 | 16401 | 768.7969
66 | 16401 | 768.7969
66 | 16402 | 768.8438
66 | 16403 | 768.8906
66 | 16405 | 768.9844
66 | 16412 | 769.3125
66 | 16414 | 769.4062
66 | 16415 | 769.4531
66 | 16412 | 769.3125
66 | 16414 | 769.4062
66 | 16412 | 769.3125
66 | 16414 | 769.4062
66 | 16414 | 769.4062
66 | 16412 | 769.3125
66 | 16409 | 769.1719
66 | 16413 | 769.3594
66 | 16413 | 769.3594
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16409 | 769.1719
66 | 16409 | 769.1719
66 | 16414 | 769.4062
66 | 16419 | 769.6406
66 | 16421 | 769.7344
66 | 16421 | 769.7344
66 | 16416 | 769.5
66 | 16411 | 769.2656
66 | 16413 | 769.3594
66 | 16418 | 769.5938
66 | 16421 | 769.7344
66 | 16417 | 769.5469
66 | 16410 | 769.2188
66 | 16414 | 769.4062
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16416 | 769.5
66 | 16414 | 769.4062
66 | 16424 | 769.875
66 | 16426 | 769.9688
66 | 16415 | 769.4531
66 | 16422 | 769.7812
66 | 16427 | 770.0156
66 | 16415 | 769.4531
66 | 16425 | 769.9219
66 | 16425 | 769.9219
66 | 16416 | 769.5
66 | 16428 | 770.0625
66 | 16422 | 769.7812
66 | 16420 | 769.6875
66 | 16427 | 770.0156
66 | 16415 | 769.4531
66 | 16428 | 770.0625
66 | 16420 | 769.6875
66 | 16424 | 769.875
66 | 16427 | 770.0156
66 | 16422 | 769.7812
66 | 16431 | 770.2031
66 | 16421 | 769.7344
66 | 16431 | 770.2031
66 | 16422 | 769.7812
66 | 16433 | 770.2969
66 | 16422 | 769.7812
66 | 16431 | 770.2031
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16425 | 769.9219
66 | 16429 | 770.1094
66 | 16430 | 770.1562
66 | 16424 | 769.875
66 | 16433 | 770.2969
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16434 | 770.3438
66 | 16424 | 769.875
66 | 16435 | 770.3906
66 | 16436 | 770.4375
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16437 | 770.4844
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16435 | 770.3906
66 | 16428 | 770.0625
66 | 16430 | 770.1562
66 | 16437 | 770.4844
66 | 16435 | 770.3906
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16441 | 770.6719
66 | 16441 | 770.6719
66 | 16437 | 770.4844
66 | 16436 | 770.4375
66 | 16432 | 770.25
66 | 16439 | 770.5781
66 | 16443 | 770.7656
66 | 16442 | 770.7188
66 | 16443 | 770.7656
66 | 16439 | 770.5781
66 | 16437 | 770.4844
66 | 16436 | 770.4375
66 | 16435 | 770.3906
66 | 16439 | 770.5781
66 | 16438 | 770.5312
66 | 16441 | 770.6719
66 | 16440 | 770.625
66 | 16442 | 770.7188
66 | 16443 | 770.7656
66 | 16446 | 770.9062
66 | 16445 | 770.8594
66 | 16445 | 770.8594
66 | 16447 | 770.9531
66 | 16448 | 771
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16446 | 770.9062
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16449 | 771.0469
66 | 16443 | 770.7656
66 | 16443 | 770.7656
66 | 16440 | 770.625
66 | 16444 | 770.8125
66 | 16446 | 770.9062
66 | 16451 | 771.1406
66 | 16452 | 771.1875
66 | 16442 | 770.7188
66 | 16444 | 770.8125
66 | 16451 | 771.1406
66 | 16451 | 771.1406
66 | 16443 | 770.7656
66 | 16444 | 770.8125
66 | 16451 | 771.1406
66 | 16452 | 771.1875
66 | 16444 | 770.8125
66 | 16447 | 770.9531
66 | 16454 | 771.2812
66 | 16444 | 770.8125
66 | 16449 | 771.0469
66 | 16455 | 771.3281
66 | 16445 | 770.8594
66 | 16453 | 771.2344
66 | 16456 | 771.375
66 | 16446 | 770.9062
66 | 16455 | 771.3281
66 | 16454 | 771.2812
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16447 | 770.9531
66 | 16457 | 771.4219
66 | 16453 | 771.2344
66 | 16455 | 771.3281
66 | 16456 | 771.375
66 | 16450 | 771.0938
66 | 16458 | 771.4688
66 | 16450 | 771.0938
66 | 16461 | 771.6094
66 | 16453 | 771.2344
66 | 16459 | 771.5156
66 | 16454 | 771.2812
66 | 16456 | 771.375
66 | 16459 | 771.5156
66 | 16452 | 771.1875
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16454 | 771.2812
66 | 16460 | 771.5625
66 | 16452 | 771.1875
66 | 16457 | 771.4219
66 | 16462 | 771.6562
66 | 16455 | 771.3281
66 | 16458 | 771.4688
66 | 16465 | 771.7969
66 | 16459 | 771.5156
66 | 16456 | 771.375
66 | 16465 | 771.7969
66 | 16458 | 771.4688
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16461 | 771.6094
66 | 16453 | 771.2344
66 | 16463 | 771.7031
66 | 16465 | 771.7969
66 | 16461 | 771.6094
66 | 16455 | 771.3281
66 | 16465 | 771.7969
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16456 | 771.375
66 | 16463 | 771.7031
66 | 16469 | 771.9844
66 | 16469 | 771.9844
66 | 16464 | 771.75
66 | 16459 | 771.5156
66 | 16463 | 771.7031
66 | 16468 | 771.9375
66 | 16468 | 771.9375
66 | 16469 | 771.9844
66 | 16467 | 771.8906
66 | 16464 | 771.75
66 | 16461 | 771.6094
66 | 16460 | 771.5625
66 | 16461 | 771.6094
66 | 16465 | 771.7969
66 | 16468 | 771.9375
66 | 16468 | 771.9375
66 | 16468 | 771.9375
66 | 16468 | 771.9375
66 | 16469 | 771.9844
66 | 16469 | 771.9844
66 | 16467 | 771.8906
66 | 16464 | 771.75
66 | 16466 | 771.8438
66 | 16464 | 771.75
66 | 16465 | 771.7969
66 | 16462 | 771.6562
66 | 16463 | 771.7031
66 | 16465 | 771.7969
66 | 16465 | 771.7969
66 | 16466 | 771.8438
66 | 16467 | 771.8906
66 | 16472 | 772.125
66 | 16472 | 772.125
66 | 16476 | 772.3125
66 | 16479 | 772.4531
66 | 16479 | 772.4531
66 | 16477 | 772.3594
66 | 16478 | 772.4062
66 | 16473 | 772.1719
66 | 16465 | 771.7969
66 | 16469 | 771.9844
66 | 16474 | 772.2188
66 | 16479 | 772.4531
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16474 | 772.2188
66 | 16482 | 772.5938
66 | 16483 | 772.6406
66 | 16473 | 772.1719
66 | 16478 | 772.4062
66 | 16483 | 772.6406
66 | 16482 | 772.5938
66 | 16475 | 772.2656
66 | 16482 | 772.5938
66 | 16481 | 772.5469
66 | 16473 | 772.1719
66 | 16482 | 772.5938
66 | 16482 | 772.5938
66 | 16476 | 772.3125
66 | 16484 | 772.6875
66 | 16482 | 772.5938
66 | 16477 | 772.3594
66 | 16485 | 772.7344
66 | 16477 | 772.3594
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16481 | 772.5469
66 | 16486 | 772.7812
66 | 16483 | 772.6406
66 | 16486 | 772.7812
66 | 16483 | 772.6406
66 | 16486 | 772.7812
66 | 16482 | 772.5938
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16483 | 772.6406
66 | 16487 | 772.8281
66 | 16485 | 772.7344
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16492 | 773.0625
66 | 16482 | 772.5938
66 | 16490 | 772.9688
66 | 16481 | 772.5469
66 | 16491 | 773.0156
66 | 16480 | 772.5
66 | 16490 | 772.9688
66 | 16478 | 772.4062
66 | 16491 | 773.0156
66 | 16481 | 772.5469
66 | 16490 | 772.9688
66 | 16482 | 772.5938
66 | 16487 | 772.8281
66 | 16494 | 773.1562
66 | 16483 | 772.6406
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16484 | 772.6875
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16484 | 772.6875
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16495 | 773.2031
66 | 16497 | 773.2969
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16492 | 773.0625
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16498 | 773.3438
66 | 16496 | 773.25
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16496 | 773.25
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16497 | 773.2969
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16501 | 773.4844
66 | 16497 | 773.2969
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16496 | 773.25
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16506 | 773.7188
66 | 16509 | 773.8594
66 | 16509 | 773.8594
66 | 16506 | 773.7188
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16510 | 773.9062
66 | 16513 | 774.0469
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16514 | 774.0938
66 | 16512 | 774
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16514 | 774.0938
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16513 | 774.0469
66 | 16503 | 773.5781
66 | 16508 | 773.8125
66 | 16513 | 774.0469
66 | 16503 | 773.5781
66 | 16516 | 774.1875
66 | 16513 | 774.0469
66 | 16508 | 773.8125
66 | 16517 | 774.2344
66 | 16507 | 773.7656
66 | 16517 | 774.2344
66 | 16507 | 773.7656
66 | 16516 | 774.1875
66 | 16509 | 773.8594
66 | 16514 | 774.0938
66 | 16511 | 773.9531
66 | 16515 | 774.1406
66 | 16511 | 773.9531
66 | 16517 | 774.2344
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16513 | 774.0469
66 | 16518 | 774.2812
66 | 16513 | 774.0469
66 | 16519 | 774.3281
66 | 16516 | 774.1875
66 | 16517 | 774.2344
66 | 16516 | 774.1875
66 | 16519 | 774.3281
66 | 16514 | 774.0938
66 | 16515 | 774.1406
66 | 16515 | 774.1406
66 | 16515 | 774.1406
66 | 16519 | 774.3281
66 | 16515 | 774.1406
66 | 16518 | 774.2812
66 | 16515 | 774.1406
66 | 16521 | 774.4219
66 | 16515 | 774.1406
66 | 16524 | 774.5625
66 | 16512 | 774
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16515 | 774.1406
66 | 16524 | 774.5625
66 | 16516 | 774.1875
66 | 16517 | 774.2344
66 | 16522 | 774.4688
66 | 16512 | 774
66 | 16521 | 774.4219
66 | 16521 | 774.4219
66 | 16515 | 774.1406
66 | 16525 | 774.6094
66 | 16523 | 774.5156
66 | 16516 | 774.1875
66 | 16527 | 774.7031
66 | 16527 | 774.7031
66 | 16518 | 774.2812
66 | 16518 | 774.2812
66 | 16526 | 774.6562
66 | 16526 | 774.6562
66 | 16518 | 774.2812
66 | 16517 | 774.2344
66 | 16525 | 774.6094
66 | 16527 | 774.7031
66 | 16524 | 774.5625
66 | 16517 | 774.2344
66 | 16521 | 774.4219
66 | 16528 | 774.75
66 | 16530 | 774.8438
66 | 16528 | 774.75
66 | 16524 | 774.5625
66 | 16519 | 774.3281
66 | 16522 | 774.4688
66 | 16529 | 774.7969
66 | 16530 | 774.8438
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16520 | 774.375
66 | 16520 | 774.375
66 | 16526 | 774.6562
66 | 16529 | 774.7969
66 | 16532 | 774.9375
66 | 16532 | 774.9375
66 | 16533 | 774.9844
66 | 16533 | 774.9844
66 | 16531 | 774.8906
66 | 16528 | 774.75
66 | 16525 | 774.6094
66 | 16523 | 774.5156
66 | 16523 | 774.5156
66 | 16523 | 774.5156
66 | 16522 | 774.4688
66 | 16522 | 774.4688
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16525 | 774.6094
66 | 16525 | 774.6094
66 | 16525 | 774.6094
66 | 16528 | 774.75
66 | 16529 | 774.7969
66 | 16530 | 774.8438
66 | 16530 | 774.8438
66 | 16528 | 774.75
66 | 16529 | 774.7969
66 | 16531 | 774.8906
66 | 16529 | 774.7969
66 | 16533 | 774.9844
66 | 16533 | 774.9844
66 | 16533 | 774.9844
66 | 16534 | 775.0312
66 | 16535 | 775.0781
66 | 16537 | 775.1719
66 | 16538 | 775.2188
66 | 16540 | 775.3125
66 | 16537 | 775.1719
66 | 16538 | 775.2188
66 | 16537 | 775.1719
66 | 16530 | 774.8438
66 | 16528 | 774.75
66 | 16530 | 774.8438
66 | 16535 | 775.0781
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16538 | 775.2188
66 | 16535 | 775.0781
66 | 16530 | 774.8438
66 | 16530 | 774.8438
66 | 16537 | 775.1719
66 | 16542 | 775.4062
66 | 16543 | 775.4531
66 | 16540 | 775.3125
66 | 16533 | 774.9844
66 | 16532 | 774.9375
66 | 16537 | 775.1719
66 | 16543 | 775.4531
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16532 | 774.9375
66 | 16539 | 775.2656
66 | 16543 | 775.4531
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16544 | 775.5
66 | 16542 | 775.4062
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16547 | 775.6406
66 | 16542 | 775.4062
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16547 | 775.6406
66 | 16545 | 775.5469
66 | 16541 | 775.3594
66 | 16548 | 775.6875
66 | 16544 | 775.5
66 | 16544 | 775.5
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16546 | 775.5938
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16541 | 775.3594
66 | 16544 | 775.5
66 | 16549 | 775.7344
66 | 16539 | 775.2656
66 | 16551 | 775.8281
66 | 16541 | 775.3594
66 | 16547 | 775.6406
66 | 16549 | 775.7344
66 | 16548 | 775.6875
66 | 16552 | 775.875
66 | 16542 | 775.4062
66 | 16553 | 775.9219
66 | 16541 | 775.3594
66 | 16548 | 775.6875
66 | 16546 | 775.5938
66 | 16546 | 775.5938
66 | 16548 | 775.6875
66 | 16544 | 775.5
66 | 16550 | 775.7812
66 | 16547 | 775.6406
66 | 16550 | 775.7812
66 | 16550 | 775.7812
66 | 16550 | 775.7812
66 | 16552 | 775.875
66 | 16553 | 775.9219
66 | 16551 | 775.8281
66 | 16552 | 775.875
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16555 | 776.0156
66 | 16544 | 775.5
66 | 16556 | 776.0625
66 | 16544 | 775.5
66 | 16554 | 775.9688
66 | 16551 | 775.8281
66 | 16550 | 775.7812
66 | 16556 | 776.0625
66 | 16546 | 775.5938
66 | 16557 | 776.1094
66 | 16554 | 775.9688
66 | 16550 | 775.7812
66 | 16559 | 776.2031
66 | 16549 | 775.7344
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16547 | 775.6406
66 | 16556 | 776.0625
66 | 16554 | 775.9688
66 | 16546 | 775.5938
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16548 | 775.6875
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16551 | 775.8281
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16549 | 775.7344
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16552 | 775.875
66 | 16553 | 775.9219
66 | 16560 | 776.25
66 | 16556 | 776.0625
66 | 16549 | 775.7344
66 | 16557 | 776.1094
66 | 16562 | 776.3438
66 | 16555 | 776.0156
66 | 16555 | 776.0156
66 | 16563 | 776.3906
66 | 16561 | 776.2969
66 | 16555 | 776.0156
66 | 16560 | 776.25
66 | 16564 | 776.4375
66 | 16558 | 776.1562
66 | 16554 | 775.9688
66 | 16563 | 776.3906
66 | 16564 | 776.4375
66 | 16560 | 776.25
66 | 16552 | 775.875
66 | 16561 | 776.2969
66 | 16565 | 776.4844
66 | 16564 | 776.4375
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16566 | 776.5312
66 | 16569 | 776.6719
66 | 16560 | 776.25
66 | 16557 | 776.1094
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16563 | 776.3906
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16562 | 776.3438
66 | 16565 | 776.4844
66 | 16568 | 776.625
66 | 16566 | 776.5312
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16562 | 776.3438
66 | 16569 | 776.6719
66 | 16567 | 776.5781
66 | 16571 | 776.7656
66 | 16568 | 776.625
66 | 16563 | 776.3906
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16561 | 776.2969
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16569 | 776.6719
66 | 16569 | 776.6719
66 | 16571 | 776.7656
66 | 16573 | 776.8594
66 | 16570 | 776.7188
66 | 16569 | 776.6719
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16564 | 776.4375
66 | 16562 | 776.3438
66 | 16564 | 776.4375
66 | 16568 | 776.625
66 | 16570 | 776.7188
66 | 16569 | 776.6719
66 | 16572 | 776.8125
66 | 16572 | 776.8125
66 | 16573 | 776.8594
66 | 16574 | 776.9062
66 | 16574 | 776.9062
66 | 16573 | 776.8594
66 | 16574 | 776.9062
66 | 16574 | 776.9062
66 | 16573 | 776.8594
66 | 16574 | 776.9062
66 | 16574 | 776.9062
66 | 16572 | 776.8125
66 | 16572 | 776.8125
66 | 16571 | 776.7656
66 | 16572 | 776.8125
66 | 16572 | 776.8125
66 | 16574 | 776.9062
66 | 16574 | 776.9062
66 | 16574 | 776.9062
66 | 16572 | 776.8125
66 | 16571 | 776.7656
66 | 16568 | 776.625
66 | 16568 | 776.625
66 | 16567 | 776.5781
66 | 16568 | 776.625
66 | 16568 | 776.625
66 | 16568 | 776.625
66 | 16571 | 776.7656
66 | 16570 | 776.7188
66 | 16574 | 776.9062
66 | 16576 | 777
66 | 16578 | 777.0938
66 | 16576 | 777
66 | 16580 | 777.1875
66 | 16579 | 777.1406
66 | 16581 | 777.2344
66 | 16581 | 777.2344
66 | 16579 | 777.1406
66 | 16579 | 777.1406
66 | 16574 | 776.9062
66 | 16571 | 776.7656
66 | 16574 | 776.9062
66 | 16575 | 776.9531
66 | 16579 | 777.1406
66 | 16577 | 777.0469
66 | 16581 | 777.2344
66 | 16580 | 777.1875
66 | 16578 | 777.0938
66 | 16575 | 776.9531
66 | 16572 | 776.8125
66 | 16575 | 776.9531
66 | 16582 | 777.2812
66 | 16583 | 777.3281
66 | 16583 | 777.3281
66 | 16581 | 777.2344
66 | 16572 | 776.8125
66 | 16573 | 776.8594
66 | 16578 | 777.0938
66 | 16581 | 777.2344
66 | 16583 | 777.3281
66 | 16581 | 777.2344
66 | 16572 | 776.8125
66 | 16574 | 776.9062
66 | 16581 | 777.2344
66 | 16584 | 777.375
66 | 16582 | 777.2812
66 | 16581 | 777.2344
66 | 16576 | 777
66 | 16584 | 777.375
66 | 16586 | 777.4688
66 | 16585 | 777.4219
66 | 16579 | 777.1406
66 | 16575 | 776.9531
66 | 16582 | 777.2812
66 | 16585 | 777.4219
66 | 16583 | 777.3281
66 | 16576 | 777
66 | 16579 | 777.1406
66 | 16586 | 777.4688
66 | 16585 | 777.4219
66 | 16580 | 777.1875
66 | 16582 | 777.2812
66 | 16588 | 777.5625
66 | 16587 | 777.5156
66 | 16582 | 777.2812
66 | 16585 | 777.4219
66 | 16589 | 777.6094
66 | 16588 | 777.5625
66 | 16578 | 777.0938
66 | 16586 | 777.4688
66 | 16589 | 777.6094
66 | 16582 | 777.2812
66 | 16577 | 777.0469
66 | 16588 | 777.5625
66 | 16589 | 777.6094
66 | 16582 | 777.2812
66 | 16580 | 777.1875
66 | 16589 | 777.6094
66 | 16589 | 777.6094
66 | 16580 | 777.1875
66 | 16585 | 777.4219
66 | 16593 | 777.7969
66 | 16588 | 777.5625
66 | 16582 | 777.2812
66 | 16589 | 777.6094
66 | 16590 | 777.6562
66 | 16582 | 777.2812
66 | 16587 | 777.5156
66 | 16591 | 777.7031
66 | 16585 | 777.4219
66 | 16582 | 777.2812
66 | 16593 | 777.7969
66 | 16590 | 777.6562
66 | 16581 | 777.2344
66 | 16590 | 777.6562
66 | 16591 | 777.7031
66 | 16581 | 777.2344
66 | 16594 | 777.8438
66 | 16592 | 777.75
66 | 16584 | 777.375
66 | 16593 | 777.7969
66 | 16592 | 777.75
66 | 16582 | 777.2812
66 | 16594 | 777.8438
66 | 16591 | 777.7031
66 | 16583 | 777.3281
66 | 16592 | 777.75
66 | 16592 | 777.75
66 | 16584 | 777.375
66 | 16594 | 777.8438
66 | 16592 | 777.75
66 | 16587 | 777.5156
66 | 16597 | 777.9844
66 | 16588 | 777.5625
66 | 16594 | 777.8438
66 | 16597 | 777.9844
66 | 16587 | 777.5156
66 | 16596 | 777.9375
66 | 16595 | 777.8906
66 | 16588 | 777.5625
66 | 16597 | 777.9844
66 | 16587 | 777.5156
66 | 16594 | 777.8438
66 | 16594 | 777.8438
66 | 16587 | 777.5156
66 | 16598 | 778.0312
66 | 16591 | 777.7031
66 | 16597 | 777.9844
66 | 16598 | 778.0312
66 | 16591 | 777.7031
66 | 16599 | 778.0781
66 | 16591 | 777.7031
66 | 16597 | 777.9844
66 | 16596 | 777.9375
66 | 16588 | 777.5625
66 | 16600 | 778.125
66 | 16590 | 777.6562
66 | 16597 | 777.9844
66 | 16596 | 777.9375
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16590 | 777.6562
66 | 16599 | 778.0781
66 | 16593 | 777.7969
66 | 16600 | 778.125
66 | 16599 | 778.0781
66 | 16594 | 777.8438
66 | 16602 | 778.2188
66 | 16594 | 777.8438
66 | 16600 | 778.125
66 | 16596 | 777.9375
66 | 16597 | 777.9844
66 | 16599 | 778.0781
66 | 16590 | 777.6562
66 | 16602 | 778.2188
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16595 | 777.8906
66 | 16600 | 778.125
66 | 16598 | 778.0312
66 | 16602 | 778.2188
66 | 16596 | 777.9375
66 | 16604 | 778.3125
66 | 16596 | 777.9375
66 | 16601 | 778.1719
66 | 16594 | 777.8438
66 | 16603 | 778.2656
66 | 16593 | 777.7969
66 | 16604 | 778.3125
66 | 16592 | 777.75
66 | 16604 | 778.3125
66 | 16593 | 777.7969
66 | 16605 | 778.3594
66 | 16594 | 777.8438
66 | 16606 | 778.4062
66 | 16594 | 777.8438
66 | 16607 | 778.4531
66 | 16596 | 777.9375
66 | 16608 | 778.5
66 | 16596 | 777.9375
66 | 16606 | 778.4062
66 | 16596 | 777.9375
66 | 16607 | 778.4531
66 | 16595 | 777.8906
66 | 16606 | 778.4062
66 | 16595 | 777.8906
66 | 16606 | 778.4062
66 | 16595 | 777.8906
66 | 16608 | 778.5
66 | 16597 | 777.9844
66 | 16607 | 778.4531
66 | 16598 | 778.0312
66 | 16609 | 778.5469
66 | 16598 | 778.0312
66 | 16610 | 778.5938
66 | 16599 | 778.0781
66 | 16609 | 778.5469
66 | 16597 | 777.9844
66 | 16608 | 778.5
66 | 16597 | 777.9844
66 | 16608 | 778.5
66 | 16597 | 777.9844
66 | 16610 | 778.5938
66 | 16597 | 777.9844
66 | 16609 | 778.5469
66 | 16599 | 778.0781
66 | 16610 | 778.5938
66 | 16603 | 778.2656
66 | 16610 | 778.5938
66 | 16604 | 778.3125
66 | 16609 | 778.5469
66 | 16605 | 778.3594
66 | 16610 | 778.5938
66 | 16606 | 778.4062
66 | 16608 | 778.5
66 | 16604 | 778.3125
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16607 | 778.4531
66 | 16614 | 778.7812
66 | 16603 | 778.2656
66 | 16613 | 778.7344
66 | 16605 | 778.3594
66 | 16611 | 778.6406
66 | 16605 | 778.3594
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16605 | 778.3594
66 | 16617 | 778.9219
66 | 16605 | 778.3594
66 | 16616 | 778.875
66 | 16611 | 778.6406
66 | 16614 | 778.7812
66 | 16612 | 778.6875
66 | 16615 | 778.8281
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16616 | 778.875
66 | 16606 | 778.4062
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16613 | 778.7344
66 | 16611 | 778.6406
66 | 16614 | 778.7812
66 | 16615 | 778.8281
66 | 16609 | 778.5469
66 | 16621 | 779.1094
66 | 16610 | 778.5938
66 | 16618 | 778.9688
66 | 16614 | 778.7812
66 | 16617 | 778.9219
66 | 16614 | 778.7812
66 | 16616 | 778.875
66 | 16616 | 778.875
66 | 16607 | 778.4531
66 | 16616 | 778.875
66 | 16612 | 778.6875
66 | 16615 | 778.8281
66 | 16615 | 778.8281
66 | 16607 | 778.4531
66 | 16620 | 779.0625
66 | 16614 | 778.7812
66 | 16619 | 779.0156
66 | 16618 | 778.9688
66 | 16618 | 778.9688
66 | 16619 | 779.0156
66 | 16609 | 778.5469
66 | 16618 | 778.9688
66 | 16615 | 778.8281
66 | 16616 | 778.875
66 | 16615 | 778.8281
66 | 16609 | 778.5469
66 | 16620 | 779.0625
66 | 16612 | 778.6875
66 | 16619 | 779.0156
66 | 16617 | 778.9219
66 | 16612 | 778.6875
66 | 16621 | 779.1094
66 | 16613 | 778.7344
66 | 16619 | 779.0156
66 | 16619 | 779.0156
66 | 16612 | 778.6875
66 | 16620 | 779.0625
66 | 16619 | 779.0156
66 | 16620 | 779.0625
66 | 16621 | 779.1094
66 | 16612 | 778.6875
66 | 16620 | 779.0625
66 | 16616 | 778.875
66 | 16619 | 779.0156
66 | 16623 | 779.2031
66 | 16612 | 778.6875
66 | 16620 | 779.0625
66 | 16620 | 779.0625
66 | 16620 | 779.0625
66 | 16624 | 779.25
66 | 16615 | 778.8281
66 | 16624 | 779.25
66 | 16622 | 779.1562
66 | 16614 | 778.7812
66 | 16624 | 779.25
66 | 16620 | 779.0625
66 | 16621 | 779.1094
66 | 16624 | 779.25
66 | 16613 | 778.7344
66 | 16622 | 779.1562
66 | 16620 | 779.0625
66 | 16616 | 778.875
66 | 16625 | 779.2969
66 | 16617 | 778.9219
66 | 16624 | 779.25
66 | 16624 | 779.25
66 | 16618 | 778.9688
66 | 16627 | 779.3906
66 | 16621 | 779.1094
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16617 | 778.9219
66 | 16624 | 779.25
66 | 16622 | 779.1562
66 | 16621 | 779.1094
66 | 16625 | 779.2969
66 | 16615 | 778.8281
66 | 16625 | 779.2969
66 | 16623 | 779.2031
66 | 16618 | 778.9688
66 | 16628 | 779.4375
66 | 16625 | 779.2969
66 | 16627 | 779.3906
66 | 16629 | 779.4844
66 | 16619 | 779.0156
66 | 16627 | 779.3906
66 | 16623 | 779.2031
66 | 16625 | 779.2969
66 | 16625 | 779.2969
66 | 16618 | 778.9688
66 | 16626 | 779.3438
66 | 16625 | 779.2969
66 | 16620 | 779.0625
66 | 16627 | 779.3906
66 | 16623 | 779.2031
66 | 16626 | 779.3438
66 | 16630 | 779.5312
66 | 16621 | 779.1094
66 | 16631 | 779.5781
66 | 16632 | 779.625
66 | 16622 | 779.1562
66 | 16629 | 779.4844
66 | 16628 | 779.4375
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16627 | 779.3906
66 | 16626 | 779.3438
66 | 16630 | 779.5312
66 | 16618 | 778.9688
66 | 16627 | 779.3906
66 | 16626 | 779.3438
66 | 16625 | 779.2969
66 | 16632 | 779.625
66 | 16628 | 779.4375
66 | 16630 | 779.5312
66 | 16633 | 779.6719
66 | 16629 | 779.4844
66 | 16628 | 779.4375
66 | 16633 | 779.6719
66 | 16625 | 779.2969
66 | 16632 | 779.625
66 | 16633 | 779.6719
66 | 16622 | 779.1562
66 | 16630 | 779.5312
66 | 16630 | 779.5312
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16630 | 779.5312
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16629 | 779.4844
66 | 16628 | 779.4375
66 | 16635 | 779.7656
66 | 16632 | 779.625
66 | 16631 | 779.5781
66 | 16637 | 779.8594
66 | 16625 | 779.2969
66 | 16631 | 779.5781
66 | 16633 | 779.6719
66 | 16625 | 779.2969
66 | 16631 | 779.5781
66 | 16634 | 779.7188
66 | 16624 | 779.25
66 | 16633 | 779.6719
66 | 16632 | 779.625
66 | 16625 | 779.2969
66 | 16635 | 779.7656
66 | 16635 | 779.7656
66 | 16630 | 779.5312
66 | 16637 | 779.8594
66 | 16633 | 779.6719
66 | 16632 | 779.625
66 | 16638 | 779.9062
66 | 16630 | 779.5312
66 | 16634 | 779.7188
66 | 16636 | 779.8125
66 | 16625 | 779.2969
66 | 16634 | 779.7188
66 | 16636 | 779.8125
66 | 16626 | 779.3438
66 | 16636 | 779.8125
66 | 16635 | 779.7656
66 | 16629 | 779.4844
66 | 16638 | 779.9062
66 | 16636 | 779.8125
66 | 16634 | 779.7188
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16637 | 779.8594
66 | 16638 | 779.9062
66 | 16628 | 779.4375
66 | 16638 | 779.9062
66 | 16635 | 779.7656
66 | 16630 | 779.5312
66 | 16639 | 779.9531
66 | 16635 | 779.7656
66 | 16637 | 779.8594
66 | 16640 | 780
66 | 16630 | 779.5312
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16639 | 779.9531
66 | 16639 | 779.9531
66 | 16641 | 780.0469
66 | 16632 | 779.625
66 | 16639 | 779.9531
66 | 16638 | 779.9062
66 | 16629 | 779.4844
66 | 16637 | 779.8594
66 | 16637 | 779.8594
66 | 16632 | 779.625
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16637 | 779.8594
66 | 16642 | 780.0938
66 | 16638 | 779.9062
66 | 16638 | 779.9062
66 | 16645 | 780.2344
66 | 16637 | 779.8594
66 | 16640 | 780
66 | 16643 | 780.1406
66 | 16632 | 779.625
66 | 16641 | 780.0469
66 | 16640 | 780
66 | 16635 | 779.7656
66 | 16642 | 780.0938
66 | 16638 | 779.9062
66 | 16637 | 779.8594
66 | 16644 | 780.1875
66 | 16637 | 779.8594
66 | 16643 | 780.1406
66 | 16645 | 780.2344
66 | 16636 | 779.8125
66 | 16645 | 780.2344
66 | 16643 | 780.1406
66 | 16635 | 779.7656
66 | 16645 | 780.2344
66 | 16642 | 780.0938
66 | 16642 | 780.0938
66 | 16646 | 780.2812
66 | 16633 | 779.6719
66 | 16640 | 780
66 | 16643 | 780.1406
66 | 16634 | 779.7188
66 | 16645 | 780.2344
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16646 | 780.2812
66 | 16648 | 780.375
66 | 16639 | 779.9531
66 | 16646 | 780.2812
66 | 16648 | 780.375
66 | 16636 | 779.8125
66 | 16645 | 780.2344
66 | 16646 | 780.2812
66 | 16636 | 779.8125
66 | 16644 | 780.1875
66 | 16646 | 780.2812
66 | 16637 | 779.8594
66 | 16645 | 780.2344
66 | 16646 | 780.2812
66 | 16637 | 779.8594
66 | 16648 | 780.375
66 | 16647 | 780.3281
66 | 16644 | 780.1875
66 | 16649 | 780.4219
66 | 16643 | 780.1406
66 | 16646 | 780.2812
66 | 16647 | 780.3281
66 | 16638 | 779.9062
66 | 16646 | 780.2812
66 | 16644 | 780.1875
66 | 16643 | 780.1406
66 | 16647 | 780.3281
66 | 16642 | 780.0938
66 | 16646 | 780.2812
66 | 16647 | 780.3281
66 | 16638 | 779.9062
66 | 16649 | 780.4219
66 | 16646 | 780.2812
66 | 16648 | 780.375
66 | 16650 | 780.4688
66 | 16640 | 780
66 | 16649 | 780.4219
66 | 16647 | 780.3281
66 | 16640 | 780
66 | 16650 | 780.4688
66 | 16645 | 780.2344
66 | 16647 | 780.3281
66 | 16649 | 780.4219
66 | 16639 | 779.9531
66 | 16647 | 780.3281
66 | 16650 | 780.4688
66 | 16641 | 780.0469
66 | 16650 | 780.4688
66 | 16649 | 780.4219
66 | 16646 | 780.2812
66 | 16651 | 780.5156
66 | 16648 | 780.375
66 | 16649 | 780.4219
66 | 16654 | 780.6562
66 | 16641 | 780.0469
66 | 16651 | 780.5156
66 | 16648 | 780.375
66 | 16643 | 780.1406
66 | 16650 | 780.4688
66 | 16645 | 780.2344
66 | 16649 | 780.4219
66 | 16648 | 780.375
66 | 16642 | 780.0938
66 | 16653 | 780.6094
66 | 16649 | 780.4219
66 | 16651 | 780.5156
66 | 16654 | 780.6562
66 | 16645 | 780.2344
66 | 16652 | 780.5625
66 | 16651 | 780.5156
66 | 16645 | 780.2344
66 | 16653 | 780.6094
66 | 16646 | 780.2812
66 | 16651 | 780.5156
66 | 16651 | 780.5156
66 | 16642 | 780.0938
66 | 16653 | 780.6094
66 | 16649 | 780.4219
66 | 16651 | 780.5156
66 | 16654 | 780.6562
66 | 16644 | 780.1875
66 | 16652 | 780.5625
66 | 16652 | 780.5625
66 | 16654 | 780.6562
66 | 16653 | 780.6094
66 | 16653 | 780.6094
66 | 16657 | 780.7969
66 | 16645 | 780.2344
66 | 16654 | 780.6562
66 | 16649 | 780.4219
66 | 16651 | 780.5156
66 | 16651 | 780.5156
66 | 16645 | 780.2344
66 | 16657 | 780.7969
66 | 16648 | 780.375
66 | 16655 | 780.7031
66 | 16651 | 780.5156
66 | 16656 | 780.75
66 | 16654 | 780.6562
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16651 | 780.5156
66 | 16660 | 780.9375
66 | 16649 | 780.4219
66 | 16655 | 780.7031
66 | 16653 | 780.6094
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16651 | 780.5156
66 | 16656 | 780.75
66 | 16646 | 780.2812
66 | 16658 | 780.8438
66 | 16651 | 780.5156
66 | 16655 | 780.7031
66 | 16656 | 780.75
66 | 16658 | 780.8438
66 | 16655 | 780.7031
66 | 16658 | 780.8438
66 | 16656 | 780.75
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16650 | 780.4688
66 | 16659 | 780.8906
66 | 16648 | 780.375
66 | 16657 | 780.7969
66 | 16651 | 780.5156
66 | 16658 | 780.8438
66 | 16653 | 780.6094
66 | 16660 | 780.9375
66 | 16656 | 780.75
66 | 16659 | 780.8906
66 | 16656 | 780.75
66 | 16660 | 780.9375
66 | 16656 | 780.75
66 | 16661 | 780.9844
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16656 | 780.75
66 | 16658 | 780.8438
66 | 16656 | 780.75
66 | 16659 | 780.8906
66 | 16654 | 780.6562
66 | 16659 | 780.8906
66 | 16656 | 780.75
66 | 16661 | 780.9844
66 | 16657 | 780.7969
66 | 16662 | 781.0312
66 | 16656 | 780.75
66 | 16662 | 781.0312
66 | 16656 | 780.75
66 | 16661 | 780.9844
66 | 16654 | 780.6562
66 | 16661 | 780.9844
66 | 16653 | 780.6094
66 | 16660 | 780.9375
66 | 16653 | 780.6094
66 | 16661 | 780.9844
66 | 16651 | 780.5156
66 | 16662 | 781.0312
66 | 16654 | 780.6562
66 | 16663 | 781.0781
66 | 16657 | 780.7969
66 | 16663 | 781.0781
66 | 16658 | 780.8438
66 | 16664 | 781.125
66 | 16659 | 780.8906
66 | 16664 | 781.125
66 | 16660 | 780.9375
66 | 16663 | 781.0781
66 | 16656 | 780.75
66 | 16663 | 781.0781
66 | 16656 | 780.75
66 | 16664 | 781.125
66 | 16655 | 780.7031
66 | 16663 | 781.0781
66 | 16655 | 780.7031
66 | 16664 | 781.125
66 | 16654 | 780.6562
66 | 16665 | 781.1719
66 | 16657 | 780.7969
66 | 16667 | 781.2656
66 | 16657 | 780.7969
66 | 16666 | 781.2188
66 | 16658 | 780.8438
66 | 16666 | 781.2188
66 | 16657 | 780.7969
66 | 16666 | 781.2188
66 | 16655 | 780.7031
66 | 16666 | 781.2188
66 | 16654 | 780.6562
66 | 16667 | 781.2656
66 | 16656 | 780.75
66 | 16667 | 781.2656
66 | 16658 | 780.8438
66 | 16669 | 781.3594
66 | 16658 | 780.8438
66 | 16669 | 781.3594
66 | 16659 | 780.8906
66 | 16670 | 781.4062
66 | 16659 | 780.8906
66 | 16669 | 781.3594
66 | 16659 | 780.8906
66 | 16669 | 781.3594
66 | 16658 | 780.8438
66 | 16669 | 781.3594
66 | 16660 | 780.9375
66 | 16667 | 781.2656
66 | 16660 | 780.9375
66 | 16668 | 781.3125
66 | 16660 | 780.9375
66 | 16669 | 781.3594
66 | 16659 | 780.8906
66 | 16670 | 781.4062
66 | 16662 | 781.0312
66 | 16670 | 781.4062
66 | 16663 | 781.0781
66 | 16670 | 781.4062
66 | 16663 | 781.0781
66 | 16670 | 781.4062
66 | 16667 | 781.2656
66 | 16663 | 781.0781
66 | 16668 | 781.3125
66 | 16659 | 780.8906
66 | 16669 | 781.3594
66 | 16661 | 780.9844
66 | 16669 | 781.3594
66 | 16663 | 781.0781
66 | 16669 | 781.3594
66 | 16665 | 781.1719
66 | 16668 | 781.3125
66 | 16671 | 781.4531
66 | 16661 | 780.9844
66 | 16674 | 781.5938
66 | 16662 | 781.0312
66 | 16673 | 781.5469
66 | 16666 | 781.2188
66 | 16671 | 781.4531
66 | 16669 | 781.3594
66 | 16664 | 781.125
66 | 16673 | 781.5469
66 | 16661 | 780.9844
66 | 16671 | 781.4531
66 | 16663 | 781.0781
66 | 16670 | 781.4062
66 | 16673 | 781.5469
66 | 16664 | 781.125
66 | 16676 | 781.6875
66 | 16671 | 781.4531
66 | 16669 | 781.3594
66 | 16672 | 781.5
66 | 16663 | 781.0781
66 | 16674 | 781.5938
66 | 16665 | 781.1719
66 | 16670 | 781.4062
66 | 16671 | 781.4531
66 | 16661 | 780.9844
66 | 16674 | 781.5938
66 | 16665 | 781.1719
66 | 16673 | 781.5469
66 | 16672 | 781.5
66 | 16666 | 781.2188
66 | 16675 | 781.6406
66 | 16665 | 781.1719
66 | 16672 | 781.5
66 | 16671 | 781.4531
66 | 16671 | 781.4531
66 | 16674 | 781.5938
66 | 16665 | 781.1719
66 | 16675 | 781.6406
66 | 16672 | 781.5
66 | 16669 | 781.3594
66 | 16675 | 781.6406
66 | 16663 | 781.0781
66 | 16671 | 781.4531
66 | 16672 | 781.5
66 | 16664 | 781.125
66 | 16675 | 781.6406
66 | 16672 | 781.5
66 | 16672 | 781.5
66 | 16677 | 781.7344
66 | 16666 | 781.2188
66 | 16674 | 781.5938
66 | 16674 | 781.5938
66 | 16666 | 781.2188
66 | 16676 | 781.6875
66 | 16675 | 781.6406
66 | 16667 | 781.2656
66 | 16675 | 781.6406
66 | 16665 | 781.1719
66 | 16670 | 781.4062
66 | 16672 | 781.5
66 | 16666 | 781.2188
66 | 16675 | 781.6406
66 | 16674 | 781.5938
66 | 16669 | 781.3594
66 | 16678 | 781.7812
66 | 16673 | 781.5469
66 | 16672 | 781.5
66 | 16678 | 781.7812
66 | 16668 | 781.3125
66 | 16677 | 781.7344
66 | 16675 | 781.6406
66 | 16669 | 781.3594
66 | 16677 | 781.7344
66 | 16669 | 781.3594
66 | 16672 | 781.5
66 | 16676 | 781.6875
66 | 16666 | 781.2188
66 | 16675 | 781.6406
66 | 16677 | 781.7344
66 | 16668 | 781.3125
66 | 16678 | 781.7812
66 | 16678 | 781.7812
66 | 16669 | 781.3594
66 | 16679 | 781.8281
66 | 16678 | 781.7812
66 | 16668 | 781.3125
66 | 16677 | 781.7344
66 | 16677 | 781.7344
66 | 16668 | 781.3125
66 | 16677 | 781.7344
66 | 16678 | 781.7812
66 | 16668 | 781.3125
66 | 16677 | 781.7344
66 | 16680 | 781.875
66 | 16670 | 781.4062
66 | 16679 | 781.8281
66 | 16681 | 781.9219
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16682 | 781.9688
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16680 | 781.875
66 | 16669 | 781.3594
66 | 16675 | 781.6406
66 | 16679 | 781.8281
66 | 16673 | 781.5469
66 | 16675 | 781.6406
66 | 16682 | 781.9688
66 | 16673 | 781.5469
66 | 16677 | 781.7344
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16677 | 781.7344
66 | 16685 | 782.1094
66 | 16679 | 781.8281
66 | 16673 | 781.5469
66 | 16683 | 782.0156
66 | 16681 | 781.9219
66 | 16671 | 781.4531
66 | 16678 | 781.7812
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16672 | 781.5
66 | 16678 | 781.7812
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16677 | 781.7344
66 | 16683 | 782.0156
66 | 16681 | 781.9219
66 | 16675 | 781.6406
66 | 16680 | 781.875
66 | 16685 | 782.1094
66 | 16680 | 781.875
66 | 16674 | 781.5938
66 | 16679 | 781.8281
66 | 16683 | 782.0156
66 | 16681 | 781.9219
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16683 | 782.0156
66 | 16684 | 782.0625
66 | 16679 | 781.8281
66 | 16677 | 781.7344
66 | 16684 | 782.0625
66 | 16684 | 782.0625
66 | 16678 | 781.7812
66 | 16679 | 781.8281
66 | 16686 | 782.1562
66 | 16684 | 782.0625
66 | 16678 | 781.7812
66 | 16678 | 781.7812
66 | 16684 | 782.0625
66 | 16683 | 782.0156
66 | 16676 | 781.6875
66 | 16680 | 781.875
66 | 16686 | 782.1562
66 | 16684 | 782.0625
66 | 16676 | 781.6875
66 | 16682 | 781.9688
66 | 16686 | 782.1562
66 | 16686 | 782.1562
66 | 16681 | 781.9219
66 | 16681 | 781.9219
66 | 16685 | 782.1094
66 | 16686 | 782.1562
66 | 16683 | 782.0156
66 | 16676 | 781.6875
66 | 16680 | 781.875
66 | 16685 | 782.1094
66 | 16685 | 782.1094
66 | 16681 | 781.9219
66 | 16677 | 781.7344
66 | 16682 | 781.9688
66 | 16687 | 782.2031
66 | 16687 | 782.2031
66 | 16683 | 782.0156
66 | 16680 | 781.875
66 | 16684 | 782.0625
66 | 16690 | 782.3438
66 | 16690 | 782.3438
66 | 16686 | 782.1562
66 | 16679 | 781.8281
66 | 16679 | 781.8281
66 | 16684 | 782.0625
66 | 16687 | 782.2031
66 | 16687 | 782.2031
66 | 16684 | 782.0625
66 | 16680 | 781.875
66 | 16682 | 781.9688
66 | 16687 | 782.2031
66 | 16689 | 782.2969
66 | 16688 | 782.25
66 | 16686 | 782.1562
66 | 16682 | 781.9688
66 | 16684 | 782.0625
66 | 16688 | 782.25
66 | 16690 | 782.3438
66 | 16687 | 782.2031
66 | 16680 | 781.875
66 | 16676 | 781.6875
66 | 16683 | 782.0156
66 | 16687 | 782.2031
66 | 16688 | 782.25
66 | 16689 | 782.2969
66 | 16686 | 782.1562
66 | 16680 | 781.875
66 | 16684 | 782.0625
66 | 16688 | 782.25
66 | 16693 | 782.4844
66 | 16690 | 782.3438
66 | 16685 | 782.1094
66 | 16681 | 781.9219
66 | 16682 | 781.9688
66 | 16688 | 782.25
66 | 16689 | 782.2969
66 | 16691 | 782.3906
66 | 16688 | 782.25
66 | 16685 | 782.1094
66 | 16678 | 781.7812
66 | 16681 | 781.9219
66 | 16686 | 782.1562
66 | 16688 | 782.25
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16690 | 782.3438
66 | 16688 | 782.25
66 | 16682 | 781.9688
66 | 16682 | 781.9688
66 | 16684 | 782.0625
66 | 16685 | 782.1094
66 | 16689 | 782.2969
66 | 16691 | 782.3906
66 | 16691 | 782.3906
66 | 16691 | 782.3906
66 | 16689 | 782.2969
66 | 16689 | 782.2969
66 | 16682 | 781.9688
66 | 16681 | 781.9219
66 | 16685 | 782.1094
66 | 16689 | 782.2969
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16693 | 782.4844
66 | 16689 | 782.2969
66 | 16685 | 782.1094
66 | 16682 | 781.9688
66 | 16683 | 782.0156
66 | 16688 | 782.25
66 | 16689 | 782.2969
66 | 16691 | 782.3906
66 | 16692 | 782.4375
66 | 16693 | 782.4844
66 | 16693 | 782.4844
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16686 | 782.1562
66 | 16684 | 782.0625
66 | 16686 | 782.1562
66 | 16687 | 782.2031
66 | 16689 | 782.2969
66 | 16694 | 782.5312
66 | 16692 | 782.4375
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16694 | 782.5312
66 | 16692 | 782.4375
66 | 16692 | 782.4375
66 | 16691 | 782.3906
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16694 | 782.5312
66 | 16694 | 782.5312
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16692 | 782.4375
66 | 16691 | 782.3906
66 | 16693 | 782.4844
66 | 16691 | 782.3906
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16696 | 782.625
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16691 | 782.3906
66 | 16694 | 782.5312
66 | 16690 | 782.3438
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16696 | 782.625
66 | 16694 | 782.5312
66 | 16694 | 782.5312
66 | 16695 | 782.5781
66 | 16697 | 782.6719
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16694 | 782.5312
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16692 | 782.4375
66 | 16694 | 782.5312
66 | 16691 | 782.3906
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16694 | 782.5312
66 | 16697 | 782.6719
66 | 16696 | 782.625
66 | 16698 | 782.7188
66 | 16697 | 782.6719
66 | 16699 | 782.7656
66 | 16695 | 782.5781
66 | 16697 | 782.6719
66 | 16697 | 782.6719
66 | 16696 | 782.625
66 | 16694 | 782.5312
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16696 | 782.625
66 | 16695 | 782.5781
66 | 16698 | 782.7188
66 | 16696 | 782.625
66 | 16696 | 782.625
66 | 16697 | 782.6719
66 | 16698 | 782.7188
66 | 16696 | 782.625
66 | 16698 | 782.7188
66 | 16698 | 782.7188
66 | 16696 | 782.625
66 | 16696 | 782.625
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16697 | 782.6719
66 | 16695 | 782.5781
66 | 16697 | 782.6719
66 | 16694 | 782.5312
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16696 | 782.625
66 | 16702 | 782.9062
66 | 16701 | 782.8594
66 | 16703 | 782.9531
66 | 16704 | 783
66 | 16702 | 782.9062
66 | 16703 | 782.9531
66 | 16702 | 782.9062
66 | 16699 | 782.7656
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16699 | 782.7656
66 | 16701 | 782.8594
66 | 16704 | 783
66 | 16705 | 783.0469
66 | 16705 | 783.0469
66 | 16707 | 783.1406
66 | 16706 | 783.0938
66 | 16703 | 782.9531
66 | 16698 | 782.7188
66 | 16697 | 782.6719
66 | 16695 | 782.5781
66 | 16693 | 782.4844
66 | 16694 | 782.5312
66 | 16697 | 782.6719
66 | 16702 | 782.9062
66 | 16703 | 782.9531
66 | 16705 | 783.0469
66 | 16706 | 783.0938
66 | 16704 | 783
66 | 16702 | 782.9062
66 | 16697 | 782.6719
66 | 16697 | 782.6719
66 | 16700 | 782.8125
66 | 16701 | 782.8594
66 | 16704 | 783
66 | 16704 | 783
66 | 16707 | 783.1406
66 | 16704 | 783
66 | 16700 | 782.8125
66 | 16697 | 782.6719
66 | 16695 | 782.5781
66 | 16698 | 782.7188
66 | 16704 | 783
66 | 16705 | 783.0469
66 | 16707 | 783.1406
66 | 16709 | 783.2344
66 | 16703 | 782.9531
66 | 16702 | 782.9062
66 | 16698 | 782.7188
66 | 16696 | 782.625
66 | 16701 | 782.8594
66 | 16704 | 783
66 | 16705 | 783.0469
66 | 16706 | 783.0938
66 | 16705 | 783.0469
66 | 16702 | 782.9062
66 | 16696 | 782.625
66 | 16697 | 782.6719
66 | 16703 | 782.9531
66 | 16705 | 783.0469
66 | 16707 | 783.1406
66 | 16707 | 783.1406
66 | 16703 | 782.9531
66 | 16700 | 782.8125
66 | 16699 | 782.7656
66 | 16704 | 783
66 | 16706 | 783.0938
66 | 16708 | 783.1875
66 | 16704 | 783
66 | 16700 | 782.8125
66 | 16698 | 782.7188
66 | 16702 | 782.9062
66 | 16706 | 783.0938
66 | 16707 | 783.1406
66 | 16703 | 782.9531
66 | 16698 | 782.7188
66 | 16701 | 782.8594
66 | 16707 | 783.1406
66 | 16710 | 783.2812
66 | 16710 | 783.2812
66 | 16703 | 782.9531
66 | 16698 | 782.7188
66 | 16703 | 782.9531
66 | 16707 | 783.1406
66 | 16709 | 783.2344
66 | 16708 | 783.1875
66 | 16698 | 782.7188
66 | 16699 | 782.7656
66 | 16706 | 783.0938
66 | 16706 | 783.0938
66 | 16708 | 783.1875
66 | 16702 | 782.9062
66 | 16702 | 782.9062
66 | 16706 | 783.0938
66 | 16710 | 783.2812
66 | 16709 | 783.2344
66 | 16702 | 782.9062
66 | 16701 | 782.8594
66 | 16707 | 783.1406
66 | 16708 | 783.1875
66 | 16711 | 783.3281
66 | 16705 | 783.0469
66 | 16698 | 782.7188
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16710 | 783.2812
66 | 16702 | 782.9062
66 | 16697 | 782.6719
66 | 16705 | 783.0469
66 | 16711 | 783.3281
66 | 16711 | 783.3281
66 | 16705 | 783.0469
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16711 | 783.3281
66 | 16712 | 783.375
66 | 16702 | 782.9062
66 | 16701 | 782.8594
66 | 16708 | 783.1875
66 | 16710 | 783.2812
66 | 16705 | 783.0469
66 | 16699 | 782.7656
66 | 16704 | 783
66 | 16708 | 783.1875
66 | 16710 | 783.2812
66 | 16703 | 782.9531
66 | 16704 | 783
66 | 16711 | 783.3281
66 | 16714 | 783.4688
66 | 16707 | 783.1406
66 | 16703 | 782.9531
66 | 16711 | 783.3281
66 | 16713 | 783.4219
66 | 16710 | 783.2812
66 | 16701 | 782.8594
66 | 16709 | 783.2344
66 | 16712 | 783.375
66 | 16709 | 783.2344
66 | 16701 | 782.8594
66 | 16706 | 783.0938
66 | 16711 | 783.3281
66 | 16708 | 783.1875
66 | 16701 | 782.8594
66 | 16708 | 783.1875
66 | 16714 | 783.4688
66 | 16710 | 783.2812
66 | 16703 | 782.9531
66 | 16711 | 783.3281
66 | 16714 | 783.4688
66 | 16708 | 783.1875
66 | 16702 | 782.9062
66 | 16712 | 783.375
66 | 16710 | 783.2812
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16713 | 783.4219
66 | 16704 | 783
66 | 16706 | 783.0938
66 | 16714 | 783.4688
66 | 16711 | 783.3281
66 | 16703 | 782.9531
66 | 16713 | 783.4219
66 | 16715 | 783.5156
66 | 16709 | 783.2344
66 | 16709 | 783.2344
66 | 16714 | 783.4688
66 | 16710 | 783.2812
66 | 16702 | 782.9062
66 | 16711 | 783.3281
66 | 16715 | 783.5156
66 | 16703 | 782.9531
66 | 16711 | 783.3281
66 | 16713 | 783.4219
66 | 16703 | 782.9531
66 | 16712 | 783.375
66 | 16716 | 783.5625
66 | 16705 | 783.0469
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16706 | 783.0938
66 | 16714 | 783.4688
66 | 16717 | 783.6094
66 | 16705 | 783.0469
66 | 16715 | 783.5156
66 | 16715 | 783.5156
66 | 16706 | 783.0938
66 | 16713 | 783.4219
66 | 16715 | 783.5156
66 | 16702 | 782.9062
66 | 16712 | 783.375
66 | 16713 | 783.4219
66 | 16704 | 783
66 | 16714 | 783.4688
66 | 16716 | 783.5625
66 | 16705 | 783.0469
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16706 | 783.0938
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16704 | 783
66 | 16714 | 783.4688
66 | 16712 | 783.375
66 | 16704 | 783
66 | 16713 | 783.4219
66 | 16711 | 783.3281
66 | 16705 | 783.0469
66 | 16714 | 783.4688
66 | 16713 | 783.4219
66 | 16706 | 783.0938
66 | 16715 | 783.5156
66 | 16717 | 783.6094
66 | 16706 | 783.0938
66 | 16715 | 783.5156
66 | 16715 | 783.5156
66 | 16706 | 783.0938
66 | 16714 | 783.4688
66 | 16715 | 783.5156
66 | 16706 | 783.0938
66 | 16714 | 783.4688
66 | 16710 | 783.2812
66 | 16705 | 783.0469
66 | 16716 | 783.5625
66 | 16713 | 783.4219
66 | 16711 | 783.3281
66 | 16717 | 783.6094
66 | 16707 | 783.1406
66 | 16718 | 783.6562
66 | 16714 | 783.4688
66 | 16708 | 783.1875
66 | 16718 | 783.6562
66 | 16713 | 783.4219
66 | 16714 | 783.4688
66 | 16719 | 783.7031
66 | 16708 | 783.1875
66 | 16717 | 783.6094
66 | 16714 | 783.4688
66 | 16707 | 783.1406
66 | 16717 | 783.6094
66 | 16712 | 783.375
66 | 16710 | 783.2812
66 | 16717 | 783.6094
66 | 16709 | 783.2344
66 | 16718 | 783.6562
66 | 16715 | 783.5156
66 | 16709 | 783.2344
66 | 16721 | 783.7969
66 | 16712 | 783.375
66 | 16718 | 783.6562
66 | 16715 | 783.5156
66 | 16711 | 783.3281
66 | 16719 | 783.7031
66 | 16708 | 783.1875
66 | 16715 | 783.5156
66 | 16711 | 783.3281
66 | 16712 | 783.375
66 | 16719 | 783.7031
66 | 16708 | 783.1875
66 | 16719 | 783.7031
66 | 16712 | 783.375
66 | 16719 | 783.7031
66 | 16717 | 783.6094
66 | 16715 | 783.5156
66 | 16719 | 783.7031
66 | 16711 | 783.3281
66 | 16719 | 783.7031
66 | 16710 | 783.2812
66 | 16719 | 783.7031
66 | 16714 | 783.4688
66 | 16713 | 783.4219
66 | 16716 | 783.5625
66 | 16709 | 783.2344
66 | 16720 | 783.75
66 | 16708 | 783.1875
66 | 16719 | 783.7031
66 | 16716 | 783.5625
66 | 16717 | 783.6094
66 | 16720 | 783.75
66 | 16713 | 783.4219
66 | 16723 | 783.8906
66 | 16711 | 783.3281
66 | 16721 | 783.7969
66 | 16711 | 783.3281
66 | 16719 | 783.7031
66 | 16713 | 783.4219
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16714 | 783.4688
66 | 16719 | 783.7031
66 | 16710 | 783.2812
66 | 16722 | 783.8438
66 | 16711 | 783.3281
66 | 16723 | 783.8906
66 | 16712 | 783.375
66 | 16723 | 783.8906
66 | 16711 | 783.3281
66 | 16724 | 783.9375
66 | 16711 | 783.3281
66 | 16722 | 783.8438
66 | 16710 | 783.2812
66 | 16721 | 783.7969
66 | 16709 | 783.2344
66 | 16722 | 783.8438
66 | 16710 | 783.2812
66 | 16722 | 783.8438
66 | 16711 | 783.3281
66 | 16723 | 783.8906
66 | 16712 | 783.375
66 | 16722 | 783.8438
66 | 16713 | 783.4219
66 | 16724 | 783.9375
66 | 16712 | 783.375
66 | 16724 | 783.9375
66 | 16712 | 783.375
66 | 16725 | 783.9844
66 | 16712 | 783.375
66 | 16721 | 783.7969
66 | 16710 | 783.2812
66 | 16720 | 783.75
66 | 16711 | 783.3281
66 | 16722 | 783.8438
66 | 16712 | 783.375
66 | 16722 | 783.8438
66 | 16714 | 783.4688
66 | 16721 | 783.7969
66 | 16713 | 783.4219
66 | 16723 | 783.8906
66 | 16715 | 783.5156
66 | 16721 | 783.7969
66 | 16717 | 783.6094
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16717 | 783.6094
66 | 16719 | 783.7031
66 | 16716 | 783.5625
66 | 16721 | 783.7969
66 | 16713 | 783.4219
66 | 16720 | 783.75
66 | 16713 | 783.4219
66 | 16723 | 783.8906
66 | 16712 | 783.375
66 | 16724 | 783.9375
66 | 16715 | 783.5156
66 | 16725 | 783.9844
66 | 16717 | 783.6094
66 | 16724 | 783.9375
66 | 16717 | 783.6094
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16718 | 783.6562
66 | 16721 | 783.7969
66 | 16715 | 783.5156
66 | 16724 | 783.9375
66 | 16712 | 783.375
66 | 16725 | 783.9844
66 | 16716 | 783.5625
66 | 16721 | 783.7969
66 | 16725 | 783.9844
66 | 16716 | 783.5625
66 | 16727 | 784.0781
66 | 16716 | 783.5625
66 | 16726 | 784.0312
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16717 | 783.6094
66 | 16726 | 784.0312
66 | 16714 | 783.4688
66 | 16722 | 783.8438
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16724 | 783.9375
66 | 16716 | 783.5625
66 | 16727 | 784.0781
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16727 | 784.0781
66 | 16716 | 783.5625
66 | 16725 | 783.9844
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16726 | 784.0312
66 | 16717 | 783.6094
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16715 | 783.5156
66 | 16726 | 784.0312
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16725 | 783.9844
66 | 16720 | 783.75
66 | 16728 | 784.125
66 | 16720 | 783.75
66 | 16723 | 783.8906
66 | 16726 | 784.0312
66 | 16719 | 783.7031
66 | 16728 | 784.125
66 | 16720 | 783.75
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16717 | 783.6094
66 | 16728 | 784.125
66 | 16718 | 783.6562
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16719 | 783.7031
66 | 16729 | 784.1719
66 | 16725 | 783.9844
66 | 16722 | 783.8438
66 | 16729 | 784.1719
66 | 16718 | 783.6562
66 | 16724 | 783.9375
66 | 16724 | 783.9375
66 | 16717 | 783.6094
66 | 16728 | 784.125
66 | 16723 | 783.8906
66 | 16717 | 783.6094
66 | 16727 | 784.0781
66 | 16722 | 783.8438
66 | 16719 | 783.7031
66 | 16728 | 784.125
66 | 16721 | 783.7969
66 | 16718 | 783.6562
66 | 16728 | 784.125
66 | 16726 | 784.0312
66 | 16721 | 783.7969
66 | 16729 | 784.1719
66 | 16725 | 783.9844
66 | 16719 | 783.7031
66 | 16727 | 784.0781
66 | 16726 | 784.0312
66 | 16717 | 783.6094
66 | 16725 | 783.9844
66 | 16727 | 784.0781
66 | 16716 | 783.5625
66 | 16724 | 783.9375
66 | 16728 | 784.125
66 | 16720 | 783.75
66 | 16725 | 783.9844
66 | 16727 | 784.0781
66 | 16721 | 783.7969
66 | 16726 | 784.0312
66 | 16731 | 784.2656
66 | 16727 | 784.0781
66 | 16719 | 783.7031
66 | 16726 | 784.0312
66 | 16728 | 784.125
66 | 16718 | 783.6562
66 | 16724 | 783.9375
66 | 16728 | 784.125
66 | 16719 | 783.7031
66 | 16723 | 783.8906
66 | 16729 | 784.1719
66 | 16726 | 784.0312
66 | 16720 | 783.75
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16720 | 783.75
66 | 16727 | 784.0781
66 | 16731 | 784.2656
66 | 16722 | 783.8438
66 | 16727 | 784.0781
66 | 16729 | 784.1719
66 | 16719 | 783.7031
66 | 16725 | 783.9844
66 | 16727 | 784.0781
66 | 16720 | 783.75
66 | 16727 | 784.0781
66 | 16728 | 784.125
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16731 | 784.2656
66 | 16725 | 783.9844
66 | 16722 | 783.8438
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16721 | 783.7969
66 | 16728 | 784.125
66 | 16732 | 784.3125
66 | 16723 | 783.8906
66 | 16721 | 783.7969
66 | 16728 | 784.125
66 | 16728 | 784.125
66 | 16718 | 783.6562
66 | 16724 | 783.9375
66 | 16730 | 784.2188
66 | 16726 | 784.0312
66 | 16721 | 783.7969
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16722 | 783.8438
66 | 16728 | 784.125
66 | 16733 | 784.3594
66 | 16731 | 784.2656
66 | 16722 | 783.8438
66 | 16728 | 784.125
66 | 16732 | 784.3125
66 | 16728 | 784.125
66 | 16721 | 783.7969
66 | 16726 | 784.0312
66 | 16731 | 784.2656
66 | 16729 | 784.1719
66 | 16722 | 783.8438
66 | 16727 | 784.0781
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16723 | 783.8906
66 | 16728 | 784.125
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16728 | 784.125
66 | 16724 | 783.9375
66 | 16730 | 784.2188
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16724 | 783.9375
66 | 16726 | 784.0312
66 | 16730 | 784.2188
66 | 16732 | 784.3125
66 | 16728 | 784.125
66 | 16722 | 783.8438
66 | 16728 | 784.125
66 | 16731 | 784.2656
66 | 16734 | 784.4062
66 | 16731 | 784.2656
66 | 16724 | 783.9375
66 | 16730 | 784.2188
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16732 | 784.3125
66 | 16726 | 784.0312
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16729 | 784.1719
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16728 | 784.125
66 | 16726 | 784.0312
66 | 16731 | 784.2656
66 | 16734 | 784.4062
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16729 | 784.1719
66 | 16725 | 783.9844
66 | 16730 | 784.2188
66 | 16733 | 784.3594
66 | 16735 | 784.4531
66 | 16733 | 784.3594
66 | 16728 | 784.125
66 | 16724 | 783.9375
66 | 16727 | 784.0781
66 | 16731 | 784.2656
66 | 16735 | 784.4531
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16727 | 784.0781
66 | 16725 | 783.9844
66 | 16732 | 784.3125
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16735 | 784.4531
66 | 16733 | 784.3594
66 | 16727 | 784.0781
66 | 16725 | 783.9844
66 | 16730 | 784.2188
66 | 16733 | 784.3594
66 | 16735 | 784.4531
66 | 16733 | 784.3594
66 | 16728 | 784.125
66 | 16731 | 784.2656
66 | 16725 | 783.9844
66 | 16730 | 784.2188
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16734 | 784.4062
66 | 16727 | 784.0781
66 | 16728 | 784.125
66 | 16732 | 784.3125
66 | 16734 | 784.4062
66 | 16738 | 784.5938
66 | 16736 | 784.5
66 | 16733 | 784.3594
66 | 16727 | 784.0781
66 | 16724 | 783.9375
66 | 16728 | 784.125
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16728 | 784.125
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16737 | 784.5469
66 | 16738 | 784.5938
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16730 | 784.2188
66 | 16727 | 784.0781
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16729 | 784.1719
66 | 16730 | 784.2188
66 | 16735 | 784.4531
66 | 16734 | 784.4062
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16726 | 784.0312
66 | 16726 | 784.0312
66 | 16729 | 784.1719
66 | 16728 | 784.125
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16735 | 784.4531
66 | 16731 | 784.2656
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16732 | 784.3125
66 | 16734 | 784.4062
66 | 16736 | 784.5
66 | 16737 | 784.5469
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16741 | 784.7344
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16739 | 784.6406
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16738 | 784.5938
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16738 | 784.5938
66 | 16740 | 784.6875
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16742 | 784.7812
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16742 | 784.7812
66 | 16743 | 784.8281
66 | 16742 | 784.7812
66 | 16743 | 784.8281
66 | 16744 | 784.875
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16742 | 784.7812
66 | 16744 | 784.875
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16742 | 784.7812
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16740 | 784.6875
66 | 16736 | 784.5
66 | 16735 | 784.4531
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16735 | 784.4531
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16737 | 784.5469
66 | 16739 | 784.6406
66 | 16739 | 784.6406
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16742 | 784.7812
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16737 | 784.5469
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16737 | 784.5469
66 | 16739 | 784.6406
66 | 16743 | 784.8281
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16737 | 784.5469
66 | 16734 | 784.4062
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16739 | 784.6406
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16743 | 784.8281
66 | 16743 | 784.8281
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16734 | 784.4062
66 | 16736 | 784.5
66 | 16742 | 784.7812
66 | 16745 | 784.9219
66 | 16746 | 784.9688
66 | 16744 | 784.875
66 | 16744 | 784.875
66 | 16742 | 784.7812
66 | 16737 | 784.5469
66 | 16733 | 784.3594
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16740 | 784.6875
66 | 16742 | 784.7812
66 | 16745 | 784.9219
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16737 | 784.5469
66 | 16734 | 784.4062
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16745 | 784.9219
66 | 16746 | 784.9688
66 | 16745 | 784.9219
66 | 16742 | 784.7812
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16738 | 784.5938
66 | 16742 | 784.7812
66 | 16744 | 784.875
66 | 16743 | 784.8281
66 | 16740 | 784.6875
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16739 | 784.6406
66 | 16736 | 784.5
66 | 16741 | 784.7344
66 | 16746 | 784.9688
66 | 16745 | 784.9219
66 | 16742 | 784.7812
66 | 16738 | 784.5938
66 | 16733 | 784.3594
66 | 16737 | 784.5469
66 | 16742 | 784.7812
66 | 16746 | 784.9688
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16744 | 784.875
66 | 16748 | 785.0625
66 | 16745 | 784.9219
66 | 16745 | 784.9219
66 | 16737 | 784.5469
66 | 16740 | 784.6875
66 | 16745 | 784.9219
66 | 16746 | 784.9688
66 | 16741 | 784.7344
66 | 16734 | 784.4062
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16747 | 785.0156
66 | 16742 | 784.7812
66 | 16737 | 784.5469
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16748 | 785.0625
66 | 16743 | 784.8281
66 | 16736 | 784.5
66 | 16739 | 784.6406
66 | 16745 | 784.9219
66 | 16747 | 785.0156
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16740 | 784.6875
66 | 16745 | 784.9219
66 | 16746 | 784.9688
66 | 16742 | 784.7812
66 | 16736 | 784.5
66 | 16740 | 784.6875
66 | 16746 | 784.9688
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16739 | 784.6406
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16748 | 785.0625
66 | 16743 | 784.8281
66 | 16737 | 784.5469
66 | 16741 | 784.7344
66 | 16746 | 784.9688
66 | 16745 | 784.9219
66 | 16741 | 784.7344
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16747 | 785.0156
66 | 16743 | 784.8281
66 | 16739 | 784.6406
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16748 | 785.0625
66 | 16745 | 784.9219
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16747 | 785.0156
66 | 16747 | 785.0156
66 | 16740 | 784.6875
66 | 16738 | 784.5938
66 | 16744 | 784.875
66 | 16747 | 785.0156
66 | 16744 | 784.875
66 | 16737 | 784.5469
66 | 16742 | 784.7812
66 | 16749 | 785.1094
66 | 16748 | 785.0625
66 | 16745 | 784.9219
66 | 16741 | 784.7344
66 | 16749 | 785.1094
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16737 | 784.5469
66 | 16741 | 784.7344
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16748 | 785.0625
66 | 16747 | 785.0156
66 | 16743 | 784.8281
66 | 16740 | 784.6875
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16746 | 784.9688
66 | 16738 | 784.5938
66 | 16744 | 784.875
66 | 16748 | 785.0625
66 | 16742 | 784.7812
66 | 16738 | 784.5938
66 | 16745 | 784.9219
66 | 16745 | 784.9219
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16747 | 785.0156
66 | 16747 | 785.0156
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16749 | 785.1094
66 | 16749 | 785.1094
66 | 16745 | 784.9219
66 | 16744 | 784.875
66 | 16748 | 785.0625
66 | 16744 | 784.875
66 | 16738 | 784.5938
66 | 16745 | 784.9219
66 | 16747 | 785.0156
66 | 16741 | 784.7344
66 | 16739 | 784.6406
66 | 16747 | 785.0156
66 | 16745 | 784.9219
66 | 16738 | 784.5938
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16746 | 784.9688
66 | 16742 | 784.7812
66 | 16750 | 785.1562
66 | 16745 | 784.9219
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16747 | 785.0156
66 | 16740 | 784.6875
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16746 | 784.9688
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16750 | 785.1562
66 | 16743 | 784.8281
66 | 16745 | 784.9219
66 | 16753 | 785.2969
66 | 16745 | 784.9219
66 | 16745 | 784.9219
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16743 | 784.8281
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16740 | 784.6875
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16740 | 784.6875
66 | 16751 | 785.2031
66 | 16749 | 785.1094
66 | 16741 | 784.7344
66 | 16749 | 785.1094
66 | 16752 | 785.25
66 | 16745 | 784.9219
66 | 16745 | 784.9219
66 | 16750 | 785.1562
66 | 16744 | 784.875
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16743 | 784.8281
66 | 16745 | 784.9219
66 | 16750 | 785.1562
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16751 | 785.2031
66 | 16742 | 784.7812
66 | 16750 | 785.1562
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16748 | 785.0625
66 | 16746 | 784.9688
66 | 16752 | 785.25
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16748 | 785.0625
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16748 | 785.0625
66 | 16740 | 784.6875
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16742 | 784.7812
66 | 16744 | 784.875
66 | 16752 | 785.25
66 | 16742 | 784.7812
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16741 | 784.7344
66 | 16750 | 785.1562
66 | 16745 | 784.9219
66 | 16743 | 784.8281
66 | 16751 | 785.2031
66 | 16743 | 784.8281
66 | 16747 | 785.0156
66 | 16749 | 785.1094
66 | 16742 | 784.7812
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16742 | 784.7812
66 | 16753 | 785.2969
66 | 16746 | 784.9688
66 | 16752 | 785.25
66 | 16748 | 785.0625
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16740 | 784.6875
66 | 16750 | 785.1562
66 | 16746 | 784.9688
66 | 16745 | 784.9219
66 | 16748 | 785.0625
66 | 16743 | 784.8281
66 | 16753 | 785.2969
66 | 16744 | 784.875
66 | 16751 | 785.2031
66 | 16751 | 785.2031
66 | 16749 | 785.1094
66 | 16752 | 785.25
66 | 16747 | 785.0156
66 | 16756 | 785.4375
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16748 | 785.0625
66 | 16744 | 784.875
66 | 16749 | 785.1094
66 | 16746 | 784.9688
66 | 16749 | 785.1094
66 | 16741 | 784.7344
66 | 16754 | 785.3438
66 | 16745 | 784.9219
66 | 16751 | 785.2031
66 | 16749 | 785.1094
66 | 16749 | 785.1094
66 | 16752 | 785.25
66 | 16745 | 784.9219
66 | 16755 | 785.3906
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16748 | 785.0625
66 | 16750 | 785.1562
66 | 16747 | 785.0156
66 | 16745 | 784.9219
66 | 16748 | 785.0625
66 | 16747 | 785.0156
66 | 16747 | 785.0156
66 | 16751 | 785.2031
66 | 16749 | 785.1094
66 | 16752 | 785.25
66 | 16750 | 785.1562
66 | 16754 | 785.3438
66 | 16750 | 785.1562
66 | 16751 | 785.2031
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16749 | 785.1094
66 | 16748 | 785.0625
66 | 16748 | 785.0625
66 | 16745 | 784.9219
66 | 16750 | 785.1562
66 | 16746 | 784.9688
66 | 16752 | 785.25
66 | 16748 | 785.0625
66 | 16751 | 785.2031
66 | 16748 | 785.0625
66 | 16753 | 785.2969
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16745 | 784.9219
66 | 16753 | 785.2969
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16746 | 784.9688
66 | 16752 | 785.25
66 | 16746 | 784.9688
66 | 16754 | 785.3438
66 | 16745 | 784.9219
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16748 | 785.0625
66 | 16756 | 785.4375
66 | 16749 | 785.1094
66 | 16756 | 785.4375
66 | 16749 | 785.1094
66 | 16754 | 785.3438
66 | 16746 | 784.9688
66 | 16752 | 785.25
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16746 | 784.9688
66 | 16754 | 785.3438
66 | 16748 | 785.0625
66 | 16755 | 785.3906
66 | 16746 | 784.9688
66 | 16757 | 785.4844
66 | 16748 | 785.0625
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16756 | 785.4375
66 | 16746 | 784.9688
66 | 16756 | 785.4375
66 | 16746 | 784.9688
66 | 16755 | 785.3906
66 | 16745 | 784.9219
66 | 16755 | 785.3906
66 | 16744 | 784.875
66 | 16755 | 785.3906
66 | 16745 | 784.9219
66 | 16757 | 785.4844
66 | 16745 | 784.9219
66 | 16756 | 785.4375
66 | 16748 | 785.0625
66 | 16757 | 785.4844
66 | 16747 | 785.0156
66 | 16757 | 785.4844
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16746 | 784.9688
66 | 16755 | 785.3906
66 | 16746 | 784.9688
66 | 16756 | 785.4375
66 | 16746 | 784.9688
66 | 16755 | 785.3906
66 | 16745 | 784.9219
66 | 16755 | 785.3906
66 | 16747 | 785.0156
66 | 16757 | 785.4844
66 | 16747 | 785.0156
66 | 16758 | 785.5312
66 | 16746 | 784.9688
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16759 | 785.5781
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16745 | 784.9219
66 | 16756 | 785.4375
66 | 16743 | 784.8281
66 | 16756 | 785.4375
66 | 16745 | 784.9219
66 | 16756 | 785.4375
66 | 16746 | 784.9688
66 | 16758 | 785.5312
66 | 16745 | 784.9219
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16760 | 785.625
66 | 16749 | 785.1094
66 | 16757 | 785.4844
66 | 16748 | 785.0625
66 | 16757 | 785.4844
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16748 | 785.0625
66 | 16757 | 785.4844
66 | 16749 | 785.1094
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16757 | 785.4844
66 | 16748 | 785.0625
66 | 16757 | 785.4844
66 | 16745 | 784.9219
66 | 16757 | 785.4844
66 | 16746 | 784.9688
66 | 16756 | 785.4375
66 | 16746 | 784.9688
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16759 | 785.5781
66 | 16751 | 785.2031
66 | 16755 | 785.3906
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16756 | 785.4375
66 | 16748 | 785.0625
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16746 | 784.9688
66 | 16758 | 785.5312
66 | 16750 | 785.1562
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16750 | 785.1562
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16748 | 785.0625
66 | 16759 | 785.5781
66 | 16747 | 785.0156
66 | 16757 | 785.4844
66 | 16750 | 785.1562
66 | 16750 | 785.1562
66 | 16756 | 785.4375
66 | 16747 | 785.0156
66 | 16758 | 785.5312
66 | 16748 | 785.0625
66 | 16759 | 785.5781
66 | 16752 | 785.25
66 | 16753 | 785.2969
66 | 16762 | 785.7188
66 | 16750 | 785.1562
66 | 16758 | 785.5312
66 | 16759 | 785.5781
66 | 16750 | 785.1562
66 | 16760 | 785.625
66 | 16748 | 785.0625
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16748 | 785.0625
66 | 16759 | 785.5781
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16759 | 785.5781
66 | 16749 | 785.1094
66 | 16759 | 785.5781
66 | 16759 | 785.5781
66 | 16751 | 785.2031
66 | 16760 | 785.625
66 | 16754 | 785.3438
66 | 16751 | 785.2031
66 | 16760 | 785.625
66 | 16751 | 785.2031
66 | 16755 | 785.3906
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16757 | 785.4844
66 | 16751 | 785.2031
66 | 16750 | 785.1562
66 | 16759 | 785.5781
66 | 16752 | 785.25
66 | 16754 | 785.3438
66 | 16762 | 785.7188
66 | 16749 | 785.1094
66 | 16760 | 785.625
66 | 16760 | 785.625
66 | 16750 | 785.1562
66 | 16760 | 785.625
66 | 16753 | 785.2969
66 | 16754 | 785.3438
66 | 16759 | 785.5781
66 | 16749 | 785.1094
66 | 16752 | 785.25
66 | 16758 | 785.5312
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16759 | 785.5781
66 | 16749 | 785.1094
66 | 16759 | 785.5781
66 | 16759 | 785.5781
66 | 16750 | 785.1562
66 | 16760 | 785.625
66 | 16753 | 785.2969
66 | 16753 | 785.2969
66 | 16760 | 785.625
66 | 16750 | 785.1562
66 | 16756 | 785.4375
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16747 | 785.0156
66 | 16758 | 785.5312
66 | 16749 | 785.1094
66 | 16750 | 785.1562
66 | 16761 | 785.6719
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16762 | 785.7188
66 | 16750 | 785.1562
66 | 16760 | 785.625
66 | 16753 | 785.2969
66 | 16752 | 785.25
66 | 16760 | 785.625
66 | 16751 | 785.2031
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16750 | 785.1562
66 | 16756 | 785.4375
66 | 16759 | 785.5781
66 | 16748 | 785.0625
66 | 16758 | 785.5312
66 | 16760 | 785.625
66 | 16751 | 785.2031
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16751 | 785.2031
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16752 | 785.25
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16751 | 785.2031
66 | 16761 | 785.6719
66 | 16753 | 785.2969
66 | 16751 | 785.2031
66 | 16762 | 785.7188
66 | 16752 | 785.25
66 | 16752 | 785.25
66 | 16763 | 785.7656
66 | 16754 | 785.3438
66 | 16757 | 785.4844
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16753 | 785.2969
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16754 | 785.3438
66 | 16762 | 785.7188
66 | 16753 | 785.2969
66 | 16750 | 785.1562
66 | 16760 | 785.625
66 | 16753 | 785.2969
66 | 16752 | 785.25
66 | 16762 | 785.7188
66 | 16751 | 785.2031
66 | 16749 | 785.1094
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16751 | 785.2031
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16751 | 785.2031
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16760 | 785.625
66 | 16751 | 785.2031
66 | 16753 | 785.2969
66 | 16760 | 785.625
66 | 16751 | 785.2031
66 | 16752 | 785.25
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16750 | 785.1562
66 | 16761 | 785.6719
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16761 | 785.6719
66 | 16753 | 785.2969
66 | 16750 | 785.1562
66 | 16759 | 785.5781
66 | 16761 | 785.6719
66 | 16752 | 785.25
66 | 16754 | 785.3438
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16752 | 785.25
66 | 16758 | 785.5312
66 | 16763 | 785.7656
66 | 16756 | 785.4375
66 | 16751 | 785.2031
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16751 | 785.2031
66 | 16761 | 785.6719
66 | 16760 | 785.625
66 | 16753 | 785.2969
66 | 16750 | 785.1562
66 | 16759 | 785.5781
66 | 16762 | 785.7188
66 | 16756 | 785.4375
66 | 16751 | 785.2031
66 | 16758 | 785.5312
66 | 16764 | 785.8125
66 | 16756 | 785.4375
66 | 16752 | 785.25
66 | 16760 | 785.625
66 | 16762 | 785.7188
66 | 16756 | 785.4375
66 | 16753 | 785.2969
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16755 | 785.3906
66 | 16750 | 785.1562
66 | 16756 | 785.4375
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16752 | 785.25
66 | 16757 | 785.4844
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16753 | 785.2969
66 | 16759 | 785.5781
66 | 16765 | 785.8594
66 | 16758 | 785.5312
66 | 16753 | 785.2969
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16753 | 785.2969
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16751 | 785.2031
66 | 16759 | 785.5781
66 | 16763 | 785.7656
66 | 16758 | 785.5312
66 | 16754 | 785.3438
66 | 16755 | 785.3906
66 | 16762 | 785.7188
66 | 16763 | 785.7656
66 | 16755 | 785.3906
66 | 16752 | 785.25
66 | 16758 | 785.5312
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16761 | 785.6719
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16752 | 785.25
66 | 16760 | 785.625
66 | 16766 | 785.9062
66 | 16761 | 785.6719
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16763 | 785.7656
66 | 16766 | 785.9062
66 | 16764 | 785.8125
66 | 16756 | 785.4375
66 | 16752 | 785.25
66 | 16759 | 785.5781
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16754 | 785.3438
66 | 16752 | 785.25
66 | 16758 | 785.5312
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16753 | 785.2969
66 | 16759 | 785.5781
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16756 | 785.4375
66 | 16754 | 785.3438
66 | 16758 | 785.5312
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16757 | 785.4844
66 | 16755 | 785.3906
66 | 16752 | 785.25
66 | 16759 | 785.5781
66 | 16765 | 785.8594
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16759 | 785.5781
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16761 | 785.6719
66 | 16765 | 785.8594
66 | 16761 | 785.6719
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16759 | 785.5781
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16755 | 785.3906
66 | 16759 | 785.5781
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16757 | 785.4844
66 | 16754 | 785.3438
66 | 16754 | 785.3438
66 | 16760 | 785.625
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16764 | 785.8125
66 | 16759 | 785.5781
66 | 16758 | 785.5312
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16760 | 785.625
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16763 | 785.7656
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16764 | 785.8125
66 | 16768 | 786
66 | 16768 | 786
66 | 16763 | 785.7656
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16759 | 785.5781
66 | 16763 | 785.7656
66 | 16766 | 785.9062
66 | 16765 | 785.8594
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16762 | 785.7188
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16760 | 785.625
66 | 16758 | 785.5312
66 | 16759 | 785.5781
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16762 | 785.7188
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16762 | 785.7188
66 | 16760 | 785.625
66 | 16759 | 785.5781
66 | 16758 | 785.5312
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16765 | 785.8594
66 | 16761 | 785.6719
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16758 | 785.5312
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16761 | 785.6719
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16758 | 785.5312
66 | 16758 | 785.5312
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16760 | 785.625
66 | 16760 | 785.625
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16760 | 785.625
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16762 | 785.7188
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16759 | 785.5781
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16761 | 785.6719
66 | 16763 | 785.7656
66 | 16760 | 785.625
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16763 | 785.7656
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16766 | 785.9062
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16768 | 786
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16768 | 786
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16765 | 785.8594
66 | 16762 | 785.7188
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16770 | 786.0938
66 | 16763 | 785.7656
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16766 | 785.9062
66 | 16762 | 785.7188
66 | 16766 | 785.9062
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16763 | 785.7656
66 | 16768 | 786
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16770 | 786.0938
66 | 16768 | 786
66 | 16768 | 786
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16769 | 786.0469
66 | 16768 | 786
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16768 | 786
66 | 16770 | 786.0938
66 | 16766 | 785.9062
66 | 16772 | 786.1875
66 | 16768 | 786
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16768 | 786
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16768 | 786
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16775 | 786.3281
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16776 | 786.375
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16771 | 786.1406
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16773 | 786.2344
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16771 | 786.1406
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16776 | 786.375
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16775 | 786.3281
66 | 16778 | 786.4688
66 | 16775 | 786.3281
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16775 | 786.3281
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16780 | 786.5625
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16782 | 786.6562
66 | 16782 | 786.6562
66 | 16782 | 786.6562
66 | 16782 | 786.6562
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16781 | 786.6094
66 | 16782 | 786.6562
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16780 | 786.5625
66 | 16782 | 786.6562
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16781 | 786.6094
66 | 16780 | 786.5625
66 | 16779 | 786.5156
66 | 16780 | 786.5625
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16768 | 786
66 | 16770 | 786.0938
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16776 | 786.375
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16772 | 786.1875
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16777 | 786.4219
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16773 | 786.2344
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16770 | 786.0938
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16768 | 786
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16773 | 786.2344
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16779 | 786.5156
66 | 16779 | 786.5156
66 | 16775 | 786.3281
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16767 | 785.9531
66 | 16773 | 786.2344
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16770 | 786.0938
66 | 16766 | 785.9062
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16774 | 786.2812
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16770 | 786.0938
66 | 16767 | 785.9531
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16765 | 785.8594
66 | 16766 | 785.9062
66 | 16769 | 786.0469
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16774 | 786.2812
66 | 16768 | 786
66 | 16765 | 785.8594
66 | 16764 | 785.8125
66 | 16770 | 786.0938
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16776 | 786.375
66 | 16769 | 786.0469
66 | 16766 | 785.9062
66 | 16769 | 786.0469
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16769 | 786.0469
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16771 | 786.1406
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16778 | 786.4688
66 | 16774 | 786.2812
66 | 16768 | 786
66 | 16767 | 785.9531
66 | 16769 | 786.0469
66 | 16777 | 786.4219
66 | 16778 | 786.4688
66 | 16778 | 786.4688
66 | 16775 | 786.3281
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16769 | 786.0469
66 | 16764 | 785.8125
66 | 16768 | 786
66 | 16775 | 786.3281
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16773 | 786.2344
66 | 16766 | 785.9062
66 | 16766 | 785.9062
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16770 | 786.0938
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16777 | 786.4219
66 | 16776 | 786.375
66 | 16769 | 786.0469
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16770 | 786.0938
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16770 | 786.0938
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16776 | 786.375
66 | 16777 | 786.4219
66 | 16777 | 786.4219
66 | 16774 | 786.2812
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16774 | 786.2812
66 | 16770 | 786.0938
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16774 | 786.2812
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16771 | 786.1406
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16770 | 786.0938
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16772 | 786.1875
66 | 16765 | 785.8594
66 | 16765 | 785.8594
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16769 | 786.0469
66 | 16762 | 785.7188
66 | 16764 | 785.8125
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16769 | 786.0469
66 | 16764 | 785.8125
66 | 16768 | 786
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16775 | 786.3281
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16770 | 786.0938
66 | 16762 | 785.7188
66 | 16762 | 785.7188
66 | 16772 | 786.1875
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16771 | 786.1406
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16776 | 786.375
66 | 16776 | 786.375
66 | 16771 | 786.1406
66 | 16763 | 785.7656
66 | 16766 | 785.9062
66 | 16774 | 786.2812
66 | 16774 | 786.2812
66 | 16767 | 785.9531
66 | 16761 | 785.6719
66 | 16767 | 785.9531
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16767 | 785.9531
66 | 16775 | 786.3281
66 | 16776 | 786.375
66 | 16770 | 786.0938
66 | 16764 | 785.8125
66 | 16768 | 786
66 | 16775 | 786.3281
66 | 16774 | 786.2812
66 | 16768 | 786
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16773 | 786.2344
66 | 16776 | 786.375
66 | 16772 | 786.1875
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16777 | 786.4219
66 | 16774 | 786.2812
66 | 16768 | 786
66 | 16762 | 785.7188
66 | 16770 | 786.0938
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16769 | 786.0469
66 | 16761 | 785.6719
66 | 16763 | 785.7656
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16768 | 786
66 | 16763 | 785.7656
66 | 16768 | 786
66 | 16774 | 786.2812
66 | 16775 | 786.3281
66 | 16771 | 786.1406
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16772 | 786.1875
66 | 16772 | 786.1875
66 | 16768 | 786
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16770 | 786.0938
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16770 | 786.0938
66 | 16763 | 785.7656
66 | 16763 | 785.7656
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16770 | 786.0938
66 | 16764 | 785.8125
66 | 16763 | 785.7656
66 | 16774 | 786.2812
66 | 16773 | 786.2344
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16770 | 786.0938
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16769 | 786.0469
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16772 | 786.1875
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16773 | 786.2344
66 | 16773 | 786.2344
66 | 16772 | 786.1875
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16765 | 785.8594
66 | 16758 | 785.5312
66 | 16763 | 785.7656
66 | 16771 | 786.1406
66 | 16772 | 786.1875
66 | 16765 | 785.8594
66 | 16763 | 785.7656
66 | 16771 | 786.1406
66 | 16774 | 786.2812
66 | 16771 | 786.1406
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16769 | 786.0469
66 | 16763 | 785.7656
66 | 16759 | 785.5781
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16771 | 786.1406
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16771 | 786.1406
66 | 16773 | 786.2344
66 | 16768 | 786
66 | 16761 | 785.6719
66 | 16765 | 785.8594
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16765 | 785.8594
66 | 16760 | 785.625
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16769 | 786.0469
66 | 16760 | 785.625
66 | 16762 | 785.7188
66 | 16772 | 786.1875
66 | 16770 | 786.0938
66 | 16762 | 785.7188
66 | 16762 | 785.7188
66 | 16771 | 786.1406
66 | 16771 | 786.1406
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16763 | 785.7656
66 | 16760 | 785.625
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16764 | 785.8125
66 | 16760 | 785.625
66 | 16769 | 786.0469
66 | 16772 | 786.1875
66 | 16771 | 786.1406
66 | 16762 | 785.7188
66 | 16764 | 785.8125
66 | 16773 | 786.2344
66 | 16770 | 786.0938
66 | 16761 | 785.6719
66 | 16763 | 785.7656
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16761 | 785.6719
66 | 16770 | 786.0938
66 | 16766 | 785.9062
66 | 16759 | 785.5781
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16772 | 786.1875
66 | 16770 | 786.0938
66 | 16761 | 785.6719
66 | 16761 | 785.6719
66 | 16770 | 786.0938
66 | 16768 | 786
66 | 16757 | 785.4844
66 | 16764 | 785.8125
66 | 16769 | 786.0469
66 | 16766 | 785.9062
66 | 16759 | 785.5781
66 | 16768 | 786
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16770 | 786.0938
66 | 16772 | 786.1875
66 | 16766 | 785.9062
66 | 16760 | 785.625
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16760 | 785.625
66 | 16759 | 785.5781
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16757 | 785.4844
66 | 16762 | 785.7188
66 | 16769 | 786.0469
66 | 16769 | 786.0469
66 | 16758 | 785.5312
66 | 16764 | 785.8125
66 | 16771 | 786.1406
66 | 16767 | 785.9531
66 | 16761 | 785.6719
66 | 16771 | 786.1406
66 | 16770 | 786.0938
66 | 16760 | 785.625
66 | 16759 | 785.5781
66 | 16771 | 786.1406
66 | 16768 | 786
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16769 | 786.0469
66 | 16767 | 785.9531
66 | 16758 | 785.5312
66 | 16764 | 785.8125
66 | 16770 | 786.0938
66 | 16768 | 786
66 | 16758 | 785.5312
66 | 16769 | 786.0469
66 | 16771 | 786.1406
66 | 16762 | 785.7188
66 | 16760 | 785.625
66 | 16770 | 786.0938
66 | 16768 | 786
66 | 16757 | 785.4844
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16760 | 785.625
66 | 16770 | 786.0938
66 | 16766 | 785.9062
66 | 16761 | 785.6719
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16766 | 785.9062
66 | 16762 | 785.7188
66 | 16769 | 786.0469
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16768 | 786
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16767 | 785.9531
66 | 16768 | 786
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16769 | 786.0469
66 | 16766 | 785.9062
66 | 16762 | 785.7188
66 | 16769 | 786.0469
66 | 16770 | 786.0938
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16768 | 786
66 | 16763 | 785.7656
66 | 16757 | 785.4844
66 | 16766 | 785.9062
66 | 16767 | 785.9531
66 | 16757 | 785.4844
66 | 16762 | 785.7188
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16770 | 786.0938
66 | 16764 | 785.8125
66 | 16762 | 785.7188
66 | 16771 | 786.1406
66 | 16767 | 785.9531
66 | 16762 | 785.7188
66 | 16769 | 786.0469
66 | 16765 | 785.8594
66 | 16759 | 785.5781
66 | 16767 | 785.9531
66 | 16765 | 785.8594
66 | 16757 | 785.4844
66 | 16763 | 785.7656
66 | 16769 | 786.0469
66 | 16763 | 785.7656
66 | 16762 | 785.7188
66 | 16769 | 786.0469
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16769 | 786.0469
66 | 16766 | 785.9062
66 | 16759 | 785.5781
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16767 | 785.9531
66 | 16760 | 785.625
66 | 16759 | 785.5781
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16757 | 785.4844
66 | 16766 | 785.9062
66 | 16768 | 786
66 | 16765 | 785.8594
66 | 16761 | 785.6719
66 | 16769 | 786.0469
66 | 16765 | 785.8594
66 | 16759 | 785.5781
66 | 16768 | 786
66 | 16763 | 785.7656
66 | 16760 | 785.625
66 | 16766 | 785.9062
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16766 | 785.9062
66 | 16761 | 785.6719
66 | 16763 | 785.7656
66 | 16768 | 786
66 | 16762 | 785.7188
66 | 16760 | 785.625
66 | 16769 | 786.0469
66 | 16761 | 785.6719
66 | 16762 | 785.7188
66 | 16768 | 786
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16767 | 785.9531
66 | 16764 | 785.8125
66 | 16758 | 785.5312
66 | 16766 | 785.9062
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16766 | 785.9062
66 | 16756 | 785.4375
66 | 16763 | 785.7656
66 | 16768 | 786
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16766 | 785.9062
66 | 16758 | 785.5312
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16756 | 785.4375
66 | 16760 | 785.625
66 | 16764 | 785.8125
66 | 16756 | 785.4375
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16756 | 785.4375
66 | 16764 | 785.8125
66 | 16767 | 785.9531
66 | 16757 | 785.4844
66 | 16765 | 785.8594
66 | 16769 | 786.0469
66 | 16760 | 785.625
66 | 16768 | 786
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16764 | 785.8125
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16765 | 785.8594
66 | 16754 | 785.3438
66 | 16761 | 785.6719
66 | 16766 | 785.9062
66 | 16755 | 785.3906
66 | 16765 | 785.8594
66 | 16767 | 785.9531
66 | 16756 | 785.4375
66 | 16764 | 785.8125
66 | 16766 | 785.9062
66 | 16758 | 785.5312
66 | 16763 | 785.7656
66 | 16765 | 785.8594
66 | 16754 | 785.3438
66 | 16760 | 785.625
66 | 16764 | 785.8125
66 | 16753 | 785.2969
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16753 | 785.2969
66 | 16762 | 785.7188
66 | 16765 | 785.8594
66 | 16756 | 785.4375
66 | 16767 | 785.9531
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16767 | 785.9531
66 | 16757 | 785.4844
66 | 16764 | 785.8125
66 | 16764 | 785.8125
66 | 16756 | 785.4375
66 | 16766 | 785.9062
66 | 16759 | 785.5781
66 | 16759 | 785.5781
66 | 16764 | 785.8125
66 | 16756 | 785.4375
66 | 16767 | 785.9531
66 | 16759 | 785.5781
66 | 16759 | 785.5781
66 | 16766 | 785.9062
66 | 16756 | 785.4375
66 | 16767 | 785.9531
66 | 16766 | 785.9062
66 | 16758 | 785.5312
66 | 16766 | 785.9062
66 | 16756 | 785.4375
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16755 | 785.3906
66 | 16762 | 785.7188
66 | 16762 | 785.7188
66 | 16755 | 785.3906
66 | 16765 | 785.8594
66 | 16762 | 785.7188
66 | 16758 | 785.5312
66 | 16765 | 785.8594
66 | 16755 | 785.3906
66 | 16764 | 785.8125
66 | 16765 | 785.8594
66 | 16758 | 785.5312
66 | 16766 | 785.9062
66 | 16755 | 785.3906
66 | 16762 | 785.7188
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16762 | 785.7188
66 | 16752 | 785.25
66 | 16762 | 785.7188
66 | 16764 | 785.8125
66 | 16756 | 785.4375
66 | 16765 | 785.8594
66 | 16755 | 785.3906
66 | 16765 | 785.8594
66 | 16762 | 785.7188
66 | 16759 | 785.5781
66 | 16765 | 785.8594
66 | 16756 | 785.4375
66 | 16766 | 785.9062
66 | 16758 | 785.5312
66 | 16760 | 785.625
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16764 | 785.8125
66 | 16752 | 785.25
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16754 | 785.3438
66 | 16763 | 785.7656
66 | 16764 | 785.8125
66 | 16759 | 785.5781
66 | 16766 | 785.9062
66 | 16758 | 785.5312
66 | 16766 | 785.9062
66 | 16755 | 785.3906
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16756 | 785.4375
66 | 16762 | 785.7188
66 | 16753 | 785.2969
66 | 16762 | 785.7188
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16764 | 785.8125
66 | 16755 | 785.3906
66 | 16765 | 785.8594
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16756 | 785.4375
66 | 16766 | 785.9062
66 | 16754 | 785.3438
66 | 16763 | 785.7656
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16762 | 785.7188
66 | 16753 | 785.2969
66 | 16757 | 785.4844
66 | 16762 | 785.7188
66 | 16754 | 785.3438
66 | 16765 | 785.8594
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16764 | 785.8125
66 | 16754 | 785.3438
66 | 16765 | 785.8594
66 | 16761 | 785.6719
66 | 16755 | 785.3906
66 | 16763 | 785.7656
66 | 16752 | 785.25
66 | 16763 | 785.7656
66 | 16754 | 785.3438
66 | 16759 | 785.5781
66 | 16762 | 785.7188
66 | 16758 | 785.5312
66 | 16763 | 785.7656
66 | 16755 | 785.3906
66 | 16764 | 785.8125
66 | 16753 | 785.2969
66 | 16764 | 785.8125
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16763 | 785.7656
66 | 16752 | 785.25
66 | 16763 | 785.7656
66 | 16756 | 785.4375
66 | 16759 | 785.5781
66 | 16760 | 785.625
66 | 16754 | 785.3438
66 | 16761 | 785.6719
66 | 16755 | 785.3906
66 | 16761 | 785.6719
66 | 16756 | 785.4375
66 | 16761 | 785.6719
66 | 16758 | 785.5312
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16764 | 785.8125
66 | 16757 | 785.4844
66 | 16763 | 785.7656
66 | 16754 | 785.3438
66 | 16762 | 785.7188
66 | 16753 | 785.2969
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16754 | 785.3438
66 | 16761 | 785.6719
66 | 16755 | 785.3906
66 | 16763 | 785.7656
66 | 16756 | 785.4375
66 | 16762 | 785.7188
66 | 16760 | 785.625
66 | 16763 | 785.7656
66 | 16758 | 785.5312
66 | 16762 | 785.7188
66 | 16758 | 785.5312
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16756 | 785.4375
66 | 16762 | 785.7188
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16758 | 785.5312
66 | 16762 | 785.7188
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16758 | 785.5312
66 | 16760 | 785.625
66 | 16756 | 785.4375
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16756 | 785.4375
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16759 | 785.5781
66 | 16761 | 785.6719
66 | 16758 | 785.5312
66 | 16761 | 785.6719
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16755 | 785.3906
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16759 | 785.5781
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16758 | 785.5312
66 | 16759 | 785.5781
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16762 | 785.7188
66 | 16758 | 785.5312
66 | 16763 | 785.7656
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16759 | 785.5781
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16755 | 785.3906
66 | 16759 | 785.5781
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16761 | 785.6719
66 | 16756 | 785.4375
66 | 16760 | 785.625
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16755 | 785.3906
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16758 | 785.5312
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16757 | 785.4844
66 | 16760 | 785.625
66 | 16758 | 785.5312
66 | 16759 | 785.5781
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16755 | 785.3906
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16757 | 785.4844
66 | 16759 | 785.5781
66 | 16758 | 785.5312
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16754 | 785.3438
66 | 16758 | 785.5312
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16759 | 785.5781
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16755 | 785.3906
66 | 16756 | 785.4375
66 | 16754 | 785.3438
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16754 | 785.3438
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16758 | 785.5312
66 | 16756 | 785.4375
66 | 16758 | 785.5312
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16757 | 785.4844
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16753 | 785.2969
66 | 16752 | 785.25
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16756 | 785.4375
66 | 16757 | 785.4844
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16753 | 785.2969
66 | 16754 | 785.3438
66 | 16753 | 785.2969
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16754 | 785.3438
66 | 16754 | 785.3438
66 | 16752 | 785.25
66 | 16752 | 785.25
66 | 16752 | 785.25
66 | 16752 | 785.25
66 | 16752 | 785.25
66 | 16752 | 785.25
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16752 | 785.25
66 | 16756 | 785.4375
66 | 16755 | 785.3906
66 | 16757 | 785.4844
66 | 16753 | 785.2969
66 | 16756 | 785.4375
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16755 | 785.3906
66 | 16756 | 785.4375
66 | 16751 | 785.2031
66 | 16753 | 785.2969
66 | 16751 | 785.2031
66 | 16754 | 785.3438
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16756 | 785.4375
66 | 16753 | 785.2969
66 | 16758 | 785.5312
66 | 16754 | 785.3438
66 | 16756 | 785.4375
66 | 16754 | 785.3438
66 | 16756 | 785.4375
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16750 | 785.1562
66 | 16754 | 785.3438
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16753 | 785.2969
66 | 16756 | 785.4375
66 | 16752 | 785.25
66 | 16756 | 785.4375
66 | 16752 | 785.25
66 | 16756 | 785.4375
66 | 16750 | 785.1562
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16752 | 785.25
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16750 | 785.1562
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16753 | 785.2969
66 | 16756 | 785.4375
66 | 16753 | 785.2969
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16753 | 785.2969
66 | 16749 | 785.1094
66 | 16752 | 785.25
66 | 16749 | 785.1094
66 | 16754 | 785.3438
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16750 | 785.1562
66 | 16754 | 785.3438
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16752 | 785.25
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16755 | 785.3906
66 | 16751 | 785.2031
66 | 16755 | 785.3906
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16752 | 785.25
66 | 16748 | 785.0625
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16752 | 785.25
66 | 16750 | 785.1562
66 | 16754 | 785.3438
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16751 | 785.2031
66 | 16756 | 785.4375
66 | 16750 | 785.1562
66 | 16754 | 785.3438
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16748 | 785.0625
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16747 | 785.0156
66 | 16755 | 785.3906
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16753 | 785.2969
66 | 16746 | 784.9688
66 | 16751 | 785.2031
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16753 | 785.2969
66 | 16749 | 785.1094
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16754 | 785.3438
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16746 | 784.9688
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16748 | 785.0625
66 | 16754 | 785.3438
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16752 | 785.25
66 | 16746 | 784.9688
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16753 | 785.2969
66 | 16744 | 784.875
66 | 16751 | 785.2031
66 | 16744 | 784.875
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16745 | 784.9219
66 | 16751 | 785.2031
66 | 16747 | 785.0156
66 | 16751 | 785.2031
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16747 | 785.0156
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16743 | 784.8281
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16751 | 785.2031
66 | 16746 | 784.9688
66 | 16753 | 785.2969
66 | 16745 | 784.9219
66 | 16753 | 785.2969
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16744 | 784.875
66 | 16751 | 785.2031
66 | 16742 | 784.7812
66 | 16750 | 785.1562
66 | 16742 | 784.7812
66 | 16750 | 785.1562
66 | 16743 | 784.8281
66 | 16750 | 785.1562
66 | 16744 | 784.875
66 | 16752 | 785.25
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16746 | 784.9688
66 | 16752 | 785.25
66 | 16745 | 784.9219
66 | 16751 | 785.2031
66 | 16743 | 784.8281
66 | 16751 | 785.2031
66 | 16743 | 784.8281
66 | 16748 | 785.0625
66 | 16742 | 784.7812
66 | 16749 | 785.1094
66 | 16743 | 784.8281
66 | 16751 | 785.2031
66 | 16743 | 784.8281
66 | 16751 | 785.2031
66 | 16745 | 784.9219
66 | 16752 | 785.25
66 | 16744 | 784.875
66 | 16751 | 785.2031
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16743 | 784.8281
66 | 16749 | 785.1094
66 | 16739 | 784.6406
66 | 16749 | 785.1094
66 | 16740 | 784.6875
66 | 16748 | 785.0625
66 | 16741 | 784.7344
66 | 16750 | 785.1562
66 | 16742 | 784.7812
66 | 16751 | 785.2031
66 | 16742 | 784.7812
66 | 16751 | 785.2031
66 | 16744 | 784.875
66 | 16751 | 785.2031
66 | 16743 | 784.8281
66 | 16750 | 785.1562
66 | 16742 | 784.7812
66 | 16750 | 785.1562
66 | 16741 | 784.7344
66 | 16747 | 785.0156
66 | 16738 | 784.5938
66 | 16749 | 785.1094
66 | 16738 | 784.5938
66 | 16750 | 785.1562
66 | 16740 | 784.6875
66 | 16744 | 784.875
66 | 16750 | 785.1562
66 | 16740 | 784.6875
66 | 16751 | 785.2031
66 | 16742 | 784.7812
66 | 16750 | 785.1562
66 | 16742 | 784.7812
66 | 16744 | 784.875
66 | 16749 | 785.1094
66 | 16740 | 784.6875
66 | 16748 | 785.0625
66 | 16738 | 784.5938
66 | 16749 | 785.1094
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16741 | 784.7344
66 | 16750 | 785.1562
66 | 16740 | 784.6875
66 | 16750 | 785.1562
66 | 16749 | 785.1094
66 | 16743 | 784.8281
66 | 16747 | 785.0156
66 | 16738 | 784.5938
66 | 16748 | 785.0625
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16747 | 785.0156
66 | 16738 | 784.5938
66 | 16748 | 785.0625
66 | 16739 | 784.6406
66 | 16750 | 785.1562
66 | 16739 | 784.6406
66 | 16747 | 785.0156
66 | 16750 | 785.1562
66 | 16741 | 784.7344
66 | 16749 | 785.1094
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16744 | 784.875
66 | 16739 | 784.6406
66 | 16748 | 785.0625
66 | 16735 | 784.4531
66 | 16749 | 785.1094
66 | 16741 | 784.7344
66 | 16742 | 784.7812
66 | 16748 | 785.0625
66 | 16741 | 784.7344
66 | 16750 | 785.1562
66 | 16739 | 784.6406
66 | 16749 | 785.1094
66 | 16744 | 784.875
66 | 16739 | 784.6406
66 | 16749 | 785.1094
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16746 | 784.9688
66 | 16738 | 784.5938
66 | 16747 | 785.0156
66 | 16738 | 784.5938
66 | 16741 | 784.7344
66 | 16747 | 785.0156
66 | 16740 | 784.6875
66 | 16750 | 785.1562
66 | 16739 | 784.6406
66 | 16745 | 784.9219
66 | 16748 | 785.0625
66 | 16740 | 784.6875
66 | 16748 | 785.0625
66 | 16736 | 784.5
66 | 16741 | 784.7344
66 | 16747 | 785.0156
66 | 16737 | 784.5469
66 | 16746 | 784.9688
66 | 16736 | 784.5
66 | 16748 | 785.0625
66 | 16746 | 784.9688
66 | 16740 | 784.6875
66 | 16748 | 785.0625
66 | 16738 | 784.5938
66 | 16748 | 785.0625
66 | 16738 | 784.5938
66 | 16742 | 784.7812
66 | 16747 | 785.0156
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16736 | 784.5
66 | 16747 | 785.0156
66 | 16735 | 784.4531
66 | 16745 | 784.9219
66 | 16747 | 785.0156
66 | 16737 | 784.5469
66 | 16745 | 784.9219
66 | 16738 | 784.5938
66 | 16749 | 785.1094
66 | 16738 | 784.5938
66 | 16751 | 785.2031
66 | 16738 | 784.5938
66 | 16743 | 784.8281
66 | 16747 | 785.0156
66 | 16737 | 784.5469
66 | 16746 | 784.9688
66 | 16734 | 784.4062
66 | 16745 | 784.9219
66 | 16744 | 784.875
66 | 16736 | 784.5
66 | 16746 | 784.9688
66 | 16737 | 784.5469
66 | 16747 | 785.0156
66 | 16735 | 784.4531
66 | 16743 | 784.8281
66 | 16747 | 785.0156
66 | 16738 | 784.5938
66 | 16748 | 785.0625
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16747 | 785.0156
66 | 16737 | 784.5469
66 | 16746 | 784.9688
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16744 | 784.875
66 | 16746 | 784.9688
66 | 16739 | 784.6406
66 | 16749 | 785.1094
66 | 16740 | 784.6875
66 | 16738 | 784.5938
66 | 16747 | 785.0156
66 | 16737 | 784.5469
66 | 16740 | 784.6875
66 | 16745 | 784.9219
66 | 16734 | 784.4062
66 | 16746 | 784.9688
66 | 16738 | 784.5938
66 | 16735 | 784.4531
66 | 16745 | 784.9219
66 | 16736 | 784.5
66 | 16741 | 784.7344
66 | 16745 | 784.9219
66 | 16737 | 784.5469
66 | 16747 | 785.0156
66 | 16746 | 784.9688
66 | 16738 | 784.5938
66 | 16747 | 785.0156
66 | 16740 | 784.6875
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16744 | 784.875
66 | 16742 | 784.7812
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16737 | 784.5469
66 | 16746 | 784.9688
66 | 16735 | 784.4531
66 | 16740 | 784.6875
66 | 16746 | 784.9688
66 | 16736 | 784.5
66 | 16745 | 784.9219
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16744 | 784.875
66 | 16734 | 784.4062
66 | 16738 | 784.5938
66 | 16744 | 784.875
66 | 16736 | 784.5
66 | 16739 | 784.6406
66 | 16745 | 784.9219
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16745 | 784.9219
66 | 16739 | 784.6406
66 | 16745 | 784.9219
66 | 16742 | 784.7812
66 | 16734 | 784.4062
66 | 16743 | 784.8281
66 | 16742 | 784.7812
66 | 16734 | 784.4062
66 | 16744 | 784.875
66 | 16742 | 784.7812
66 | 16733 | 784.3594
66 | 16743 | 784.8281
66 | 16742 | 784.7812
66 | 16736 | 784.5
66 | 16745 | 784.9219
66 | 16745 | 784.9219
66 | 16734 | 784.4062
66 | 16744 | 784.875
66 | 16743 | 784.8281
66 | 16733 | 784.3594
66 | 16743 | 784.8281
66 | 16743 | 784.8281
66 | 16732 | 784.3125
66 | 16741 | 784.7344
66 | 16744 | 784.875
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16745 | 784.9219
66 | 16737 | 784.5469
66 | 16745 | 784.9219
66 | 16745 | 784.9219
66 | 16734 | 784.4062
66 | 16744 | 784.875
66 | 16745 | 784.9219
66 | 16733 | 784.3594
66 | 16740 | 784.6875
66 | 16743 | 784.8281
66 | 16732 | 784.3125
66 | 16740 | 784.6875
66 | 16742 | 784.7812
66 | 16734 | 784.4062
66 | 16738 | 784.5938
66 | 16746 | 784.9688
66 | 16739 | 784.6406
66 | 16736 | 784.5
66 | 16745 | 784.9219
66 | 16744 | 784.875
66 | 16734 | 784.4062
66 | 16744 | 784.875
66 | 16743 | 784.8281
66 | 16734 | 784.4062
66 | 16742 | 784.7812
66 | 16743 | 784.8281
66 | 16732 | 784.3125
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16741 | 784.7344
66 | 16733 | 784.3594
66 | 16744 | 784.875
66 | 16744 | 784.875
66 | 16734 | 784.4062
66 | 16740 | 784.6875
66 | 16745 | 784.9219
66 | 16738 | 784.5938
66 | 16735 | 784.4531
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16731 | 784.2656
66 | 16739 | 784.6406
66 | 16741 | 784.7344
66 | 16731 | 784.2656
66 | 16735 | 784.4531
66 | 16743 | 784.8281
66 | 16741 | 784.7344
66 | 16734 | 784.4062
66 | 16743 | 784.8281
66 | 16744 | 784.875
66 | 16733 | 784.3594
66 | 16739 | 784.6406
66 | 16744 | 784.875
66 | 16736 | 784.5
66 | 16733 | 784.3594
66 | 16742 | 784.7812
66 | 16742 | 784.7812
66 | 16730 | 784.2188
66 | 16734 | 784.4062
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16732 | 784.3125
66 | 16735 | 784.4531
66 | 16744 | 784.875
66 | 16743 | 784.8281
66 | 16733 | 784.3594
66 | 16736 | 784.5
66 | 16743 | 784.8281
66 | 16739 | 784.6406
66 | 16732 | 784.3125
66 | 16742 | 784.7812
66 | 16740 | 784.6875
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16731 | 784.2656
66 | 16735 | 784.4531
66 | 16742 | 784.7812
66 | 16741 | 784.7344
66 | 16733 | 784.3594
66 | 16737 | 784.5469
66 | 16744 | 784.875
66 | 16743 | 784.8281
66 | 16732 | 784.3125
66 | 16737 | 784.5469
66 | 16742 | 784.7812
66 | 16740 | 784.6875
66 | 16730 | 784.2188
66 | 16734 | 784.4062
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16730 | 784.2188
66 | 16735 | 784.4531
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16731 | 784.2656
66 | 16736 | 784.5
66 | 16743 | 784.8281
66 | 16742 | 784.7812
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16736 | 784.5
66 | 16730 | 784.2188
66 | 16738 | 784.5938
66 | 16741 | 784.7344
66 | 16737 | 784.5469
66 | 16729 | 784.1719
66 | 16738 | 784.5938
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16732 | 784.3125
66 | 16735 | 784.4531
66 | 16743 | 784.8281
66 | 16741 | 784.7344
66 | 16733 | 784.3594
66 | 16730 | 784.2188
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16729 | 784.1719
66 | 16733 | 784.3594
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16733 | 784.3594
66 | 16731 | 784.2656
66 | 16739 | 784.6406
66 | 16741 | 784.7344
66 | 16740 | 784.6875
66 | 16731 | 784.2656
66 | 16736 | 784.5
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16734 | 784.4062
66 | 16728 | 784.125
66 | 16733 | 784.3594
66 | 16740 | 784.6875
66 | 16738 | 784.5938
66 | 16728 | 784.125
66 | 16731 | 784.2656
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16730 | 784.2188
66 | 16735 | 784.4531
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16735 | 784.4531
66 | 16728 | 784.125
66 | 16732 | 784.3125
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16732 | 784.3125
66 | 16730 | 784.2188
66 | 16736 | 784.5
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16731 | 784.2656
66 | 16734 | 784.4062
66 | 16742 | 784.7812
66 | 16740 | 784.6875
66 | 16735 | 784.4531
66 | 16730 | 784.2188
66 | 16733 | 784.3594
66 | 16739 | 784.6406
66 | 16739 | 784.6406
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16730 | 784.2188
66 | 16732 | 784.3125
66 | 16741 | 784.7344
66 | 16741 | 784.7344
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16739 | 784.6406
66 | 16740 | 784.6875
66 | 16735 | 784.4531
66 | 16729 | 784.1719
66 | 16734 | 784.4062
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16730 | 784.2188
66 | 16727 | 784.0781
66 | 16736 | 784.5
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16735 | 784.4531
66 | 16730 | 784.2188
66 | 16733 | 784.3594
66 | 16740 | 784.6875
66 | 16740 | 784.6875
66 | 16733 | 784.3594
66 | 16728 | 784.125
66 | 16733 | 784.3594
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16732 | 784.3125
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16739 | 784.6406
66 | 16739 | 784.6406
66 | 16737 | 784.5469
66 | 16734 | 784.4062
66 | 16728 | 784.125
66 | 16732 | 784.3125
66 | 16740 | 784.6875
66 | 16741 | 784.7344
66 | 16739 | 784.6406
66 | 16734 | 784.4062
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16732 | 784.3125
66 | 16727 | 784.0781
66 | 16728 | 784.125
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16736 | 784.5
66 | 16729 | 784.1719
66 | 16731 | 784.2656
66 | 16737 | 784.5469
66 | 16738 | 784.5938
66 | 16739 | 784.6406
66 | 16736 | 784.5
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16734 | 784.4062
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16735 | 784.4531
66 | 16740 | 784.6875
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16732 | 784.3125
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16735 | 784.4531
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16725 | 783.9844
66 | 16726 | 784.0312
66 | 16730 | 784.2188
66 | 16737 | 784.5469
66 | 16739 | 784.6406
66 | 16739 | 784.6406
66 | 16738 | 784.5938
66 | 16740 | 784.6875
66 | 16732 | 784.3125
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16733 | 784.3594
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16724 | 783.9375
66 | 16727 | 784.0781
66 | 16731 | 784.2656
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16728 | 784.125
66 | 16723 | 783.8906
66 | 16725 | 783.9844
66 | 16726 | 784.0312
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16738 | 784.5938
66 | 16738 | 784.5938
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16727 | 784.0781
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16731 | 784.2656
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16728 | 784.125
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16726 | 784.0312
66 | 16730 | 784.2188
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16733 | 784.3594
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16736 | 784.5
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16724 | 783.9375
66 | 16727 | 784.0781
66 | 16729 | 784.1719
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16737 | 784.5469
66 | 16737 | 784.5469
66 | 16736 | 784.5
66 | 16737 | 784.5469
66 | 16735 | 784.4531
66 | 16736 | 784.5
66 | 16735 | 784.4531
66 | 16735 | 784.4531
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16736 | 784.5
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16735 | 784.4531
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16735 | 784.4531
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16734 | 784.4062
66 | 16734 | 784.4062
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16731 | 784.2656
66 | 16733 | 784.3594
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16733 | 784.3594
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16731 | 784.2656
66 | 16733 | 784.3594
66 | 16731 | 784.2656
66 | 16733 | 784.3594
66 | 16733 | 784.3594
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16731 | 784.2656
66 | 16729 | 784.1719
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16732 | 784.3125
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16729 | 784.1719
66 | 16731 | 784.2656
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16729 | 784.1719
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16729 | 784.1719
66 | 16729 | 784.1719
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16732 | 784.3125
66 | 16730 | 784.2188
66 | 16731 | 784.2656
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16729 | 784.1719
66 | 16728 | 784.125
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16727 | 784.0781
66 | 16729 | 784.1719
66 | 16727 | 784.0781
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16728 | 784.125
66 | 16730 | 784.2188
66 | 16730 | 784.2188
66 | 16728 | 784.125
66 | 16729 | 784.1719
66 | 16728 | 784.125
66 | 16728 | 784.125
66 | 16728 | 784.125
66 | 16726 | 784.0312
66 | 16726 | 784.0312
66 | 16726 | 784.0312
66 | 16724 | 783.9375
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16726 | 784.0312
66 | 16725 | 783.9844
66 | 16726 | 784.0312
66 | 16726 | 784.0312
66 | 16728 | 784.125
66 | 16726 | 784.0312
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16725 | 783.9844
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16726 | 784.0312
66 | 16724 | 783.9375
66 | 16727 | 784.0781
66 | 16724 | 783.9375
66 | 16727 | 784.0781
66 | 16725 | 783.9844
66 | 16727 | 784.0781
66 | 16724 | 783.9375
66 | 16726 | 784.0312
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16723 | 783.8906
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16725 | 783.9844
66 | 16724 | 783.9375
66 | 16724 | 783.9375
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16721 | 783.7969
66 | 16725 | 783.9844
66 | 16723 | 783.8906
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16722 | 783.8438
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16719 | 783.7031
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16721 | 783.7969
66 | 16722 | 783.8438
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16717 | 783.6094
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16719 | 783.7031
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16719 | 783.7031
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16718 | 783.6562
66 | 16716 | 783.5625
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16715 | 783.5156
66 | 16717 | 783.6094
66 | 16716 | 783.5625
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16718 | 783.6562
66 | 16717 | 783.6094
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16717 | 783.6094
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16717 | 783.6094
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16715 | 783.5156
66 | 16714 | 783.4688
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16717 | 783.6094
66 | 16716 | 783.5625
66 | 16715 | 783.5156
66 | 16717 | 783.6094
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16716 | 783.5625
66 | 16717 | 783.6094
66 | 16716 | 783.5625
66 | 16717 | 783.6094
66 | 16716 | 783.5625
66 | 16715 | 783.5156
66 | 16714 | 783.4688
66 | 16713 | 783.4219
66 | 16713 | 783.4219
66 | 16715 | 783.5156
66 | 16713 | 783.4219
66 | 16715 | 783.5156
66 | 16713 | 783.4219
66 | 16714 | 783.4688
66 | 16714 | 783.4688
66 | 16720 | 783.75
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16720 | 783.75
66 | 16722 | 783.8438
66 | 16720 | 783.75
66 | 16722 | 783.8438
66 | 16721 | 783.7969
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16724 | 783.9375
66 | 16723 | 783.8906
66 | 16724 | 783.9375
66 | 16718 | 783.6562
66 | 16716 | 783.5625
66 | 16715 | 783.5156
66 | 16714 | 783.4688
66 | 16711 | 783.3281
66 | 16712 | 783.375
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16724 | 783.9375
66 | 16720 | 783.75
66 | 16715 | 783.5156
66 | 16714 | 783.4688
66 | 16714 | 783.4688
66 | 16712 | 783.375
66 | 16710 | 783.2812
66 | 16714 | 783.4688
66 | 16718 | 783.6562
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16723 | 783.8906
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16724 | 783.9375
66 | 16718 | 783.6562
66 | 16714 | 783.4688
66 | 16710 | 783.2812
66 | 16712 | 783.375
66 | 16719 | 783.7031
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16714 | 783.4688
66 | 16712 | 783.375
66 | 16711 | 783.3281
66 | 16710 | 783.2812
66 | 16714 | 783.4688
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16722 | 783.8438
66 | 16722 | 783.8438
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16711 | 783.3281
66 | 16712 | 783.375
66 | 16710 | 783.2812
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16711 | 783.3281
66 | 16711 | 783.3281
66 | 16713 | 783.4219
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16722 | 783.8438
66 | 16714 | 783.4688
66 | 16711 | 783.3281
66 | 16708 | 783.1875
66 | 16713 | 783.4219
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16712 | 783.375
66 | 16710 | 783.2812
66 | 16709 | 783.2344
66 | 16713 | 783.4219
66 | 16718 | 783.6562
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16720 | 783.75
66 | 16714 | 783.4688
66 | 16710 | 783.2812
66 | 16709 | 783.2344
66 | 16713 | 783.4219
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16715 | 783.5156
66 | 16711 | 783.3281
66 | 16709 | 783.2344
66 | 16714 | 783.4688
66 | 16719 | 783.7031
66 | 16721 | 783.7969
66 | 16721 | 783.7969
66 | 16715 | 783.5156
66 | 16710 | 783.2812
66 | 16712 | 783.375
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16711 | 783.3281
66 | 16709 | 783.2344
66 | 16711 | 783.3281
66 | 16717 | 783.6094
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16711 | 783.3281
66 | 16709 | 783.2344
66 | 16715 | 783.5156
66 | 16720 | 783.75
66 | 16721 | 783.7969
66 | 16719 | 783.7031
66 | 16712 | 783.375
66 | 16709 | 783.2344
66 | 16714 | 783.4688
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16718 | 783.6562
66 | 16709 | 783.2344
66 | 16708 | 783.1875
66 | 16709 | 783.2344
66 | 16714 | 783.4688
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16716 | 783.5625
66 | 16711 | 783.3281
66 | 16709 | 783.2344
66 | 16711 | 783.3281
66 | 16719 | 783.7031
66 | 16720 | 783.75
66 | 16720 | 783.75
66 | 16714 | 783.4688
66 | 16709 | 783.2344
66 | 16708 | 783.1875
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16715 | 783.5156
66 | 16709 | 783.2344
66 | 16707 | 783.1406
66 | 16712 | 783.375
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16713 | 783.4219
66 | 16708 | 783.1875
66 | 16713 | 783.4219
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16709 | 783.2344
66 | 16707 | 783.1406
66 | 16711 | 783.3281
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16709 | 783.2344
66 | 16706 | 783.0938
66 | 16714 | 783.4688
66 | 16718 | 783.6562
66 | 16719 | 783.7031
66 | 16716 | 783.5625
66 | 16709 | 783.2344
66 | 16707 | 783.1406
66 | 16715 | 783.5156
66 | 16719 | 783.7031
66 | 16719 | 783.7031
66 | 16712 | 783.375
66 | 16708 | 783.1875
66 | 16709 | 783.2344
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16710 | 783.2812
66 | 16707 | 783.1406
66 | 16712 | 783.375
66 | 16716 | 783.5625
66 | 16719 | 783.7031
66 | 16711 | 783.3281
66 | 16708 | 783.1875
66 | 16715 | 783.5156
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16709 | 783.2344
66 | 16707 | 783.1406
66 | 16712 | 783.375
66 | 16715 | 783.5156
66 | 16715 | 783.5156
66 | 16708 | 783.1875
66 | 16705 | 783.0469
66 | 16711 | 783.3281
66 | 16717 | 783.6094
66 | 16717 | 783.6094
66 | 16708 | 783.1875
66 | 16707 | 783.1406
66 | 16714 | 783.4688
66 | 16718 | 783.6562
66 | 16718 | 783.6562
66 | 16711 | 783.3281
66 | 16707 | 783.1406
66 | 16712 | 783.375
66 | 16716 | 783.5625
66 | 16715 | 783.5156
66 | 16709 | 783.2344
66 | 16705 | 783.0469
66 | 16710 | 783.2812
66 | 16714 | 783.4688
66 | 16715 | 783.5156
66 | 16709 | 783.2344
66 | 16706 | 783.0938
66 | 16710 | 783.2812
66 | 16716 | 783.5625
66 | 16718 | 783.6562
66 | 16710 | 783.2812
66 | 16708 | 783.1875
66 | 16708 | 783.1875
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16710 | 783.2812
66 | 16705 | 783.0469
66 | 16707 | 783.1406
66 | 16714 | 783.4688
66 | 16715 | 783.5156
66 | 16707 | 783.1406
66 | 16705 | 783.0469
66 | 16711 | 783.3281
66 | 16715 | 783.5156
66 | 16716 | 783.5625
66 | 16710 | 783.2812
66 | 16706 | 783.0938
66 | 16711 | 783.3281
66 | 16715 | 783.5156
66 | 16715 | 783.5156
66 | 16706 | 783.0938
66 | 16703 | 782.9531
66 | 16713 | 783.4219
66 | 16714 | 783.4688
66 | 16715 | 783.5156
66 | 16708 | 783.1875
66 | 16704 | 783
66 | 16713 | 783.4219
66 | 16717 | 783.6094
66 | 16711 | 783.3281
66 | 16705 | 783.0469
66 | 16713 | 783.4219
66 | 16718 | 783.6562
66 | 16715 | 783.5156
66 | 16707 | 783.1406
66 | 16705 | 783.0469
66 | 16712 | 783.375
66 | 16716 | 783.5625
66 | 16711 | 783.3281
66 | 16704 | 783
66 | 16707 | 783.1406
66 | 16713 | 783.4219
66 | 16716 | 783.5625
66 | 16708 | 783.1875
66 | 16703 | 782.9531
66 | 16710 | 783.2812
66 | 16716 | 783.5625
66 | 16716 | 783.5625
66 | 16707 | 783.1406
66 | 16704 | 783
66 | 16712 | 783.375
66 | 16715 | 783.5156
66 | 16714 | 783.4688
66 | 16706 | 783.0938
66 | 16703 | 782.9531
66 | 16712 | 783.375
66 | 16714 | 783.4688
66 | 16707 | 783.1406
66 | 16703 | 782.9531
66 | 16711 | 783.3281
66 | 16715 | 783.5156
66 | 16714 | 783.4688
66 | 16705 | 783.0469
66 | 16707 | 783.1406
66 | 16714 | 783.4688
66 | 16714 | 783.4688
66 | 16707 | 783.1406
66 | 16704 | 783
66 | 16711 | 783.3281
66 | 16713 | 783.4219
66 | 16706 | 783.0938
66 | 16703 | 782.9531
66 | 16710 | 783.2812
66 | 16712 | 783.375
66 | 16714 | 783.4688
66 | 16705 | 783.0469
66 | 16710 | 783.2812
66 | 16715 | 783.5156
66 | 16712 | 783.375
66 | 16705 | 783.0469
66 | 16712 | 783.375
66 | 16715 | 783.5156
66 | 16707 | 783.1406
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16713 | 783.4219
66 | 16711 | 783.3281
66 | 16704 | 783
66 | 16706 | 783.0938
66 | 16712 | 783.375
66 | 16713 | 783.4219
66 | 16705 | 783.0469
66 | 16708 | 783.1875
66 | 16714 | 783.4688
66 | 16712 | 783.375
66 | 16704 | 783
66 | 16711 | 783.3281
66 | 16714 | 783.4688
66 | 16703 | 782.9531
66 | 16703 | 782.9531
66 | 16712 | 783.375
66 | 16713 | 783.4219
66 | 16702 | 782.9062
66 | 16707 | 783.1406
66 | 16712 | 783.375
66 | 16703 | 782.9531
66 | 16705 | 783.0469
66 | 16712 | 783.375
66 | 16712 | 783.375
66 | 16704 | 783
66 | 16711 | 783.3281
66 | 16715 | 783.5156
66 | 16705 | 783.0469
66 | 16707 | 783.1406
66 | 16714 | 783.4688
66 | 16704 | 783
66 | 16705 | 783.0469
66 | 16710 | 783.2812
66 | 16709 | 783.2344
66 | 16700 | 782.8125
66 | 16710 | 783.2812
66 | 16713 | 783.4219
66 | 16703 | 782.9531
66 | 16710 | 783.2812
66 | 16714 | 783.4688
66 | 16706 | 783.0938
66 | 16709 | 783.2344
66 | 16714 | 783.4688
66 | 16706 | 783.0938
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16711 | 783.3281
66 | 16703 | 782.9531
66 | 16706 | 783.0938
66 | 16712 | 783.375
66 | 16703 | 782.9531
66 | 16709 | 783.2344
66 | 16713 | 783.4219
66 | 16704 | 783
66 | 16708 | 783.1875
66 | 16713 | 783.4219
66 | 16705 | 783.0469
66 | 16710 | 783.2812
66 | 16713 | 783.4219
66 | 16702 | 782.9062
66 | 16708 | 783.1875
66 | 16710 | 783.2812
66 | 16700 | 782.8125
66 | 16708 | 783.1875
66 | 16710 | 783.2812
66 | 16700 | 782.8125
66 | 16707 | 783.1406
66 | 16712 | 783.375
66 | 16704 | 783
66 | 16710 | 783.2812
66 | 16712 | 783.375
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16712 | 783.375
66 | 16701 | 782.8594
66 | 16710 | 783.2812
66 | 16712 | 783.375
66 | 16700 | 782.8125
66 | 16707 | 783.1406
66 | 16709 | 783.2344
66 | 16702 | 782.9062
66 | 16709 | 783.2344
66 | 16713 | 783.4219
66 | 16701 | 782.8594
66 | 16711 | 783.3281
66 | 16712 | 783.375
66 | 16703 | 782.9531
66 | 16710 | 783.2812
66 | 16712 | 783.375
66 | 16701 | 782.8594
66 | 16708 | 783.1875
66 | 16709 | 783.2344
66 | 16700 | 782.8125
66 | 16707 | 783.1406
66 | 16710 | 783.2812
66 | 16699 | 782.7656
66 | 16708 | 783.1875
66 | 16711 | 783.3281
66 | 16702 | 782.9062
66 | 16712 | 783.375
66 | 16711 | 783.3281
66 | 16701 | 782.8594
66 | 16708 | 783.1875
66 | 16711 | 783.3281
66 | 16700 | 782.8125
66 | 16707 | 783.1406
66 | 16709 | 783.2344
66 | 16697 | 782.6719
66 | 16706 | 783.0938
66 | 16709 | 783.2344
66 | 16699 | 782.7656
66 | 16709 | 783.2344
66 | 16711 | 783.3281
66 | 16700 | 782.8125
66 | 16709 | 783.2344
66 | 16710 | 783.2812
66 | 16701 | 782.8594
66 | 16710 | 783.2812
66 | 16710 | 783.2812
66 | 16698 | 782.7188
66 | 16706 | 783.0938
66 | 16708 | 783.1875
66 | 16698 | 782.7188
66 | 16707 | 783.1406
66 | 16709 | 783.2344
66 | 16700 | 782.8125
66 | 16707 | 783.1406
66 | 16709 | 783.2344
66 | 16698 | 782.7188
66 | 16711 | 783.3281
66 | 16711 | 783.3281
66 | 16701 | 782.8594
66 | 16710 | 783.2812
66 | 16709 | 783.2344
66 | 16700 | 782.8125
66 | 16709 | 783.2344
66 | 16707 | 783.1406
66 | 16698 | 782.7188
66 | 16706 | 783.0938
66 | 16706 | 783.0938
66 | 16697 | 782.6719
66 | 16708 | 783.1875
66 | 16709 | 783.2344
66 | 16699 | 782.7656
66 | 16709 | 783.2344
66 | 16710 | 783.2812
66 | 16699 | 782.7656
66 | 16709 | 783.2344
66 | 16710 | 783.2812
66 | 16698 | 782.7188
66 | 16708 | 783.1875
66 | 16708 | 783.1875
66 | 16698 | 782.7188
66 | 16708 | 783.1875
66 | 16708 | 783.1875
66 | 16696 | 782.625
66 | 16706 | 783.0938
66 | 16707 | 783.1406
66 | 16697 | 782.6719
66 | 16709 | 783.2344
66 | 16710 | 783.2812
66 | 16699 | 782.7656
66 | 16710 | 783.2812
66 | 16709 | 783.2344
66 | 16698 | 782.7188
66 | 16707 | 783.1406
66 | 16707 | 783.1406
66 | 16695 | 782.5781
66 | 16706 | 783.0938
66 | 16707 | 783.1406
66 | 16697 | 782.6719
66 | 16706 | 783.0938
66 | 16708 | 783.1875
66 | 16699 | 782.7656
66 | 16708 | 783.1875
66 | 16710 | 783.2812
66 | 16698 | 782.7188
66 | 16709 | 783.2344
66 | 16709 | 783.2344
66 | 16698 | 782.7188
66 | 16706 | 783.0938
66 | 16706 | 783.0938
66 | 16696 | 782.625
66 | 16706 | 783.0938
66 | 16706 | 783.0938
66 | 16696 | 782.625
66 | 16706 | 783.0938
66 | 16706 | 783.0938
66 | 16698 | 782.7188
66 | 16707 | 783.1406
66 | 16708 | 783.1875
66 | 16699 | 782.7656
66 | 16708 | 783.1875
66 | 16706 | 783.0938
66 | 16697 | 782.6719
66 | 16705 | 783.0469
66 | 16705 | 783.0469
66 | 16697 | 782.6719
66 | 16705 | 783.0469
66 | 16704 | 783
66 | 16696 | 782.625
66 | 16706 | 783.0938
66 | 16704 | 783
66 | 16696 | 782.625
66 | 16707 | 783.1406
66 | 16707 | 783.1406
66 | 16697 | 782.6719
66 | 16708 | 783.1875
66 | 16706 | 783.0938
66 | 16696 | 782.625
66 | 16707 | 783.1406
66 | 16705 | 783.0469
66 | 16696 | 782.625
66 | 16706 | 783.0938
66 | 16705 | 783.0469
66 | 16694 | 782.5312
66 | 16704 | 783
66 | 16703 | 782.9531
66 | 16694 | 782.5312
66 | 16707 | 783.1406
66 | 16704 | 783
66 | 16696 | 782.625
66 | 16707 | 783.1406
66 | 16703 | 782.9531
66 | 16696 | 782.625
66 | 16707 | 783.1406
66 | 16706 | 783.0938
66 | 16696 | 782.625
66 | 16705 | 783.0469
66 | 16703 | 782.9531
66 | 16694 | 782.5312
66 | 16705 | 783.0469
66 | 16705 | 783.0469
66 | 16695 | 782.5781
66 | 16705 | 783.0469
66 | 16705 | 783.0469
66 | 16698 | 782.7188
66 | 16707 | 783.1406
66 | 16706 | 783.0938
66 | 16697 | 782.6719
66 | 16707 | 783.1406
66 | 16704 | 783
66 | 16695 | 782.5781
66 | 16705 | 783.0469
66 | 16701 | 782.8594
66 | 16693 | 782.4844
66 | 16703 | 782.9531
66 | 16702 | 782.9062
66 | 16694 | 782.5312
66 | 16704 | 783
66 | 16704 | 783
66 | 16695 | 782.5781
66 | 16706 | 783.0938
66 | 16701 | 782.8594
66 | 16696 | 782.625
66 | 16707 | 783.1406
66 | 16699 | 782.7656
66 | 16705 | 783.0469
66 | 16705 | 783.0469
66 | 16694 | 782.5312
66 | 16705 | 783.0469
66 | 16700 | 782.8125
66 | 16693 | 782.4844
66 | 16705 | 783.0469
66 | 16698 | 782.7188
66 | 16703 | 782.9531
66 | 16705 | 783.0469
66 | 16695 | 782.5781
66 | 16705 | 783.0469
66 | 16703 | 782.9531
66 | 16695 | 782.5781
66 | 16707 | 783.1406
66 | 16699 | 782.7656
66 | 16699 | 782.7656
66 | 16705 | 783.0469
66 | 16694 | 782.5312
66 | 16704 | 783
66 | 16701 | 782.8594
66 | 16692 | 782.4375
66 | 16704 | 783
66 | 16700 | 782.8125
66 | 16695 | 782.5781
66 | 16705 | 783.0469
66 | 16695 | 782.5781
66 | 16705 | 783.0469
66 | 16703 | 782.9531
66 | 16693 | 782.4844
66 | 16704 | 783
66 | 16702 | 782.9062
66 | 16694 | 782.5312
66 | 16704 | 783
66 | 16699 | 782.7656
66 | 16693 | 782.4844
66 | 16705 | 783.0469
66 | 16700 | 782.8125
66 | 16693 | 782.4844
66 | 16704 | 783
66 | 16694 | 782.5312
66 | 16704 | 783
66 | 16702 | 782.9062
66 | 16694 | 782.5312
66 | 16704 | 783
66 | 16700 | 782.8125
66 | 16697 | 782.6719
66 | 16705 | 783.0469
66 | 16692 | 782.4375
66 | 16703 | 782.9531
66 | 16697 | 782.6719
66 | 16691 | 782.3906
66 | 16703 | 782.9531
66 | 16695 | 782.5781
66 | 16703 | 782.9531
66 | 16700 | 782.8125
66 | 16694 | 782.5312
66 | 16705 | 783.0469
66 | 16694 | 782.5312
66 | 16705 | 783.0469
66 | 16699 | 782.7656
66 | 16695 | 782.5781
66 | 16704 | 783
66 | 16692 | 782.4375
66 | 16702 | 782.9062
66 | 16699 | 782.7656
66 | 16694 | 782.5312
66 | 16702 | 782.9062
66 | 16693 | 782.4844
66 | 16703 | 782.9531
66 | 16690 | 782.3438
66 | 16703 | 782.9531
66 | 16693 | 782.4844
66 | 16704 | 783
66 | 16699 | 782.7656
66 | 16693 | 782.4844
66 | 16705 | 783.0469
66 | 16693 | 782.4844
66 | 16704 | 783
66 | 16694 | 782.5312
66 | 16702 | 782.9062
66 | 16699 | 782.7656
66 | 16694 | 782.5312
66 | 16700 | 782.8125
66 | 16690 | 782.3438
66 | 16702 | 782.9062
66 | 16696 | 782.625
66 | 16700 | 782.8125
66 | 16700 | 782.8125
66 | 16693 | 782.4844
66 | 16704 | 783
66 | 16693 | 782.4844
66 | 16704 | 783
66 | 16698 | 782.7188
66 | 16696 | 782.625
66 | 16702 | 782.9062
66 | 16695 | 782.5781
66 | 16702 | 782.9062
66 | 16693 | 782.4844
66 | 16702 | 782.9062
66 | 16693 | 782.4844
66 | 16703 | 782.9531
66 | 16694 | 782.5312
66 | 16703 | 782.9531
66 | 16694 | 782.5312
66 | 16705 | 783.0469
66 | 16696 | 782.625
66 | 16703 | 782.9531
66 | 16699 | 782.7656
66 | 16693 | 782.4844
66 | 16701 | 782.8594
66 | 16691 | 782.3906
66 | 16701 | 782.8594
66 | 16693 | 782.4844
66 | 16700 | 782.8125
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16702 | 782.9062
66 | 16690 | 782.3438
66 | 16704 | 783
66 | 16691 | 782.3906
66 | 16702 | 782.9062
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16701 | 782.8594
66 | 16691 | 782.3906
66 | 16702 | 782.9062
66 | 16690 | 782.3438
66 | 16701 | 782.8594
66 | 16695 | 782.5781
66 | 16695 | 782.5781
66 | 16701 | 782.8594
66 | 16693 | 782.4844
66 | 16701 | 782.8594
66 | 16693 | 782.4844
66 | 16703 | 782.9531
66 | 16694 | 782.5312
66 | 16703 | 782.9531
66 | 16694 | 782.5312
66 | 16700 | 782.8125
66 | 16693 | 782.4844
66 | 16701 | 782.8594
66 | 16690 | 782.3438
66 | 16701 | 782.8594
66 | 16689 | 782.2969
66 | 16700 | 782.8125
66 | 16691 | 782.3906
66 | 16701 | 782.8594
66 | 16696 | 782.625
66 | 16691 | 782.3906
66 | 16701 | 782.8594
66 | 16692 | 782.4375
66 | 16702 | 782.9062
66 | 16690 | 782.3438
66 | 16702 | 782.9062
66 | 16690 | 782.3438
66 | 16701 | 782.8594
66 | 16689 | 782.2969
66 | 16701 | 782.8594
66 | 16686 | 782.1562
66 | 16700 | 782.8125
66 | 16688 | 782.25
66 | 16701 | 782.8594
66 | 16689 | 782.2969
66 | 16700 | 782.8125
66 | 16690 | 782.3438
66 | 16702 | 782.9062
66 | 16694 | 782.5312
66 | 16701 | 782.8594
66 | 16697 | 782.6719
66 | 16693 | 782.4844
66 | 16699 | 782.7656
66 | 16692 | 782.4375
66 | 16699 | 782.7656
66 | 16692 | 782.4375
66 | 16698 | 782.7188
66 | 16692 | 782.4375
66 | 16698 | 782.7188
66 | 16690 | 782.3438
66 | 16700 | 782.8125
66 | 16692 | 782.4375
66 | 16700 | 782.8125
66 | 16693 | 782.4844
66 | 16700 | 782.8125
66 | 16694 | 782.5312
66 | 16699 | 782.7656
66 | 16691 | 782.3906
66 | 16698 | 782.7188
66 | 16692 | 782.4375
66 | 16697 | 782.6719
66 | 16693 | 782.4844
66 | 16698 | 782.7188
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16698 | 782.7188
66 | 16695 | 782.5781
66 | 16698 | 782.7188
66 | 16696 | 782.625
66 | 16698 | 782.7188
66 | 16694 | 782.5312
66 | 16697 | 782.6719
66 | 16692 | 782.4375
66 | 16697 | 782.6719
66 | 16689 | 782.2969
66 | 16699 | 782.7656
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16699 | 782.7656
66 | 16693 | 782.4844
66 | 16697 | 782.6719
66 | 16692 | 782.4375
66 | 16698 | 782.7188
66 | 16693 | 782.4844
66 | 16698 | 782.7188
66 | 16693 | 782.4844
66 | 16697 | 782.6719
66 | 16689 | 782.2969
66 | 16696 | 782.625
66 | 16689 | 782.2969
66 | 16695 | 782.5781
66 | 16691 | 782.3906
66 | 16697 | 782.6719
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16697 | 782.6719
66 | 16694 | 782.5312
66 | 16698 | 782.7188
66 | 16692 | 782.4375
66 | 16697 | 782.6719
66 | 16691 | 782.3906
66 | 16695 | 782.5781
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16691 | 782.3906
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16697 | 782.6719
66 | 16695 | 782.5781
66 | 16696 | 782.625
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16692 | 782.4375
66 | 16696 | 782.625
66 | 16691 | 782.3906
66 | 16694 | 782.5312
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16694 | 782.5312
66 | 16697 | 782.6719
66 | 16693 | 782.4844
66 | 16697 | 782.6719
66 | 16693 | 782.4844
66 | 16696 | 782.625
66 | 16693 | 782.4844
66 | 16693 | 782.4844
66 | 16691 | 782.3906
66 | 16695 | 782.5781
66 | 16692 | 782.4375
66 | 16693 | 782.4844
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16692 | 782.4375
66 | 16694 | 782.5312
66 | 16692 | 782.4375
66 | 16695 | 782.5781
66 | 16690 | 782.3438
66 | 16694 | 782.5312
66 | 16696 | 782.625
66 | 16695 | 782.5781
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16691 | 782.3906
66 | 16693 | 782.4844
66 | 16690 | 782.3438
66 | 16694 | 782.5312
66 | 16689 | 782.2969
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16694 | 782.5312
66 | 16692 | 782.4375
66 | 16697 | 782.6719
66 | 16693 | 782.4844
66 | 16695 | 782.5781
66 | 16690 | 782.3438
66 | 16694 | 782.5312
66 | 16689 | 782.2969
66 | 16693 | 782.4844
66 | 16688 | 782.25
66 | 16693 | 782.4844
66 | 16686 | 782.1562
66 | 16694 | 782.5312
66 | 16689 | 782.2969
66 | 16693 | 782.4844
66 | 16689 | 782.2969
66 | 16696 | 782.625
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16692 | 782.4375
66 | 16695 | 782.5781
66 | 16689 | 782.2969
66 | 16694 | 782.5312
66 | 16688 | 782.25
66 | 16691 | 782.3906
66 | 16687 | 782.2031
66 | 16691 | 782.3906
66 | 16689 | 782.2969
66 | 16692 | 782.4375
66 | 16688 | 782.25
66 | 16693 | 782.4844
66 | 16692 | 782.4375
66 | 16693 | 782.4844
66 | 16690 | 782.3438
66 | 16693 | 782.4844
66 | 16691 | 782.3906
66 | 16690 | 782.3438
66 | 16690 | 782.3438
66 | 16691 | 782.3906
66 | 16689 | 782.2969
66 | 16692 | 782.4375
66 | 16689 | 782.2969
66 | 16692 | 782.4375
66 | 16690 | 782.3438
66 | 16692 | 782.4375
66 | 16692 | 782.4375
66 | 16692 | 782.4375
66 | 16691 | 782.3906
66 | 16694 | 782.5312
66 | 16693 | 782.4844
66 | 16691 | 782.3906
66 | 16696 | 782.625
66 | 16685 | 782.1094
66 | 16695 | 782.5781
66 | 16684 | 782.0625
66 | 16692 | 782.4375
66 | 16686 | 782.1562
66 | 16691 | 782.3906
66 | 16689 | 782.2969
66 | 16691 | 782.3906
66 | 16690 | 782.3438
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16690 | 782.3438
66 | 16696 | 782.625
66 | 16686 | 782.1562
66 | 16694 | 782.5312
66 | 16686 | 782.1562
66 | 16696 | 782.625
66 | 16686 | 782.1562
66 | 16691 | 782.3906
66 | 16689 | 782.2969
66 | 16690 | 782.3438
66 | 16689 | 782.2969
66 | 16691 | 782.3906
66 | 16690 | 782.3438
66 | 16692 | 782.4375
66 | 16691 | 782.3906
66 | 16692 | 782.4375
66 | 16693 | 782.4844
66 | 16689 | 782.2969
66 | 16695 | 782.5781
66 | 16688 | 782.25
66 | 16694 | 782.5312
66 | 16683 | 782.0156
66 | 16692 | 782.4375
66 | 16682 | 781.9688
66 | 16691 | 782.3906
66 | 16688 | 782.25
66 | 16688 | 782.25
66 | 16688 | 782.25
66 | 16689 | 782.2969
66 | 16695 | 782.5781
66 | 16687 | 782.2031
66 | 16695 | 782.5781
66 | 16684 | 782.0625
66 | 16694 | 782.5312
66 | 16685 | 782.1094
66 | 16691 | 782.3906
66 | 16687 | 782.2031
66 | 16688 | 782.25
66 | 16688 | 782.25
66 | 16689 | 782.2969
66 | 16688 | 782.25
66 | 16692 | 782.4375
66 | 16689 | 782.2969
66 | 16690 | 782.3438
66 | 16692 | 782.4375
66 | 16690 | 782.3438
66 | 16695 | 782.5781
66 | 16689 | 782.2969
66 | 16690 | 782.3438
66 | 16687 | 782.2031
66 | 16693 | 782.4844
66 | 16685 | 782.1094
66 | 16693 | 782.4844
66 | 16682 | 781.9688
66 | 16693 | 782.4844
66 | 16681 | 781.9219
66 | 16694 | 782.5312
66 | 16686 | 782.1562
66 | 16689 | 782.2969
66 | 16694 | 782.5312
66 | 16686 | 782.1562
66 | 16693 | 782.4844
66 | 16685 | 782.1094
66 | 16692 | 782.4375
66 | 16690 | 782.3438
66 | 16687 | 782.2031
66 | 16693 | 782.4844
66 | 16683 | 782.0156
66 | 16693 | 782.4844
66 | 16684 | 782.0625
66 | 16686 | 782.1562
66 | 16694 | 782.5312
66 | 16684 | 782.0625
66 | 16694 | 782.5312
66 | 16689 | 782.2969
66 | 16690 | 782.3438
66 | 16694 | 782.5312
66 | 16687 | 782.2031
66 | 16695 | 782.5781
66 | 16685 | 782.1094
66 | 16688 | 782.25
66 | 16692 | 782.4375
66 | 16683 | 782.0156
66 | 16693 | 782.4844
66 | 16688 | 782.25
66 | 16688 | 782.25
66 | 16692 | 782.4375
66 | 16687 | 782.2031
66 | 16695 | 782.5781
66 | 16687 | 782.2031
66 | 16695 | 782.5781
66 | 16688 | 782.25
66 | 16688 | 782.25
66 | 16693 | 782.4844
66 | 16684 | 782.0625
66 | 16694 | 782.5312
66 | 16686 | 782.1562
66 | 16686 | 782.1562
66 | 16690 | 782.3438
66 | 16683 | 782.0156
66 | 16692 | 782.4375
66 | 16684 | 782.0625
66 | 16687 | 782.2031
66 | 16693 | 782.4844
66 | 16686 | 782.1562
66 | 16693 | 782.4844
66 | 16686 | 782.1562
66 | 16691 | 782.3906
66 | 16690 | 782.3438
66 | 16684 | 782.0625
66 | 16692 | 782.4375
66 | 16683 | 782.0156
66 | 16684 | 782.0625
66 | 16691 | 782.3906
66 | 16680 | 781.875
66 | 16691 | 782.3906
66 | 16686 | 782.1562
66 | 16686 | 782.1562
66 | 16693 | 782.4844
66 | 16686 | 782.1562
66 | 16687 | 782.2031
66 | 16692 | 782.4375
66 | 16682 | 781.9688
66 | 16693 | 782.4844
66 | 16686 | 782.1562
66 | 16684 | 782.0625
66 | 16692 | 782.4375
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16690 | 782.3438
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16690 | 782.3438
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16693 | 782.4844
66 | 16685 | 782.1094
66 | 16686 | 782.1562
66 | 16693 | 782.4844
66 | 16684 | 782.0625
66 | 16685 | 782.1094
66 | 16692 | 782.4375
66 | 16685 | 782.1094
66 | 16684 | 782.0625
66 | 16691 | 782.3906
66 | 16686 | 782.1562
66 | 16681 | 781.9219
66 | 16691 | 782.3906
66 | 16685 | 782.1094
66 | 16685 | 782.1094
66 | 16693 | 782.4844
66 | 16685 | 782.1094
66 | 16684 | 782.0625
66 | 16694 | 782.5312
66 | 16686 | 782.1562
66 | 16686 | 782.1562
66 | 16691 | 782.3906
66 | 16679 | 781.8281
66 | 16690 | 782.3438
66 | 16685 | 782.1094
66 | 16681 | 781.9219
66 | 16691 | 782.3906
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16691 | 782.3906
66 | 16687 | 782.2031
66 | 16683 | 782.0156
66 | 16691 | 782.3906
66 | 16683 | 782.0156
66 | 16688 | 782.25
66 | 16692 | 782.4375
66 | 16681 | 781.9219
66 | 16691 | 782.3906
66 | 16684 | 782.0625
66 | 16682 | 781.9688
66 | 16688 | 782.25
66 | 16683 | 782.0156
66 | 16682 | 781.9688
66 | 16691 | 782.3906
66 | 16685 | 782.1094
66 | 16684 | 782.0625
66 | 16692 | 782.4375
66 | 16686 | 782.1562
66 | 16685 | 782.1094
66 | 16693 | 782.4844
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16692 | 782.4375
66 | 16685 | 782.1094
66 | 16682 | 781.9688
66 | 16690 | 782.3438
66 | 16682 | 781.9688
66 | 16681 | 781.9219
66 | 16688 | 782.25
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16691 | 782.3906
66 | 16687 | 782.2031
66 | 16683 | 782.0156
66 | 16691 | 782.3906
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16690 | 782.3438
66 | 16683 | 782.0156
66 | 16682 | 781.9688
66 | 16689 | 782.2969
66 | 16684 | 782.0625
66 | 16681 | 781.9219
66 | 16688 | 782.25
66 | 16682 | 781.9688
66 | 16680 | 781.875
66 | 16689 | 782.2969
66 | 16684 | 782.0625
66 | 16685 | 782.1094
66 | 16691 | 782.3906
66 | 16685 | 782.1094
66 | 16682 | 781.9688
66 | 16691 | 782.3906
66 | 16686 | 782.1562
66 | 16679 | 781.8281
66 | 16687 | 782.2031
66 | 16682 | 781.9688
66 | 16677 | 781.7344
66 | 16688 | 782.25
66 | 16682 | 781.9688
66 | 16679 | 781.8281
66 | 16689 | 782.2969
66 | 16685 | 782.1094
66 | 16680 | 781.875
66 | 16690 | 782.3438
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16690 | 782.3438
66 | 16684 | 782.0625
66 | 16679 | 781.8281
66 | 16687 | 782.2031
66 | 16683 | 782.0156
66 | 16679 | 781.8281
66 | 16686 | 782.1562
66 | 16682 | 781.9688
66 | 16679 | 781.8281
66 | 16689 | 782.2969
66 | 16684 | 782.0625
66 | 16679 | 781.8281
66 | 16690 | 782.3438
66 | 16686 | 782.1562
66 | 16679 | 781.8281
66 | 16690 | 782.3438
66 | 16686 | 782.1562
66 | 16678 | 781.7812
66 | 16687 | 782.2031
66 | 16684 | 782.0625
66 | 16677 | 781.7344
66 | 16686 | 782.1562
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16687 | 782.2031
66 | 16684 | 782.0625
66 | 16678 | 781.7812
66 | 16688 | 782.25
66 | 16685 | 782.1094
66 | 16682 | 781.9688
66 | 16688 | 782.25
66 | 16684 | 782.0625
66 | 16678 | 781.7812
66 | 16687 | 782.2031
66 | 16685 | 782.1094
66 | 16677 | 781.7344
66 | 16686 | 782.1562
66 | 16683 | 782.0156
66 | 16677 | 781.7344
66 | 16687 | 782.2031
66 | 16682 | 781.9688
66 | 16678 | 781.7812
66 | 16687 | 782.2031
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16689 | 782.2969
66 | 16684 | 782.0625
66 | 16678 | 781.7812
66 | 16687 | 782.2031
66 | 16682 | 781.9688
66 | 16676 | 781.6875
66 | 16686 | 782.1562
66 | 16682 | 781.9688
66 | 16678 | 781.7812
66 | 16684 | 782.0625
66 | 16682 | 781.9688
66 | 16676 | 781.6875
66 | 16687 | 782.2031
66 | 16683 | 782.0156
66 | 16679 | 781.8281
66 | 16688 | 782.25
66 | 16686 | 782.1562
66 | 16678 | 781.7812
66 | 16688 | 782.25
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16685 | 782.1094
66 | 16681 | 781.9219
66 | 16674 | 781.5938
66 | 16684 | 782.0625
66 | 16682 | 781.9688
66 | 16677 | 781.7344
66 | 16686 | 782.1562
66 | 16683 | 782.0156
66 | 16677 | 781.7344
66 | 16688 | 782.25
66 | 16686 | 782.1562
66 | 16678 | 781.7812
66 | 16686 | 782.1562
66 | 16682 | 781.9688
66 | 16675 | 781.6406
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16674 | 781.5938
66 | 16684 | 782.0625
66 | 16681 | 781.9219
66 | 16675 | 781.6406
66 | 16686 | 782.1562
66 | 16683 | 782.0156
66 | 16678 | 781.7812
66 | 16685 | 782.1094
66 | 16682 | 781.9688
66 | 16677 | 781.7344
66 | 16685 | 782.1094
66 | 16683 | 782.0156
66 | 16675 | 781.6406
66 | 16684 | 782.0625
66 | 16682 | 781.9688
66 | 16673 | 781.5469
66 | 16684 | 782.0625
66 | 16681 | 781.9219
66 | 16675 | 781.6406
66 | 16685 | 782.1094
66 | 16684 | 782.0625
66 | 16675 | 781.6406
66 | 16685 | 782.1094
66 | 16684 | 782.0625
66 | 16677 | 781.7344
66 | 16686 | 782.1562
66 | 16683 | 782.0156
66 | 16675 | 781.6406
66 | 16683 | 782.0156
66 | 16684 | 782.0625
66 | 16672 | 781.5
66 | 16682 | 781.9688
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16682 | 781.9688
66 | 16684 | 782.0625
66 | 16674 | 781.5938
66 | 16684 | 782.0625
66 | 16683 | 782.0156
66 | 16676 | 781.6875
66 | 16684 | 782.0625
66 | 16684 | 782.0625
66 | 16674 | 781.5938
66 | 16683 | 782.0156
66 | 16683 | 782.0156
66 | 16672 | 781.5
66 | 16681 | 781.9219
66 | 16682 | 781.9688
66 | 16671 | 781.4531
66 | 16682 | 781.9688
66 | 16685 | 782.1094
66 | 16673 | 781.5469
66 | 16684 | 782.0625
66 | 16684 | 782.0625
66 | 16674 | 781.5938
66 | 16686 | 782.1562
66 | 16685 | 782.1094
66 | 16674 | 781.5938
66 | 16683 | 782.0156
66 | 16681 | 781.9219
66 | 16672 | 781.5
66 | 16681 | 781.9219
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16681 | 781.9219
66 | 16682 | 781.9688
66 | 16673 | 781.5469
66 | 16681 | 781.9219
66 | 16683 | 782.0156
66 | 16673 | 781.5469
66 | 16683 | 782.0156
66 | 16684 | 782.0625
66 | 16674 | 781.5938
66 | 16683 | 782.0156
66 | 16684 | 782.0625
66 | 16673 | 781.5469
66 | 16681 | 781.9219
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16680 | 781.875
66 | 16682 | 781.9688
66 | 16671 | 781.4531
66 | 16679 | 781.8281
66 | 16683 | 782.0156
66 | 16672 | 781.5
66 | 16682 | 781.9688
66 | 16684 | 782.0625
66 | 16672 | 781.5
66 | 16682 | 781.9688
66 | 16680 | 781.875
66 | 16673 | 781.5469
66 | 16681 | 781.9219
66 | 16681 | 781.9219
66 | 16672 | 781.5
66 | 16679 | 781.8281
66 | 16683 | 782.0156
66 | 16671 | 781.4531
66 | 16680 | 781.875
66 | 16682 | 781.9688
66 | 16671 | 781.4531
66 | 16682 | 781.9688
66 | 16683 | 782.0156
66 | 16672 | 781.5
66 | 16681 | 781.9219
66 | 16684 | 782.0625
66 | 16673 | 781.5469
66 | 16681 | 781.9219
66 | 16683 | 782.0156
66 | 16671 | 781.4531
66 | 16678 | 781.7812
66 | 16682 | 781.9688
66 | 16680 | 781.875
66 | 16670 | 781.4062
66 | 16680 | 781.875
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16681 | 781.9219
66 | 16684 | 782.0625
66 | 16675 | 781.6406
66 | 16676 | 781.6875
66 | 16683 | 782.0156
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16680 | 781.875
66 | 16680 | 781.875
66 | 16670 | 781.4062
66 | 16677 | 781.7344
66 | 16680 | 781.875
66 | 16670 | 781.4062
66 | 16679 | 781.8281
66 | 16683 | 782.0156
66 | 16672 | 781.5
66 | 16680 | 781.875
66 | 16683 | 782.0156
66 | 16672 | 781.5
66 | 16679 | 781.8281
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16678 | 781.7812
66 | 16682 | 781.9688
66 | 16679 | 781.8281
66 | 16670 | 781.4062
66 | 16678 | 781.7812
66 | 16681 | 781.9219
66 | 16671 | 781.4531
66 | 16679 | 781.8281
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16679 | 781.8281
66 | 16684 | 782.0625
66 | 16681 | 781.9219
66 | 16673 | 781.5469
66 | 16680 | 781.875
66 | 16682 | 781.9688
66 | 16672 | 781.5
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16680 | 781.875
66 | 16673 | 781.5469
66 | 16672 | 781.5
66 | 16679 | 781.8281
66 | 16681 | 781.9219
66 | 16672 | 781.5
66 | 16677 | 781.7344
66 | 16684 | 782.0625
66 | 16681 | 781.9219
66 | 16673 | 781.5469
66 | 16676 | 781.6875
66 | 16682 | 781.9688
66 | 16682 | 781.9688
66 | 16671 | 781.4531
66 | 16672 | 781.5
66 | 16678 | 781.7812
66 | 16681 | 781.9219
66 | 16672 | 781.5
66 | 16670 | 781.4062
66 | 16678 | 781.7812
66 | 16681 | 781.9219
66 | 16671 | 781.4531
66 | 16674 | 781.5938
66 | 16681 | 781.9219
66 | 16683 | 782.0156
66 | 16671 | 781.4531
66 | 16679 | 781.8281
66 | 16681 | 781.9219
66 | 16671 | 781.4531
66 | 16671 | 781.4531
66 | 16678 | 781.7812
66 | 16679 | 781.8281
66 | 16670 | 781.4062
66 | 16676 | 781.6875
66 | 16681 | 781.9219
66 | 16672 | 781.5
66 | 16672 | 781.5
66 | 16681 | 781.9219
66 | 16680 | 781.875
66 | 16673 | 781.5469
66 | 16675 | 781.6406
66 | 16681 | 781.9219
66 | 16681 | 781.9219
66 | 16671 | 781.4531
66 | 16672 | 781.5
66 | 16678 | 781.7812
66 | 16678 | 781.7812
66 | 16670 | 781.4062
66 | 16671 | 781.4531
66 | 16678 | 781.7812
66 | 16679 | 781.8281
66 | 16670 | 781.4062
66 | 16677 | 781.7344
66 | 16682 | 781.9688
66 | 16675 | 781.6406
66 | 16671 | 781.4531
66 | 16677 | 781.7344
66 | 16681 | 781.9219
66 | 16670 | 781.4062
66 | 16675 | 781.6406
66 | 16680 | 781.875
66 | 16678 | 781.7812
66 | 16669 | 781.3594
66 | 16669 | 781.3594
66 | 16676 | 781.6875
66 | 16679 | 781.8281
66 | 16679 | 781.8281
66 | 16670 | 781.4062
66 | 16672 | 781.5
66 | 16678 | 781.7812
66 | 16681 | 781.9219
66 | 16671 | 781.4531
66 | 16671 | 781.4531
66 | 16676 | 781.6875
66 | 16679 | 781.8281
66 | 16669 | 781.3594
66 | 16668 | 781.3125
66 | 16675 | 781.6406
66 | 16677 | 781.7344
66 | 16670 | 781.4062
66 | 16670 | 781.4062
66 | 16677 | 781.7344
66 | 16679 | 781.8281
66 | 16674 | 781.5938
66 | 16670 | 781.4062
66 | 16676 | 781.6875
66 | 16680 | 781.875
66 | 16678 | 781.7812
66 | 16668 | 781.3125
66 | 16673 | 781.5469
66 | 16676 | 781.6875
66 | 16679 | 781.8281
66 | 16672 | 781.5
66 | 16667 | 781.2656
66 | 16674 | 781.5938
66 | 16679 | 781.8281
66 | 16679 | 781.8281
66 | 16669 | 781.3594
66 | 16670 | 781.4062
66 | 16677 | 781.7344
66 | 16681 | 781.9219
66 | 16674 | 781.5938
66 | 16670 | 781.4062
66 | 16676 | 781.6875
66 | 16680 | 781.875
66 | 16676 | 781.6875
66 | 16668 | 781.3125
66 | 16674 | 781.5938
66 | 16678 | 781.7812
66 | 16677 | 781.7344
66 | 16667 | 781.2656
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16680 | 781.875
66 | 16674 | 781.5938
66 | 16670 | 781.4062
66 | 16676 | 781.6875
66 | 16682 | 781.9688
66 | 16680 | 781.875
66 | 16672 | 781.5
66 | 16668 | 781.3125
66 | 16675 | 781.6406
66 | 16678 | 781.7812
66 | 16674 | 781.5938
66 | 16666 | 781.2188
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16679 | 781.8281
66 | 16669 | 781.3594
66 | 16671 | 781.4531
66 | 16678 | 781.7812
66 | 16679 | 781.8281
66 | 16670 | 781.4062
66 | 16669 | 781.3594
66 | 16673 | 781.5469
66 | 16679 | 781.8281
66 | 16676 | 781.6875
66 | 16667 | 781.2656
66 | 16671 | 781.4531
66 | 16676 | 781.6875
66 | 16677 | 781.7344
66 | 16669 | 781.3594
66 | 16667 | 781.2656
66 | 16675 | 781.6406
66 | 16678 | 781.7812
66 | 16672 | 781.5
66 | 16668 | 781.3125
66 | 16675 | 781.6406
66 | 16678 | 781.7812
66 | 16677 | 781.7344
66 | 16669 | 781.3594
66 | 16667 | 781.2656
66 | 16673 | 781.5469
66 | 16676 | 781.6875
66 | 16675 | 781.6406
66 | 16668 | 781.3125
66 | 16665 | 781.1719
66 | 16673 | 781.5469
66 | 16677 | 781.7344
66 | 16677 | 781.7344
66 | 16670 | 781.4062
66 | 16668 | 781.3125
66 | 16673 | 781.5469
66 | 16676 | 781.6875
66 | 16677 | 781.7344
66 | 16670 | 781.4062
66 | 16666 | 781.2188
66 | 16672 | 781.5
66 | 16676 | 781.6875
66 | 16675 | 781.6406
66 | 16669 | 781.3594
66 | 16664 | 781.125
66 | 16669 | 781.3594
66 | 16675 | 781.6406
66 | 16677 | 781.7344
66 | 16670 | 781.4062
66 | 16668 | 781.3125
66 | 16668 | 781.3125
66 | 16674 | 781.5938
66 | 16678 | 781.7812
66 | 16674 | 781.5938
66 | 16668 | 781.3125
66 | 16666 | 781.2188
66 | 16671 | 781.4531
66 | 16673 | 781.5469
66 | 16676 | 781.6875
66 | 16675 | 781.6406
66 | 16669 | 781.3594
66 | 16665 | 781.1719
66 | 16669 | 781.3594
66 | 16673 | 781.5469
66 | 16678 | 781.7812
66 | 16677 | 781.7344
66 | 16671 | 781.4531
66 | 16667 | 781.2656
66 | 16670 | 781.4062
66 | 16674 | 781.5938
66 | 16677 | 781.7344
66 | 16675 | 781.6406
66 | 16671 | 781.4531
66 | 16665 | 781.1719
66 | 16669 | 781.3594
66 | 16672 | 781.5
66 | 16675 | 781.6406
66 | 16676 | 781.6875
66 | 16670 | 781.4062
66 | 16668 | 781.3125
66 | 16666 | 781.2188
66 | 16671 | 781.4531
66 | 16674 | 781.5938
66 | 16678 | 781.7812
66 | 16673 | 781.5469
66 | 16668 | 781.3125
66 | 16664 | 781.125
66 | 16669 | 781.3594
66 | 16671 | 781.4531
66 | 16674 | 781.5938
66 | 16674 | 781.5938
66 | 16671 | 781.4531
66 | 16668 | 781.3125
66 | 16666 | 781.2188
66 | 16665 | 781.1719
66 | 16668 | 781.3125
66 | 16672 | 781.5
66 | 16677 | 781.7344
66 | 16676 | 781.6875
66 | 16671 | 781.4531
66 | 16667 | 781.2656
66 | 16665 | 781.1719
66 | 16670 | 781.4062
66 | 16674 | 781.5938
66 | 16675 | 781.6406
66 | 16668 | 781.3125
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16670 | 781.4062
66 | 16673 | 781.5469
66 | 16675 | 781.6406
66 | 16673 | 781.5469
66 | 16668 | 781.3125
66 | 16667 | 781.2656
66 | 16671 | 781.4531
66 | 16675 | 781.6406
66 | 16675 | 781.6406
66 | 16674 | 781.5938
66 | 16670 | 781.4062
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16668 | 781.3125
66 | 16669 | 781.3594
66 | 16673 | 781.5469
66 | 16672 | 781.5
66 | 16667 | 781.2656
66 | 16666 | 781.2188
66 | 16664 | 781.125
66 | 16667 | 781.2656
66 | 16672 | 781.5
66 | 16674 | 781.5938
66 | 16675 | 781.6406
66 | 16670 | 781.4062
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16667 | 781.2656
66 | 16669 | 781.3594
66 | 16672 | 781.5
66 | 16674 | 781.5938
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16666 | 781.2188
66 | 16670 | 781.4062
66 | 16675 | 781.6406
66 | 16674 | 781.5938
66 | 16667 | 781.2656
66 | 16665 | 781.1719
66 | 16664 | 781.125
66 | 16668 | 781.3125
66 | 16671 | 781.4531
66 | 16673 | 781.5469
66 | 16671 | 781.4531
66 | 16664 | 781.125
66 | 16661 | 780.9844
66 | 16666 | 781.2188
66 | 16669 | 781.3594
66 | 16672 | 781.5
66 | 16668 | 781.3125
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16669 | 781.3594
66 | 16674 | 781.5938
66 | 16675 | 781.6406
66 | 16672 | 781.5
66 | 16667 | 781.2656
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16665 | 781.1719
66 | 16668 | 781.3125
66 | 16671 | 781.4531
66 | 16672 | 781.5
66 | 16672 | 781.5
66 | 16667 | 781.2656
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16668 | 781.3125
66 | 16671 | 781.4531
66 | 16675 | 781.6406
66 | 16674 | 781.5938
66 | 16672 | 781.5
66 | 16667 | 781.2656
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16665 | 781.1719
66 | 16666 | 781.2188
66 | 16670 | 781.4062
66 | 16673 | 781.5469
66 | 16673 | 781.5469
66 | 16672 | 781.5
66 | 16669 | 781.3594
66 | 16669 | 781.3594
66 | 16666 | 781.2188
66 | 16667 | 781.2656
66 | 16667 | 781.2656
66 | 16663 | 781.0781
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16664 | 781.125
66 | 16667 | 781.2656
66 | 16667 | 781.2656
66 | 16669 | 781.3594
66 | 16671 | 781.4531
66 | 16674 | 781.5938
66 | 16674 | 781.5938
66 | 16673 | 781.5469
66 | 16669 | 781.3594
66 | 16665 | 781.1719
66 | 16666 | 781.2188
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16665 | 781.1719
66 | 16668 | 781.3125
66 | 16670 | 781.4062
66 | 16673 | 781.5469
66 | 16672 | 781.5
66 | 16670 | 781.4062
66 | 16668 | 781.3125
66 | 16667 | 781.2656
66 | 16668 | 781.3125
66 | 16664 | 781.125
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16666 | 781.2188
66 | 16666 | 781.2188
66 | 16668 | 781.3125
66 | 16670 | 781.4062
66 | 16671 | 781.4531
66 | 16672 | 781.5
66 | 16669 | 781.3594
66 | 16668 | 781.3125
66 | 16666 | 781.2188
66 | 16666 | 781.2188
66 | 16665 | 781.1719
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16668 | 781.3125
66 | 16671 | 781.4531
66 | 16672 | 781.5
66 | 16672 | 781.5
66 | 16669 | 781.3594
66 | 16667 | 781.2656
66 | 16666 | 781.2188
66 | 16665 | 781.1719
66 | 16665 | 781.1719
66 | 16664 | 781.125
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16664 | 781.125
66 | 16664 | 781.125
66 | 16666 | 781.2188
66 | 16668 | 781.3125
66 | 16671 | 781.4531
66 | 16670 | 781.4062
66 | 16670 | 781.4062
66 | 16667 | 781.2656
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16662 | 781.0312
66 | 16665 | 781.1719
66 | 16667 | 781.2656
66 | 16670 | 781.4062
66 | 16672 | 781.5
66 | 16669 | 781.3594
66 | 16666 | 781.2188
66 | 16665 | 781.1719
66 | 16663 | 781.0781
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16662 | 781.0312
66 | 16665 | 781.1719
66 | 16664 | 781.125
66 | 16668 | 781.3125
66 | 16669 | 781.3594
66 | 16670 | 781.4062
66 | 16668 | 781.3125
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16666 | 781.2188
66 | 16667 | 781.2656
66 | 16670 | 781.4062
66 | 16669 | 781.3594
66 | 16671 | 781.4531
66 | 16668 | 781.3125
66 | 16665 | 781.1719
66 | 16664 | 781.125
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16664 | 781.125
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16664 | 781.125
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16657 | 780.7969
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16664 | 781.125
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16661 | 780.9844
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16661 | 780.9844
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16663 | 781.0781
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16662 | 781.0312
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16657 | 780.7969
66 | 16659 | 780.8906
66 | 16657 | 780.7969
66 | 16657 | 780.7969
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16657 | 780.7969
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16662 | 781.0312
66 | 16661 | 780.9844
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16661 | 780.9844
66 | 16660 | 780.9375
66 | 16658 | 780.8438
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16660 | 780.9375
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16657 | 780.7969
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16660 | 780.9375
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16656 | 780.75
66 | 16655 | 780.7031
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16659 | 780.8906
66 | 16659 | 780.8906
66 | 16658 | 780.8438
66 | 16658 | 780.8438
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16651 | 780.5156
66 | 16650 | 780.4688
66 | 16649 | 780.4219
66 | 16649 | 780.4219
66 | 16647 | 780.3281
66 | 16647 | 780.3281
66 | 16646 | 780.2812
66 | 16653 | 780.6094
66 | 16655 | 780.7031
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16655 | 780.7031
66 | 16656 | 780.75
66 | 16654 | 780.6562
66 | 16653 | 780.6094
66 | 16649 | 780.4219
66 | 16649 | 780.4219
66 | 16648 | 780.375
66 | 16647 | 780.3281
66 | 16647 | 780.3281
66 | 16646 | 780.2812
66 | 16647 | 780.3281
66 | 16655 | 780.7031
66 | 16655 | 780.7031
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16653 | 780.6094
66 | 16656 | 780.75
66 | 16655 | 780.7031
66 | 16654 | 780.6562
66 | 16650 | 780.4688
66 | 16647 | 780.3281
66 | 16646 | 780.2812
66 | 16647 | 780.3281
66 | 16646 | 780.2812
66 | 16652 | 780.5625
66 | 16655 | 780.7031
66 | 16655 | 780.7031
66 | 16656 | 780.75
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16652 | 780.5625
66 | 16649 | 780.4219
66 | 16645 | 780.2344
66 | 16644 | 780.1875
66 | 16644 | 780.1875
66 | 16646 | 780.2812
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16658 | 780.8438
66 | 16657 | 780.7969
66 | 16655 | 780.7031
66 | 16652 | 780.5625
66 | 16647 | 780.3281
66 | 16645 | 780.2344
66 | 16644 | 780.1875
66 | 16643 | 780.1406
66 | 16644 | 780.1875
66 | 16648 | 780.375
66 | 16655 | 780.7031
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16658 | 780.8438
66 | 16656 | 780.75
66 | 16652 | 780.5625
66 | 16647 | 780.3281
66 | 16646 | 780.2812
66 | 16646 | 780.2812
66 | 16654 | 780.6562
66 | 16655 | 780.7031
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16652 | 780.5625
66 | 16649 | 780.4219
66 | 16646 | 780.2812
66 | 16645 | 780.2344
66 | 16645 | 780.2344
66 | 16645 | 780.2344
66 | 16650 | 780.4688
66 | 16656 | 780.75
66 | 16656 | 780.75
66 | 16657 | 780.7969
66 | 16656 | 780.75
66 | 16654 | 780.6562
66 | 16648 | 780.375
66 | 16644 | 780.1875
66 | 16643 | 780.1406
66 | 16642 | 780.0938
66 | 16650 | 780.4688
66 | 16654 | 780.6562
66 | 16653 | 780.6094
66 | 16656 | 780.75
66 | 16654 | 780.6562
66 | 16649 | 780.4219
66 | 16645 | 780.2344
66 | 16644 | 780.1875
66 | 16645 | 780.2344
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16654 | 780.6562
66 | 16650 | 780.4688
66 | 16645 | 780.2344
66 | 16642 | 780.0938
66 | 16642 | 780.0938
66 | 16645 | 780.2344
66 | 16653 | 780.6094
66 | 16654 | 780.6562
66 | 16655 | 780.7031
66 | 16655 | 780.7031
66 | 16653 | 780.6094
66 | 16650 | 780.4688
66 | 16645 | 780.2344
66 | 16643 | 780.1406
66 | 16646 | 780.2812
66 | 16653 | 780.6094
66 | 16653 | 780.6094
66 | 16651 | 780.5156
66 | 16651 | 780.5156
66 | 16644 | 780.1875
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16653 | 780.6094
66 | 16653 | 780.6094
66 | 16654 | 780.6562
66 | 16653 | 780.6094
66 | 16650 | 780.4688
66 | 16646 | 780.2812
66 | 16643 | 780.1406
66 | 16644 | 780.1875
66 | 16652 | 780.5625
66 | 16652 | 780.5625
66 | 16652 | 780.5625
66 | 16651 | 780.5156
66 | 16650 | 780.4688
66 | 16645 | 780.2344
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16642 | 780.0938
66 | 16652 | 780.5625
66 | 16653 | 780.6094
66 | 16654 | 780.6562
66 | 16652 | 780.5625
66 | 16647 | 780.3281
66 | 16644 | 780.1875
66 | 16642 | 780.0938
66 | 16646 | 780.2812
66 | 16652 | 780.5625
66 | 16652 | 780.5625
66 | 16650 | 780.4688
66 | 16644 | 780.1875
66 | 16640 | 780
66 | 16641 | 780.0469
66 | 16646 | 780.2812
66 | 16651 | 780.5156
66 | 16652 | 780.5625
66 | 16654 | 780.6562
66 | 16650 | 780.4688
66 | 16645 | 780.2344
66 | 16643 | 780.1406
66 | 16648 | 780.375
66 | 16652 | 780.5625
66 | 16651 | 780.5156
66 | 16648 | 780.375
66 | 16641 | 780.0469
66 | 16640 | 780
66 | 16650 | 780.4688
66 | 16651 | 780.5156
66 | 16648 | 780.375
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16643 | 780.1406
66 | 16652 | 780.5625
66 | 16653 | 780.6094
66 | 16650 | 780.4688
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16651 | 780.5156
66 | 16652 | 780.5625
66 | 16647 | 780.3281
66 | 16641 | 780.0469
66 | 16638 | 779.9062
66 | 16646 | 780.2812
66 | 16650 | 780.4688
66 | 16651 | 780.5156
66 | 16643 | 780.1406
66 | 16640 | 780
66 | 16646 | 780.2812
66 | 16652 | 780.5625
66 | 16651 | 780.5156
66 | 16647 | 780.3281
66 | 16642 | 780.0938
66 | 16647 | 780.3281
66 | 16652 | 780.5625
66 | 16650 | 780.4688
66 | 16642 | 780.0938
66 | 16638 | 779.9062
66 | 16649 | 780.4219
66 | 16650 | 780.4688
66 | 16649 | 780.4219
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16650 | 780.4688
66 | 16651 | 780.5156
66 | 16650 | 780.4688
66 | 16643 | 780.1406
66 | 16641 | 780.0469
66 | 16651 | 780.5156
66 | 16651 | 780.5156
66 | 16646 | 780.2812
66 | 16640 | 780
66 | 16641 | 780.0469
66 | 16650 | 780.4688
66 | 16648 | 780.375
66 | 16646 | 780.2812
66 | 16640 | 780
66 | 16641 | 780.0469
66 | 16650 | 780.4688
66 | 16649 | 780.4219
66 | 16641 | 780.0469
66 | 16639 | 779.9531
66 | 16646 | 780.2812
66 | 16651 | 780.5156
66 | 16651 | 780.5156
66 | 16642 | 780.0938
66 | 16639 | 779.9531
66 | 16648 | 780.375
66 | 16649 | 780.4219
66 | 16648 | 780.375
66 | 16640 | 780
66 | 16638 | 779.9062
66 | 16644 | 780.1875
66 | 16649 | 780.4219
66 | 16648 | 780.375
66 | 16642 | 780.0938
66 | 16640 | 780
66 | 16651 | 780.5156
66 | 16650 | 780.4688
66 | 16646 | 780.2812
66 | 16639 | 779.9531
66 | 16643 | 780.1406
66 | 16650 | 780.4688
66 | 16648 | 780.375
66 | 16639 | 779.9531
66 | 16636 | 779.8125
66 | 16643 | 780.1406
66 | 16649 | 780.4219
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16641 | 780.0469
66 | 16649 | 780.4219
66 | 16649 | 780.4219
66 | 16642 | 780.0938
66 | 16638 | 779.9062
66 | 16649 | 780.4219
66 | 16649 | 780.4219
66 | 16644 | 780.1875
66 | 16638 | 779.9062
66 | 16642 | 780.0938
66 | 16649 | 780.4219
66 | 16646 | 780.2812
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16647 | 780.3281
66 | 16648 | 780.375
66 | 16645 | 780.2344
66 | 16638 | 779.9062
66 | 16642 | 780.0938
66 | 16649 | 780.4219
66 | 16649 | 780.4219
66 | 16641 | 780.0469
66 | 16640 | 780
66 | 16650 | 780.4688
66 | 16647 | 780.3281
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16648 | 780.375
66 | 16646 | 780.2812
66 | 16640 | 780
66 | 16638 | 779.9062
66 | 16648 | 780.375
66 | 16648 | 780.375
66 | 16645 | 780.2344
66 | 16638 | 779.9062
66 | 16648 | 780.375
66 | 16650 | 780.4688
66 | 16647 | 780.3281
66 | 16638 | 779.9062
66 | 16639 | 779.9531
66 | 16648 | 780.375
66 | 16644 | 780.1875
66 | 16636 | 779.8125
66 | 16641 | 780.0469
66 | 16648 | 780.375
66 | 16643 | 780.1406
66 | 16636 | 779.8125
66 | 16641 | 780.0469
66 | 16648 | 780.375
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16645 | 780.2344
66 | 16649 | 780.4219
66 | 16645 | 780.2344
66 | 16637 | 779.8594
66 | 16640 | 780
66 | 16647 | 780.3281
66 | 16645 | 780.2344
66 | 16636 | 779.8125
66 | 16639 | 779.9531
66 | 16647 | 780.3281
66 | 16644 | 780.1875
66 | 16636 | 779.8125
66 | 16641 | 780.0469
66 | 16648 | 780.375
66 | 16645 | 780.2344
66 | 16638 | 779.9062
66 | 16645 | 780.2344
66 | 16649 | 780.4219
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16638 | 779.9062
66 | 16646 | 780.2812
66 | 16646 | 780.2812
66 | 16635 | 779.7656
66 | 16636 | 779.8125
66 | 16645 | 780.2344
66 | 16644 | 780.1875
66 | 16635 | 779.7656
66 | 16642 | 780.0938
66 | 16647 | 780.3281
66 | 16641 | 780.0469
66 | 16638 | 779.9062
66 | 16648 | 780.375
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16641 | 780.0469
66 | 16648 | 780.375
66 | 16642 | 780.0938
66 | 16635 | 779.7656
66 | 16645 | 780.2344
66 | 16644 | 780.1875
66 | 16635 | 779.7656
66 | 16642 | 780.0938
66 | 16646 | 780.2812
66 | 16642 | 780.0938
66 | 16637 | 779.8594
66 | 16647 | 780.3281
66 | 16645 | 780.2344
66 | 16637 | 779.8594
66 | 16643 | 780.1406
66 | 16647 | 780.3281
66 | 16641 | 780.0469
66 | 16636 | 779.8125
66 | 16646 | 780.2812
66 | 16644 | 780.1875
66 | 16635 | 779.7656
66 | 16640 | 780
66 | 16644 | 780.1875
66 | 16640 | 780
66 | 16636 | 779.8125
66 | 16646 | 780.2812
66 | 16641 | 780.0469
66 | 16636 | 779.8125
66 | 16648 | 780.375
66 | 16646 | 780.2812
66 | 16635 | 779.7656
66 | 16645 | 780.2344
66 | 16645 | 780.2344
66 | 16635 | 779.7656
66 | 16640 | 780
66 | 16645 | 780.2344
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16635 | 779.7656
66 | 16645 | 780.2344
66 | 16639 | 779.9531
66 | 16638 | 779.9062
66 | 16646 | 780.2812
66 | 16639 | 779.9531
66 | 16642 | 780.0938
66 | 16647 | 780.3281
66 | 16638 | 779.9062
66 | 16639 | 779.9531
66 | 16646 | 780.2812
66 | 16637 | 779.8594
66 | 16636 | 779.8125
66 | 16645 | 780.2344
66 | 16636 | 779.8125
66 | 16636 | 779.8125
66 | 16645 | 780.2344
66 | 16638 | 779.9062
66 | 16636 | 779.8125
66 | 16646 | 780.2812
66 | 16638 | 779.9062
66 | 16636 | 779.8125
66 | 16646 | 780.2812
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16646 | 780.2812
66 | 16641 | 780.0469
66 | 16633 | 779.6719
66 | 16644 | 780.1875
66 | 16640 | 780
66 | 16633 | 779.6719
66 | 16644 | 780.1875
66 | 16641 | 780.0469
66 | 16633 | 779.6719
66 | 16645 | 780.2344
66 | 16645 | 780.2344
66 | 16635 | 779.7656
66 | 16644 | 780.1875
66 | 16645 | 780.2344
66 | 16637 | 779.8594
66 | 16637 | 779.8594
66 | 16645 | 780.2344
66 | 16635 | 779.7656
66 | 16638 | 779.9062
66 | 16642 | 780.0938
66 | 16635 | 779.7656
66 | 16636 | 779.8125
66 | 16644 | 780.1875
66 | 16634 | 779.7188
66 | 16635 | 779.7656
66 | 16645 | 780.2344
66 | 16637 | 779.8594
66 | 16640 | 780
66 | 16644 | 780.1875
66 | 16638 | 779.9062
66 | 16635 | 779.7656
66 | 16644 | 780.1875
66 | 16636 | 779.8125
66 | 16638 | 779.9062
66 | 16643 | 780.1406
66 | 16633 | 779.6719
66 | 16636 | 779.8125
66 | 16642 | 780.0938
66 | 16635 | 779.7656
66 | 16638 | 779.9062
66 | 16645 | 780.2344
66 | 16636 | 779.8125
66 | 16638 | 779.9062
66 | 16646 | 780.2812
66 | 16637 | 779.8594
66 | 16640 | 780
66 | 16646 | 780.2812
66 | 16635 | 779.7656
66 | 16639 | 779.9531
66 | 16643 | 780.1406
66 | 16633 | 779.6719
66 | 16641 | 780.0469
66 | 16640 | 780
66 | 16632 | 779.625
66 | 16644 | 780.1875
66 | 16641 | 780.0469
66 | 16633 | 779.6719
66 | 16644 | 780.1875
66 | 16636 | 779.8125
66 | 16638 | 779.9062
66 | 16645 | 780.2344
66 | 16636 | 779.8125
66 | 16642 | 780.0938
66 | 16645 | 780.2344
66 | 16634 | 779.7188
66 | 16643 | 780.1406
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16643 | 780.1406
66 | 16638 | 779.9062
66 | 16632 | 779.625
66 | 16643 | 780.1406
66 | 16635 | 779.7656
66 | 16639 | 779.9531
66 | 16644 | 780.1875
66 | 16634 | 779.7188
66 | 16643 | 780.1406
66 | 16643 | 780.1406
66 | 16633 | 779.6719
66 | 16643 | 780.1406
66 | 16643 | 780.1406
66 | 16633 | 779.6719
66 | 16642 | 780.0938
66 | 16636 | 779.8125
66 | 16634 | 779.7188
66 | 16643 | 780.1406
66 | 16634 | 779.7188
66 | 16641 | 780.0469
66 | 16643 | 780.1406
66 | 16633 | 779.6719
66 | 16643 | 780.1406
66 | 16636 | 779.8125
66 | 16639 | 779.9531
66 | 16644 | 780.1875
66 | 16634 | 779.7188
66 | 16641 | 780.0469
66 | 16643 | 780.1406
66 | 16632 | 779.625
66 | 16643 | 780.1406
66 | 16633 | 779.6719
66 | 16638 | 779.9062
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16642 | 780.0938
66 | 16639 | 779.9531
66 | 16632 | 779.625
66 | 16642 | 780.0938
66 | 16635 | 779.7656
66 | 16641 | 780.0469
66 | 16642 | 780.0938
66 | 16634 | 779.7188
66 | 16643 | 780.1406
66 | 16636 | 779.8125
66 | 16638 | 779.9062
66 | 16641 | 780.0469
66 | 16630 | 779.5312
66 | 16639 | 779.9531
66 | 16638 | 779.9062
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16633 | 779.6719
66 | 16635 | 779.7656
66 | 16643 | 780.1406
66 | 16632 | 779.625
66 | 16642 | 780.0938
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16644 | 780.1875
66 | 16636 | 779.8125
66 | 16636 | 779.8125
66 | 16641 | 780.0469
66 | 16630 | 779.5312
66 | 16641 | 780.0469
66 | 16630 | 779.5312
66 | 16641 | 780.0469
66 | 16638 | 779.9062
66 | 16631 | 779.5781
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16642 | 780.0938
66 | 16638 | 779.9062
66 | 16635 | 779.7656
66 | 16643 | 780.1406
66 | 16630 | 779.5312
66 | 16641 | 780.0469
66 | 16637 | 779.8594
66 | 16632 | 779.625
66 | 16640 | 780
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16634 | 779.7188
66 | 16634 | 779.7188
66 | 16640 | 780
66 | 16629 | 779.4844
66 | 16642 | 780.0938
66 | 16633 | 779.6719
66 | 16640 | 780
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16641 | 780.0469
66 | 16631 | 779.5781
66 | 16641 | 780.0469
66 | 16630 | 779.5312
66 | 16638 | 779.9062
66 | 16637 | 779.8594
66 | 16631 | 779.5781
66 | 16638 | 779.9062
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16634 | 779.7188
66 | 16637 | 779.8594
66 | 16641 | 780.0469
66 | 16630 | 779.5312
66 | 16642 | 780.0938
66 | 16633 | 779.6719
66 | 16637 | 779.8594
66 | 16639 | 779.9531
66 | 16629 | 779.4844
66 | 16639 | 779.9531
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16638 | 779.9062
66 | 16629 | 779.4844
66 | 16640 | 780
66 | 16630 | 779.5312
66 | 16639 | 779.9531
66 | 16635 | 779.7656
66 | 16633 | 779.6719
66 | 16640 | 780
66 | 16631 | 779.5781
66 | 16641 | 780.0469
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16633 | 779.6719
66 | 16635 | 779.7656
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16638 | 779.9062
66 | 16628 | 779.4375
66 | 16639 | 779.9531
66 | 16630 | 779.5312
66 | 16638 | 779.9062
66 | 16634 | 779.7188
66 | 16635 | 779.7656
66 | 16640 | 780
66 | 16632 | 779.625
66 | 16640 | 780
66 | 16631 | 779.5781
66 | 16641 | 780.0469
66 | 16637 | 779.8594
66 | 16631 | 779.5781
66 | 16639 | 779.9531
66 | 16626 | 779.3438
66 | 16638 | 779.9062
66 | 16628 | 779.4375
66 | 16637 | 779.8594
66 | 16635 | 779.7656
66 | 16631 | 779.5781
66 | 16638 | 779.9062
66 | 16632 | 779.625
66 | 16639 | 779.9531
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16629 | 779.4844
66 | 16639 | 779.9531
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16637 | 779.8594
66 | 16627 | 779.3906
66 | 16638 | 779.9062
66 | 16628 | 779.4375
66 | 16635 | 779.7656
66 | 16634 | 779.7188
66 | 16628 | 779.4375
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16640 | 780
66 | 16630 | 779.5312
66 | 16640 | 780
66 | 16635 | 779.7656
66 | 16631 | 779.5781
66 | 16638 | 779.9062
66 | 16627 | 779.3906
66 | 16637 | 779.8594
66 | 16628 | 779.4375
66 | 16634 | 779.7188
66 | 16634 | 779.7188
66 | 16628 | 779.4375
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16638 | 779.9062
66 | 16629 | 779.4844
66 | 16640 | 780
66 | 16628 | 779.4375
66 | 16638 | 779.9062
66 | 16629 | 779.4844
66 | 16636 | 779.8125
66 | 16635 | 779.7656
66 | 16629 | 779.4844
66 | 16637 | 779.8594
66 | 16627 | 779.3906
66 | 16638 | 779.9062
66 | 16627 | 779.3906
66 | 16638 | 779.9062
66 | 16631 | 779.5781
66 | 16632 | 779.625
66 | 16637 | 779.8594
66 | 16628 | 779.4375
66 | 16637 | 779.8594
66 | 16628 | 779.4375
66 | 16639 | 779.9531
66 | 16627 | 779.3906
66 | 16636 | 779.8125
66 | 16632 | 779.625
66 | 16631 | 779.5781
66 | 16635 | 779.7656
66 | 16627 | 779.3906
66 | 16637 | 779.8594
66 | 16628 | 779.4375
66 | 16638 | 779.9062
66 | 16631 | 779.5781
66 | 16638 | 779.9062
66 | 16630 | 779.5312
66 | 16638 | 779.9062
66 | 16628 | 779.4375
66 | 16639 | 779.9531
66 | 16628 | 779.4375
66 | 16637 | 779.8594
66 | 16628 | 779.4375
66 | 16635 | 779.7656
66 | 16628 | 779.4375
66 | 16636 | 779.8125
66 | 16628 | 779.4375
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16637 | 779.8594
66 | 16630 | 779.5312
66 | 16637 | 779.8594
66 | 16630 | 779.5312
66 | 16636 | 779.8125
66 | 16630 | 779.5312
66 | 16635 | 779.7656
66 | 16629 | 779.4844
66 | 16634 | 779.7188
66 | 16627 | 779.3906
66 | 16634 | 779.7188
66 | 16627 | 779.3906
66 | 16636 | 779.8125
66 | 16628 | 779.4375
66 | 16636 | 779.8125
66 | 16628 | 779.4375
66 | 16638 | 779.9062
66 | 16629 | 779.4844
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16638 | 779.9062
66 | 16629 | 779.4844
66 | 16636 | 779.8125
66 | 16629 | 779.4844
66 | 16634 | 779.7188
66 | 16627 | 779.3906
66 | 16635 | 779.7656
66 | 16628 | 779.4375
66 | 16634 | 779.7188
66 | 16627 | 779.3906
66 | 16635 | 779.7656
66 | 16630 | 779.5312
66 | 16636 | 779.8125
66 | 16633 | 779.6719
66 | 16636 | 779.8125
66 | 16633 | 779.6719
66 | 16634 | 779.7188
66 | 16632 | 779.625
66 | 16634 | 779.7188
66 | 16632 | 779.625
66 | 16632 | 779.625
66 | 16631 | 779.5781
66 | 16632 | 779.625
66 | 16632 | 779.625
66 | 16634 | 779.7188
66 | 16634 | 779.7188
66 | 16633 | 779.6719
66 | 16635 | 779.7656
66 | 16635 | 779.7656
66 | 16635 | 779.7656
66 | 16634 | 779.7188
66 | 16634 | 779.7188
66 | 16634 | 779.7188
66 | 16633 | 779.6719
66 | 16633 | 779.6719
66 | 16632 | 779.625
66 | 16631 | 779.5781
66 | 16633 | 779.6719
66 | 16631 | 779.5781
66 | 16631 | 779.5781
66 | 16632 | 779.625
66 | 16635 | 779.7656
66 | 16631 | 779.5781
66 | 16635 | 779.7656
66 | 16633 | 779.6719
66 | 16635 | 779.7656
66 | 16632 | 779.625
66 | 16635 | 779.7656
66 | 16631 | 779.5781
66 | 16634 | 779.7188
66 | 16630 | 779.5312
66 | 16634 | 779.7188
66 | 16628 | 779.4375
66 | 16634 | 779.7188
66 | 16629 | 779.4844
66 | 16635 | 779.7656
66 | 16629 | 779.4844
66 | 16635 | 779.7656
66 | 16630 | 779.5312
66 | 16637 | 779.8594
66 | 16630 | 779.5312
66 | 16637 | 779.8594
66 | 16632 | 779.625
66 | 16637 | 779.8594
66 | 16629 | 779.4844
66 | 16636 | 779.8125
66 | 16629 | 779.4844
66 | 16635 | 779.7656
66 | 16627 | 779.3906
66 | 16633 | 779.6719
66 | 16626 | 779.3438
66 | 16635 | 779.7656
66 | 16629 | 779.4844
66 | 16635 | 779.7656
66 | 16630 | 779.5312
66 | 16637 | 779.8594
66 | 16630 | 779.5312
66 | 16637 | 779.8594
66 | 16628 | 779.4375
66 | 16635 | 779.7656
66 | 16625 | 779.2969
66 | 16635 | 779.7656
66 | 16624 | 779.25
66 | 16633 | 779.6719
66 | 16623 | 779.2031
66 | 16632 | 779.625
66 | 16624 | 779.25
66 | 16634 | 779.7188
66 | 16624 | 779.25
66 | 16630 | 779.5312
66 | 16633 | 779.6719
66 | 16628 | 779.4375
66 | 16635 | 779.7656
66 | 16626 | 779.3438
66 | 16636 | 779.8125
66 | 16625 | 779.2969
66 | 16634 | 779.7188
66 | 16625 | 779.2969
66 | 16633 | 779.6719
66 | 16624 | 779.25
66 | 16632 | 779.625
66 | 16623 | 779.2031
66 | 16630 | 779.5312
66 | 16625 | 779.2969
66 | 16632 | 779.625
66 | 16625 | 779.2969
66 | 16634 | 779.7188
66 | 16626 | 779.3438
66 | 16633 | 779.6719
66 | 16627 | 779.3906
66 | 16634 | 779.7188
66 | 16627 | 779.3906
66 | 16633 | 779.6719
66 | 16625 | 779.2969
66 | 16633 | 779.6719
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16621 | 779.1094
66 | 16632 | 779.625
66 | 16621 | 779.1094
66 | 16631 | 779.5781
66 | 16630 | 779.5312
66 | 16629 | 779.4844
66 | 16635 | 779.7656
66 | 16626 | 779.3438
66 | 16635 | 779.7656
66 | 16626 | 779.3438
66 | 16634 | 779.7188
66 | 16622 | 779.1562
66 | 16632 | 779.625
66 | 16623 | 779.2031
66 | 16626 | 779.3438
66 | 16629 | 779.4844
66 | 16623 | 779.2031
66 | 16630 | 779.5312
66 | 16621 | 779.1094
66 | 16632 | 779.625
66 | 16622 | 779.1562
66 | 16633 | 779.6719
66 | 16628 | 779.4375
66 | 16628 | 779.4375
66 | 16634 | 779.7188
66 | 16625 | 779.2969
66 | 16635 | 779.7656
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16627 | 779.3906
66 | 16627 | 779.3906
66 | 16632 | 779.625
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16623 | 779.2031
66 | 16631 | 779.5781
66 | 16622 | 779.1562
66 | 16633 | 779.6719
66 | 16623 | 779.2031
66 | 16634 | 779.7188
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16628 | 779.4375
66 | 16628 | 779.4375
66 | 16633 | 779.6719
66 | 16624 | 779.25
66 | 16632 | 779.625
66 | 16623 | 779.2031
66 | 16631 | 779.5781
66 | 16620 | 779.0625
66 | 16631 | 779.5781
66 | 16626 | 779.3438
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16624 | 779.25
66 | 16633 | 779.6719
66 | 16623 | 779.2031
66 | 16634 | 779.7188
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16625 | 779.2969
66 | 16629 | 779.4844
66 | 16633 | 779.6719
66 | 16624 | 779.25
66 | 16632 | 779.625
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16620 | 779.0625
66 | 16630 | 779.5312
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16624 | 779.25
66 | 16634 | 779.7188
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16631 | 779.5781
66 | 16626 | 779.3438
66 | 16632 | 779.625
66 | 16622 | 779.1562
66 | 16632 | 779.625
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16630 | 779.5312
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16620 | 779.0625
66 | 16632 | 779.625
66 | 16626 | 779.3438
66 | 16627 | 779.3906
66 | 16634 | 779.7188
66 | 16624 | 779.25
66 | 16634 | 779.7188
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16622 | 779.1562
66 | 16630 | 779.5312
66 | 16620 | 779.0625
66 | 16630 | 779.5312
66 | 16627 | 779.3906
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16623 | 779.2031
66 | 16626 | 779.3438
66 | 16631 | 779.5781
66 | 16623 | 779.2031
66 | 16633 | 779.6719
66 | 16620 | 779.0625
66 | 16631 | 779.5781
66 | 16628 | 779.4375
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16618 | 778.9688
66 | 16629 | 779.4844
66 | 16627 | 779.3906
66 | 16621 | 779.1094
66 | 16630 | 779.5312
66 | 16621 | 779.1094
66 | 16629 | 779.4844
66 | 16630 | 779.5312
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16621 | 779.1094
66 | 16630 | 779.5312
66 | 16626 | 779.3438
66 | 16621 | 779.1094
66 | 16628 | 779.4375
66 | 16618 | 778.9688
66 | 16627 | 779.3906
66 | 16628 | 779.4375
66 | 16622 | 779.1562
66 | 16629 | 779.4844
66 | 16619 | 779.0156
66 | 16628 | 779.4375
66 | 16629 | 779.4844
66 | 16621 | 779.1094
66 | 16631 | 779.5781
66 | 16623 | 779.2031
66 | 16624 | 779.25
66 | 16631 | 779.5781
66 | 16620 | 779.0625
66 | 16628 | 779.4375
66 | 16628 | 779.4375
66 | 16619 | 779.0156
66 | 16629 | 779.4844
66 | 16628 | 779.4375
66 | 16620 | 779.0625
66 | 16629 | 779.4844
66 | 16622 | 779.1562
66 | 16622 | 779.1562
66 | 16631 | 779.5781
66 | 16620 | 779.0625
66 | 16629 | 779.4844
66 | 16630 | 779.5312
66 | 16622 | 779.1562
66 | 16629 | 779.4844
66 | 16624 | 779.25
66 | 16621 | 779.1094
66 | 16629 | 779.4844
66 | 16618 | 778.9688
66 | 16626 | 779.3438
66 | 16629 | 779.4844
66 | 16619 | 779.0156
66 | 16626 | 779.3438
66 | 16625 | 779.2969
66 | 16619 | 779.0156
66 | 16630 | 779.5312
66 | 16619 | 779.0156
66 | 16626 | 779.3438
66 | 16631 | 779.5781
66 | 16621 | 779.1094
66 | 16629 | 779.4844
66 | 16628 | 779.4375
66 | 16619 | 779.0156
66 | 16631 | 779.5781
66 | 16629 | 779.4844
66 | 16621 | 779.1094
66 | 16630 | 779.5312
66 | 16622 | 779.1562
66 | 16620 | 779.0625
66 | 16629 | 779.4844
66 | 16618 | 778.9688
66 | 16627 | 779.3906
66 | 16629 | 779.4844
66 | 16621 | 779.1094
66 | 16630 | 779.5312
66 | 16627 | 779.3906
66 | 16621 | 779.1094
66 | 16630 | 779.5312
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16630 | 779.5312
66 | 16618 | 778.9688
66 | 16625 | 779.2969
66 | 16625 | 779.2969
66 | 16617 | 778.9219
66 | 16626 | 779.3438
66 | 16628 | 779.4375
66 | 16617 | 778.9219
66 | 16625 | 779.2969
66 | 16628 | 779.4375
66 | 16618 | 778.9688
66 | 16627 | 779.3906
66 | 16630 | 779.5312
66 | 16620 | 779.0625
66 | 16628 | 779.4375
66 | 16628 | 779.4375
66 | 16618 | 778.9688
66 | 16628 | 779.4375
66 | 16629 | 779.4844
66 | 16619 | 779.0156
66 | 16625 | 779.2969
66 | 16627 | 779.3906
66 | 16617 | 778.9219
66 | 16627 | 779.3906
66 | 16627 | 779.3906
66 | 16618 | 778.9688
66 | 16627 | 779.3906
66 | 16624 | 779.25
66 | 16618 | 778.9688
66 | 16628 | 779.4375
66 | 16626 | 779.3438
66 | 16622 | 779.1562
66 | 16629 | 779.4844
66 | 16625 | 779.2969
66 | 16619 | 779.0156
66 | 16629 | 779.4844
66 | 16625 | 779.2969
66 | 16619 | 779.0156
66 | 16627 | 779.3906
66 | 16622 | 779.1562
66 | 16616 | 778.875
66 | 16626 | 779.3438
66 | 16628 | 779.4375
66 | 16618 | 778.9688
66 | 16625 | 779.2969
66 | 16627 | 779.3906
66 | 16617 | 778.9219
66 | 16626 | 779.3438
66 | 16629 | 779.4844
66 | 16618 | 778.9688
66 | 16625 | 779.2969
66 | 16625 | 779.2969
66 | 16616 | 778.875
66 | 16625 | 779.2969
66 | 16625 | 779.2969
66 | 16616 | 778.875
66 | 16624 | 779.25
66 | 16623 | 779.2031
66 | 16616 | 778.875
66 | 16626 | 779.3438
66 | 16628 | 779.4375
66 | 16617 | 778.9219
66 | 16625 | 779.2969
66 | 16629 | 779.4844
66 | 16617 | 778.9219
66 | 16625 | 779.2969
66 | 16628 | 779.4375
66 | 16619 | 779.0156
66 | 16623 | 779.2031
66 | 16626 | 779.3438
66 | 16615 | 778.8281
66 | 16624 | 779.25
66 | 16626 | 779.3438
66 | 16615 | 778.8281
66 | 16620 | 779.0625
66 | 16625 | 779.2969
66 | 16615 | 778.8281
66 | 16623 | 779.2031
66 | 16627 | 779.3906
66 | 16616 | 778.875
66 | 16620 | 779.0625
66 | 16628 | 779.4375
66 | 16618 | 778.9688
66 | 16620 | 779.0625
66 | 16627 | 779.3906
66 | 16624 | 779.25
66 | 16616 | 778.875
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16613 | 778.7344
66 | 16620 | 779.0625
66 | 16625 | 779.2969
66 | 16615 | 778.8281
66 | 16619 | 779.0156
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16617 | 778.9219
66 | 16625 | 779.2969
66 | 16628 | 779.4375
66 | 16616 | 778.875
66 | 16625 | 779.2969
66 | 16627 | 779.3906
66 | 16616 | 778.875
66 | 16621 | 779.1094
66 | 16626 | 779.3438
66 | 16618 | 778.9688
66 | 16614 | 778.7812
66 | 16622 | 779.1562
66 | 16623 | 779.2031
66 | 16613 | 778.7344
66 | 16618 | 778.9688
66 | 16624 | 779.25
66 | 16615 | 778.8281
66 | 16620 | 779.0625
66 | 16626 | 779.3438
66 | 16624 | 779.25
66 | 16615 | 778.8281
66 | 16623 | 779.2031
66 | 16627 | 779.3906
66 | 16614 | 778.7812
66 | 16616 | 778.875
66 | 16624 | 779.25
66 | 16622 | 779.1562
66 | 16615 | 778.8281
66 | 16621 | 779.1094
66 | 16621 | 779.1094
66 | 16614 | 778.7812
66 | 16621 | 779.1094
66 | 16624 | 779.25
66 | 16614 | 778.7812
66 | 16618 | 778.9688
66 | 16628 | 779.4375
66 | 16625 | 779.2969
66 | 16616 | 778.875
66 | 16623 | 779.2031
66 | 16625 | 779.2969
66 | 16614 | 778.7812
66 | 16620 | 779.0625
66 | 16625 | 779.2969
66 | 16614 | 778.7812
66 | 16615 | 778.8281
66 | 16624 | 779.25
66 | 16624 | 779.25
66 | 16614 | 778.7812
66 | 16621 | 779.1094
66 | 16625 | 779.2969
66 | 16614 | 778.7812
66 | 16617 | 778.9219
66 | 16625 | 779.2969
66 | 16625 | 779.2969
66 | 16616 | 778.875
66 | 16622 | 779.1562
66 | 16625 | 779.2969
66 | 16619 | 779.0156
66 | 16614 | 778.7812
66 | 16620 | 779.0625
66 | 16624 | 779.25
66 | 16617 | 778.9219
66 | 16615 | 778.8281
66 | 16624 | 779.25
66 | 16624 | 779.25
66 | 16614 | 778.7812
66 | 16617 | 778.9219
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16616 | 778.875
66 | 16619 | 779.0156
66 | 16625 | 779.2969
66 | 16617 | 778.9219
66 | 16615 | 778.8281
66 | 16623 | 779.2031
66 | 16625 | 779.2969
66 | 16613 | 778.7344
66 | 16616 | 778.875
66 | 16621 | 779.1094
66 | 16623 | 779.2031
66 | 16613 | 778.7344
66 | 16615 | 778.8281
66 | 16622 | 779.1562
66 | 16623 | 779.2031
66 | 16614 | 778.7812
66 | 16619 | 779.0156
66 | 16625 | 779.2969
66 | 16623 | 779.2031
66 | 16614 | 778.7812
66 | 16620 | 779.0625
66 | 16624 | 779.25
66 | 16616 | 778.875
66 | 16613 | 778.7344
66 | 16619 | 779.0156
66 | 16621 | 779.1094
66 | 16616 | 778.875
66 | 16611 | 778.6406
66 | 16619 | 779.0156
66 | 16623 | 779.2031
66 | 16622 | 779.1562
66 | 16614 | 778.7812
66 | 16620 | 779.0625
66 | 16625 | 779.2969
66 | 16624 | 779.25
66 | 16615 | 778.8281
66 | 16619 | 779.0156
66 | 16624 | 779.25
66 | 16622 | 779.1562
66 | 16613 | 778.7344
66 | 16615 | 778.8281
66 | 16618 | 778.9688
66 | 16622 | 779.1562
66 | 16613 | 778.7344
66 | 16613 | 778.7344
66 | 16621 | 779.1094
66 | 16623 | 779.2031
66 | 16615 | 778.8281
66 | 16613 | 778.7344
66 | 16620 | 779.0625
66 | 16624 | 779.25
66 | 16617 | 778.9219
66 | 16615 | 778.8281
66 | 16621 | 779.1094
66 | 16623 | 779.2031
66 | 16615 | 778.8281
66 | 16613 | 778.7344
66 | 16618 | 778.9688
66 | 16622 | 779.1562
66 | 16614 | 778.7812
66 | 16612 | 778.6875
66 | 16616 | 778.875
66 | 16622 | 779.1562
66 | 16622 | 779.1562
66 | 16612 | 778.6875
66 | 16616 | 778.875
66 | 16623 | 779.2031
66 | 16623 | 779.2031
66 | 16613 | 778.7344
66 | 16614 | 778.7812
66 | 16623 | 779.2031
66 | 16622 | 779.1562
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16621 | 779.1094
66 | 16621 | 779.1094
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16618 | 778.9688
66 | 16622 | 779.1562
66 | 16622 | 779.1562
66 | 16614 | 778.7812
66 | 16614 | 778.7812
66 | 16618 | 778.9688
66 | 16623 | 779.2031
66 | 16620 | 779.0625
66 | 16614 | 778.7812
66 | 16617 | 778.9219
66 | 16622 | 779.1562
66 | 16618 | 778.9688
66 | 16610 | 778.5938
66 | 16614 | 778.7812
66 | 16621 | 779.1094
66 | 16619 | 779.0156
66 | 16611 | 778.6406
66 | 16615 | 778.8281
66 | 16621 | 779.1094
66 | 16615 | 778.8281
66 | 16614 | 778.7812
66 | 16617 | 778.9219
66 | 16622 | 779.1562
66 | 16621 | 779.1094
66 | 16611 | 778.6406
66 | 16615 | 778.8281
66 | 16621 | 779.1094
66 | 16619 | 779.0156
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16614 | 778.7812
66 | 16621 | 779.1094
66 | 16620 | 779.0625
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16616 | 778.875
66 | 16623 | 779.2031
66 | 16622 | 779.1562
66 | 16612 | 778.6875
66 | 16616 | 778.875
66 | 16623 | 779.2031
66 | 16621 | 779.1094
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16616 | 778.875
66 | 16620 | 779.0625
66 | 16619 | 779.0156
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16615 | 778.8281
66 | 16621 | 779.1094
66 | 16620 | 779.0625
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16617 | 778.9219
66 | 16621 | 779.1094
66 | 16621 | 779.1094
66 | 16613 | 778.7344
66 | 16608 | 778.5
66 | 16612 | 778.6875
66 | 16619 | 779.0156
66 | 16619 | 779.0156
66 | 16613 | 778.7344
66 | 16609 | 778.5469
66 | 16614 | 778.7812
66 | 16621 | 779.1094
66 | 16620 | 779.0625
66 | 16614 | 778.7812
66 | 16612 | 778.6875
66 | 16616 | 778.875
66 | 16622 | 779.1562
66 | 16621 | 779.1094
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16614 | 778.7812
66 | 16619 | 779.0156
66 | 16620 | 779.0625
66 | 16614 | 778.7812
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16619 | 779.0156
66 | 16621 | 779.1094
66 | 16614 | 778.7812
66 | 16610 | 778.5938
66 | 16615 | 778.8281
66 | 16619 | 779.0156
66 | 16620 | 779.0625
66 | 16616 | 778.875
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16619 | 779.0156
66 | 16620 | 779.0625
66 | 16617 | 778.9219
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16612 | 778.6875
66 | 16620 | 779.0625
66 | 16621 | 779.1094
66 | 16615 | 778.8281
66 | 16610 | 778.5938
66 | 16615 | 778.8281
66 | 16617 | 778.9219
66 | 16621 | 779.1094
66 | 16621 | 779.1094
66 | 16614 | 778.7812
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16614 | 778.7812
66 | 16618 | 778.9688
66 | 16618 | 778.9688
66 | 16611 | 778.6406
66 | 16607 | 778.4531
66 | 16614 | 778.7812
66 | 16618 | 778.9688
66 | 16620 | 779.0625
66 | 16618 | 778.9688
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16614 | 778.7812
66 | 16619 | 779.0156
66 | 16621 | 779.1094
66 | 16619 | 779.0156
66 | 16611 | 778.6406
66 | 16607 | 778.4531
66 | 16611 | 778.6406
66 | 16614 | 778.7812
66 | 16619 | 779.0156
66 | 16618 | 778.9688
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16613 | 778.7344
66 | 16615 | 778.8281
66 | 16620 | 779.0625
66 | 16621 | 779.1094
66 | 16615 | 778.8281
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16613 | 778.7344
66 | 16618 | 778.9688
66 | 16617 | 778.9219
66 | 16614 | 778.7812
66 | 16611 | 778.6406
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16614 | 778.7812
66 | 16619 | 779.0156
66 | 16621 | 779.1094
66 | 16619 | 779.0156
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16619 | 779.0156
66 | 16619 | 779.0156
66 | 16618 | 778.9688
66 | 16615 | 778.8281
66 | 16611 | 778.6406
66 | 16606 | 778.4062
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16615 | 778.8281
66 | 16619 | 779.0156
66 | 16619 | 779.0156
66 | 16615 | 778.8281
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16615 | 778.8281
66 | 16618 | 778.9688
66 | 16618 | 778.9688
66 | 16617 | 778.9219
66 | 16609 | 778.5469
66 | 16607 | 778.4531
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16615 | 778.8281
66 | 16618 | 778.9688
66 | 16617 | 778.9219
66 | 16613 | 778.7344
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16613 | 778.7344
66 | 16617 | 778.9219
66 | 16619 | 779.0156
66 | 16619 | 779.0156
66 | 16616 | 778.875
66 | 16610 | 778.5938
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16611 | 778.6406
66 | 16614 | 778.7812
66 | 16615 | 778.8281
66 | 16616 | 778.875
66 | 16616 | 778.875
66 | 16612 | 778.6875
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16618 | 778.9688
66 | 16619 | 779.0156
66 | 16616 | 778.875
66 | 16611 | 778.6406
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16615 | 778.8281
66 | 16615 | 778.8281
66 | 16616 | 778.875
66 | 16614 | 778.7812
66 | 16611 | 778.6406
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16613 | 778.7344
66 | 16615 | 778.8281
66 | 16618 | 778.9688
66 | 16618 | 778.9688
66 | 16617 | 778.9219
66 | 16613 | 778.7344
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16615 | 778.8281
66 | 16618 | 778.9688
66 | 16618 | 778.9688
66 | 16616 | 778.875
66 | 16610 | 778.5938
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16616 | 778.875
66 | 16615 | 778.8281
66 | 16615 | 778.8281
66 | 16611 | 778.6406
66 | 16606 | 778.4062
66 | 16604 | 778.3125
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16614 | 778.7812
66 | 16616 | 778.875
66 | 16618 | 778.9688
66 | 16618 | 778.9688
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16606 | 778.4062
66 | 16604 | 778.3125
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16615 | 778.8281
66 | 16616 | 778.875
66 | 16617 | 778.9219
66 | 16617 | 778.9219
66 | 16612 | 778.6875
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16612 | 778.6875
66 | 16613 | 778.7344
66 | 16614 | 778.7812
66 | 16613 | 778.7344
66 | 16615 | 778.8281
66 | 16613 | 778.7344
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16610 | 778.5938
66 | 16613 | 778.7344
66 | 16615 | 778.8281
66 | 16616 | 778.875
66 | 16615 | 778.8281
66 | 16615 | 778.8281
66 | 16615 | 778.8281
66 | 16613 | 778.7344
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16603 | 778.2656
66 | 16604 | 778.3125
66 | 16603 | 778.2656
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16608 | 778.5
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16616 | 778.875
66 | 16616 | 778.875
66 | 16614 | 778.7812
66 | 16614 | 778.7812
66 | 16611 | 778.6406
66 | 16608 | 778.5
66 | 16601 | 778.1719
66 | 16602 | 778.2188
66 | 16604 | 778.3125
66 | 16603 | 778.2656
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16608 | 778.5
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16604 | 778.3125
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16603 | 778.2656
66 | 16605 | 778.3594
66 | 16603 | 778.2656
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16609 | 778.5469
66 | 16607 | 778.4531
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16613 | 778.7344
66 | 16612 | 778.6875
66 | 16613 | 778.7344
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16608 | 778.5
66 | 16606 | 778.4062
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16603 | 778.2656
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16603 | 778.2656
66 | 16603 | 778.2656
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16606 | 778.4062
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16613 | 778.7344
66 | 16614 | 778.7812
66 | 16613 | 778.7344
66 | 16612 | 778.6875
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16614 | 778.7812
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16611 | 778.6406
66 | 16608 | 778.5
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16613 | 778.7344
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16613 | 778.7344
66 | 16611 | 778.6406
66 | 16613 | 778.7344
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16613 | 778.7344
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16613 | 778.7344
66 | 16613 | 778.7344
66 | 16613 | 778.7344
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16610 | 778.5938
66 | 16613 | 778.7344
66 | 16609 | 778.5469
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16607 | 778.4531
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16609 | 778.5469
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16612 | 778.6875
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16610 | 778.5938
66 | 16611 | 778.6406
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16612 | 778.6875
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16607 | 778.4531
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16609 | 778.5469
66 | 16607 | 778.4531
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16611 | 778.6406
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16609 | 778.5469
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16606 | 778.4062
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16609 | 778.5469
66 | 16607 | 778.4531
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16610 | 778.5938
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16601 | 778.1719
66 | 16598 | 778.0312
66 | 16597 | 777.9844
66 | 16600 | 778.125
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16597 | 777.9844
66 | 16596 | 777.9375
66 | 16599 | 778.0781
66 | 16600 | 778.125
66 | 16607 | 778.4531
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16606 | 778.4062
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16597 | 777.9844
66 | 16597 | 777.9844
66 | 16600 | 778.125
66 | 16603 | 778.2656
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16603 | 778.2656
66 | 16601 | 778.1719
66 | 16596 | 777.9375
66 | 16594 | 777.8438
66 | 16597 | 777.9844
66 | 16602 | 778.2188
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16609 | 778.5469
66 | 16608 | 778.5
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16603 | 778.2656
66 | 16601 | 778.1719
66 | 16596 | 777.9375
66 | 16594 | 777.8438
66 | 16596 | 777.9375
66 | 16602 | 778.2188
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16604 | 778.3125
66 | 16600 | 778.125
66 | 16596 | 777.9375
66 | 16597 | 777.9844
66 | 16604 | 778.3125
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16608 | 778.5
66 | 16606 | 778.4062
66 | 16604 | 778.3125
66 | 16602 | 778.2188
66 | 16598 | 778.0312
66 | 16594 | 777.8438
66 | 16595 | 777.8906
66 | 16601 | 778.1719
66 | 16604 | 778.3125
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16596 | 777.9375
66 | 16596 | 777.9375
66 | 16603 | 778.2656
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16604 | 778.3125
66 | 16600 | 778.125
66 | 16599 | 778.0781
66 | 16594 | 777.8438
66 | 16594 | 777.8438
66 | 16598 | 778.0312
66 | 16604 | 778.3125
66 | 16606 | 778.4062
66 | 16607 | 778.4531
66 | 16607 | 778.4531
66 | 16608 | 778.5
66 | 16604 | 778.3125
66 | 16601 | 778.1719
66 | 16595 | 777.8906
66 | 16595 | 777.8906
66 | 16601 | 778.1719
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16600 | 778.125
66 | 16595 | 777.8906
66 | 16595 | 777.8906
66 | 16599 | 778.0781
66 | 16604 | 778.3125
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16604 | 778.3125
66 | 16601 | 778.1719
66 | 16599 | 778.0781
66 | 16594 | 777.8438
66 | 16599 | 778.0781
66 | 16604 | 778.3125
66 | 16605 | 778.3594
66 | 16603 | 778.2656
66 | 16601 | 778.1719
66 | 16594 | 777.8438
66 | 16593 | 777.7969
66 | 16597 | 777.9844
66 | 16602 | 778.2188
66 | 16606 | 778.4062
66 | 16606 | 778.4062
66 | 16602 | 778.2188
66 | 16598 | 778.0312
66 | 16595 | 777.8906
66 | 16602 | 778.2188
66 | 16606 | 778.4062
66 | 16605 | 778.3594
66 | 16603 | 778.2656
66 | 16599 | 778.0781
66 | 16593 | 777.7969
66 | 16592 | 777.75
66 | 16601 | 778.1719
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16597 | 777.9844
66 | 16593 | 777.7969
66 | 16598 | 778.0312
66 | 16604 | 778.3125
66 | 16605 | 778.3594
66 | 16604 | 778.3125
66 | 16598 | 778.0312
66 | 16595 | 777.8906
66 | 16596 | 777.9375
66 | 16604 | 778.3125
66 | 16603 | 778.2656
66 | 16604 | 778.3125
66 | 16599 | 778.0781
66 | 16597 | 777.9844
66 | 16592 | 777.75
66 | 16599 | 778.0781
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16601 | 778.1719
66 | 16599 | 778.0781
66 | 16595 | 777.8906
66 | 16604 | 778.3125
66 | 16605 | 778.3594
66 | 16603 | 778.2656
66 | 16599 | 778.0781
66 | 16595 | 777.8906
66 | 16593 | 777.7969
66 | 16599 | 778.0781
66 | 16602 | 778.2188
66 | 16602 | 778.2188
66 | 16599 | 778.0781
66 | 16595 | 777.8906
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16604 | 778.3125
66 | 16602 | 778.2188
66 | 16599 | 778.0781
66 | 16594 | 777.8438
66 | 16599 | 778.0781
66 | 16605 | 778.3594
66 | 16605 | 778.3594
66 | 16600 | 778.125
66 | 16593 | 777.7969
66 | 16594 | 777.8438
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16601 | 778.1719
66 | 16598 | 778.0312
66 | 16592 | 777.75
66 | 16599 | 778.0781
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16600 | 778.125
66 | 16593 | 777.7969
66 | 16597 | 777.9844
66 | 16604 | 778.3125
66 | 16604 | 778.3125
66 | 16602 | 778.2188
66 | 16596 | 777.9375
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16603 | 778.2656
66 | 16602 | 778.2188
66 | 16599 | 778.0781
66 | 16592 | 777.75
66 | 16595 | 777.8906
66 | 16602 | 778.2188
66 | 16603 | 778.2656
66 | 16598 | 778.0312
66 | 16593 | 777.7969
66 | 16596 | 777.9375
66 | 16603 | 778.2656
66 | 16604 | 778.3125
66 | 16599 | 778.0781
66 | 16593 | 777.7969
66 | 16599 | 778.0781
66 | 16603 | 778.2656
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16591 | 777.7031
66 | 16601 | 778.1719
66 | 16603 | 778.2656
66 | 16600 | 778.125
66 | 16593 | 777.7969
66 | 16593 | 777.7969
66 | 16601 | 778.1719
66 | 16604 | 778.3125
66 | 16601 | 778.1719
66 | 16597 | 777.9844
66 | 16593 | 777.7969
66 | 16601 | 778.1719
66 | 16601 | 778.1719
66 | 16600 | 778.125
66 | 16593 | 777.7969
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16600 | 778.125
66 | 16598 | 778.0312
66 | 16591 | 777.7031
66 | 16592 | 777.75
66 | 16602 | 778.2188
66 | 16604 | 778.3125
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16595 | 777.8906
66 | 16603 | 778.2656
66 | 16604 | 778.3125
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16592 | 777.75
66 | 16599 | 778.0781
66 | 16602 | 778.2188
66 | 16598 | 778.0312
66 | 16591 | 777.7031
66 | 16595 | 777.8906
66 | 16604 | 778.3125
66 | 16602 | 778.2188
66 | 16595 | 777.8906
66 | 16593 | 777.7969
66 | 16602 | 778.2188
66 | 16603 | 778.2656
66 | 16601 | 778.1719
66 | 16594 | 777.8438
66 | 16592 | 777.75
66 | 16599 | 778.0781
66 | 16601 | 778.1719
66 | 16599 | 778.0781
66 | 16591 | 777.7031
66 | 16589 | 777.6094
66 | 16600 | 778.125
66 | 16601 | 778.1719
66 | 16598 | 778.0312
66 | 16591 | 777.7031
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16602 | 778.2188
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16593 | 777.7969
66 | 16602 | 778.2188
66 | 16601 | 778.1719
66 | 16599 | 778.0781
66 | 16590 | 777.6562
66 | 16593 | 777.7969
66 | 16599 | 778.0781
66 | 16598 | 778.0312
66 | 16593 | 777.7969
66 | 16589 | 777.6094
66 | 16601 | 778.1719
66 | 16600 | 778.125
66 | 16598 | 778.0312
66 | 16592 | 777.75
66 | 16592 | 777.75
66 | 16601 | 778.1719
66 | 16603 | 778.2656
66 | 16599 | 778.0781
66 | 16592 | 777.75
66 | 16593 | 777.7969
66 | 16600 | 778.125
66 | 16599 | 778.0781
66 | 16596 | 777.9375
66 | 16589 | 777.6094
66 | 16591 | 777.7031
66 | 16599 | 778.0781
66 | 16599 | 778.0781
66 | 16596 | 777.9375
66 | 16590 | 777.6562
66 | 16596 | 777.9375
66 | 16601 | 778.1719
66 | 16598 | 778.0312
66 | 16592 | 777.75
66 | 16594 | 777.8438
66 | 16601 | 778.1719
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16590 | 777.6562
66 | 16598 | 778.0312
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16588 | 777.5625
66 | 16597 | 777.9844
66 | 16600 | 778.125
66 | 16594 | 777.8438
66 | 16589 | 777.6094
66 | 16598 | 778.0312
66 | 16601 | 778.1719
66 | 16596 | 777.9375
66 | 16591 | 777.7031
66 | 16601 | 778.1719
66 | 16601 | 778.1719
66 | 16594 | 777.8438
66 | 16589 | 777.6094
66 | 16597 | 777.9844
66 | 16599 | 778.0781
66 | 16593 | 777.7969
66 | 16587 | 777.5156
66 | 16598 | 778.0312
66 | 16596 | 777.9375
66 | 16593 | 777.7969
66 | 16590 | 777.6562
66 | 16601 | 778.1719
66 | 16601 | 778.1719
66 | 16594 | 777.8438
66 | 16593 | 777.7969
66 | 16602 | 778.2188
66 | 16594 | 777.8438
66 | 16589 | 777.6094
66 | 16600 | 778.125
66 | 16600 | 778.125
66 | 16593 | 777.7969
66 | 16589 | 777.6094
66 | 16599 | 778.0781
66 | 16597 | 777.9844
66 | 16589 | 777.6094
66 | 16595 | 777.8906
66 | 16601 | 778.1719
66 | 16594 | 777.8438
66 | 16590 | 777.6562
66 | 16599 | 778.0781
66 | 16600 | 778.125
66 | 16592 | 777.75
66 | 16590 | 777.6562
66 | 16599 | 778.0781
66 | 16597 | 777.9844
66 | 16589 | 777.6094
66 | 16589 | 777.6094
66 | 16598 | 778.0312
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16596 | 777.9375
66 | 16599 | 778.0781
66 | 16594 | 777.8438
66 | 16589 | 777.6094
66 | 16599 | 778.0781
66 | 16599 | 778.0781
66 | 16593 | 777.7969
66 | 16590 | 777.6562
66 | 16599 | 778.0781
66 | 16598 | 778.0312
66 | 16590 | 777.6562
66 | 16591 | 777.7031
66 | 16598 | 778.0312
66 | 16593 | 777.7969
66 | 16586 | 777.4688
66 | 16596 | 777.9375
66 | 16596 | 777.9375
66 | 16589 | 777.6094
66 | 16592 | 777.75
66 | 16600 | 778.125
66 | 16595 | 777.8906
66 | 16588 | 777.5625
66 | 16600 | 778.125
66 | 16599 | 778.0781
66 | 16591 | 777.7031
66 | 16589 | 777.6094
66 | 16599 | 778.0781
66 | 16596 | 777.9375
66 | 16587 | 777.5156
66 | 16591 | 777.7031
66 | 16597 | 777.9844
66 | 16591 | 777.7031
66 | 16589 | 777.6094
66 | 16597 | 777.9844
66 | 16594 | 777.8438
66 | 16587 | 777.5156
66 | 16599 | 778.0781
66 | 16598 | 778.0312
66 | 16590 | 777.6562
66 | 16596 | 777.9375
66 | 16597 | 777.9844
66 | 16589 | 777.6094
66 | 16592 | 777.75
66 | 16595 | 777.8906
66 | 16591 | 777.7031
66 | 16587 | 777.5156
66 | 16596 | 777.9375
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16594 | 777.8438
66 | 16596 | 777.9375
66 | 16592 | 777.75
66 | 16589 | 777.6094
66 | 16598 | 778.0312
66 | 16594 | 777.8438
66 | 16589 | 777.6094
66 | 16598 | 778.0312
66 | 16596 | 777.9375
66 | 16587 | 777.5156
66 | 16592 | 777.75
66 | 16597 | 777.9844
66 | 16590 | 777.6562
66 | 16587 | 777.5156
66 | 16596 | 777.9375
66 | 16592 | 777.75
66 | 16588 | 777.5625
66 | 16599 | 778.0781
66 | 16595 | 777.8906
66 | 16587 | 777.5156
66 | 16594 | 777.8438
66 | 16597 | 777.9844
66 | 16591 | 777.7031
66 | 16589 | 777.6094
66 | 16597 | 777.9844
66 | 16591 | 777.7031
66 | 16588 | 777.5625
66 | 16596 | 777.9375
66 | 16593 | 777.7969
66 | 16584 | 777.375
66 | 16595 | 777.8906
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16594 | 777.8438
66 | 16596 | 777.9375
66 | 16589 | 777.6094
66 | 16592 | 777.75
66 | 16596 | 777.9375
66 | 16589 | 777.6094
66 | 16589 | 777.6094
66 | 16596 | 777.9375
66 | 16591 | 777.7031
66 | 16587 | 777.5156
66 | 16594 | 777.8438
66 | 16589 | 777.6094
66 | 16583 | 777.3281
66 | 16594 | 777.8438
66 | 16592 | 777.75
66 | 16584 | 777.375
66 | 16596 | 777.9375
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16594 | 777.8438
66 | 16596 | 777.9375
66 | 16588 | 777.5625
66 | 16593 | 777.7969
66 | 16595 | 777.8906
66 | 16587 | 777.5156
66 | 16591 | 777.7031
66 | 16594 | 777.8438
66 | 16584 | 777.375
66 | 16593 | 777.7969
66 | 16594 | 777.8438
66 | 16583 | 777.3281
66 | 16591 | 777.7031
66 | 16594 | 777.8438
66 | 16584 | 777.375
66 | 16593 | 777.7969
66 | 16595 | 777.8906
66 | 16586 | 777.4688
66 | 16593 | 777.7969
66 | 16595 | 777.8906
66 | 16586 | 777.4688
66 | 16593 | 777.7969
66 | 16593 | 777.7969
66 | 16584 | 777.375
66 | 16591 | 777.7031
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16589 | 777.6094
66 | 16592 | 777.75
66 | 16587 | 777.5156
66 | 16587 | 777.5156
66 | 16594 | 777.8438
66 | 16587 | 777.5156
66 | 16593 | 777.7969
66 | 16595 | 777.8906
66 | 16587 | 777.5156
66 | 16595 | 777.8906
66 | 16594 | 777.8438
66 | 16585 | 777.4219
66 | 16593 | 777.7969
66 | 16593 | 777.7969
66 | 16582 | 777.2812
66 | 16591 | 777.7031
66 | 16590 | 777.6562
66 | 16584 | 777.375
66 | 16594 | 777.8438
66 | 16587 | 777.5156
66 | 16593 | 777.7969
66 | 16593 | 777.7969
66 | 16585 | 777.4219
66 | 16595 | 777.8906
66 | 16591 | 777.7031
66 | 16589 | 777.6094
66 | 16594 | 777.8438
66 | 16585 | 777.4219
66 | 16595 | 777.8906
66 | 16591 | 777.7031
66 | 16585 | 777.4219
66 | 16593 | 777.7969
66 | 16582 | 777.2812
66 | 16592 | 777.75
66 | 16588 | 777.5625
66 | 16590 | 777.6562
66 | 16592 | 777.75
66 | 16584 | 777.375
66 | 16594 | 777.8438
66 | 16592 | 777.75
66 | 16585 | 777.4219
66 | 16595 | 777.8906
66 | 16587 | 777.5156
66 | 16592 | 777.75
66 | 16590 | 777.6562
66 | 16586 | 777.4688
66 | 16593 | 777.7969
66 | 16583 | 777.3281
66 | 16592 | 777.75
66 | 16589 | 777.6094
66 | 16585 | 777.4219
66 | 16591 | 777.7031
66 | 16584 | 777.375
66 | 16593 | 777.7969
66 | 16590 | 777.6562
66 | 16592 | 777.75
66 | 16593 | 777.7969
66 | 16586 | 777.4688
66 | 16595 | 777.8906
66 | 16584 | 777.375
66 | 16595 | 777.8906
66 | 16590 | 777.6562
66 | 16586 | 777.4688
66 | 16592 | 777.75
66 | 16582 | 777.2812
66 | 16592 | 777.75
66 | 16589 | 777.6094
66 | 16583 | 777.3281
66 | 16594 | 777.8438
66 | 16587 | 777.5156
66 | 16591 | 777.7031
66 | 16594 | 777.8438
66 | 16583 | 777.3281
66 | 16594 | 777.8438
66 | 16591 | 777.7031
66 | 16585 | 777.4219
66 | 16594 | 777.8438
66 | 16585 | 777.4219
66 | 16587 | 777.5156
66 | 16591 | 777.7031
66 | 16581 | 777.2344
66 | 16590 | 777.6562
66 | 16591 | 777.7031
66 | 16585 | 777.4219
66 | 16594 | 777.8438
66 | 16585 | 777.4219
66 | 16592 | 777.75
66 | 16592 | 777.75
66 | 16584 | 777.375
66 | 16592 | 777.75
66 | 16591 | 777.7031
66 | 16585 | 777.4219
66 | 16592 | 777.75
66 | 16583 | 777.3281
66 | 16592 | 777.75
66 | 16591 | 777.7031
66 | 16580 | 777.1875
66 | 16592 | 777.75
66 | 16585 | 777.4219
66 | 16585 | 777.4219
66 | 16593 | 777.7969
66 | 16583 | 777.3281
66 | 16592 | 777.75
66 | 16591 | 777.7031
66 | 16583 | 777.3281
66 | 16594 | 777.8438
66 | 16586 | 777.4688
66 | 16590 | 777.6562
66 | 16592 | 777.75
66 | 16581 | 777.2344
66 | 16592 | 777.75
66 | 16587 | 777.5156
66 | 16583 | 777.3281
66 | 16589 | 777.6094
66 | 16581 | 777.2344
66 | 16590 | 777.6562
66 | 16586 | 777.4688
66 | 16585 | 777.4219
66 | 16591 | 777.7031
66 | 16582 | 777.2812
66 | 16594 | 777.8438
66 | 16585 | 777.4219
66 | 16591 | 777.7031
66 | 16591 | 777.7031
66 | 16582 | 777.2812
66 | 16593 | 777.7969
66 | 16583 | 777.3281
66 | 16591 | 777.7031
66 | 16586 | 777.4688
66 | 16584 | 777.375
66 | 16590 | 777.6562
66 | 16581 | 777.2344
66 | 16592 | 777.75
66 | 16589 | 777.6094
66 | 16584 | 777.375
66 | 16593 | 777.7969
66 | 16583 | 777.3281
66 | 16592 | 777.75
66 | 16589 | 777.6094
66 | 16587 | 777.5156
66 | 16591 | 777.7031
66 | 16581 | 777.2344
66 | 16592 | 777.75
66 | 16585 | 777.4219
66 | 16588 | 777.5625
66 | 16589 | 777.6094
66 | 16579 | 777.1406
66 | 16591 | 777.7031
66 | 16583 | 777.3281
66 | 16589 | 777.6094
66 | 16588 | 777.5625
66 | 16586 | 777.4688
66 | 16589 | 777.6094
66 | 16582 | 777.2812
66 | 16593 | 777.7969
66 | 16586 | 777.4688
66 | 16588 | 777.5625
66 | 16590 | 777.6562
66 | 16580 | 777.1875
66 | 16591 | 777.7031
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16587 | 777.5156
66 | 16585 | 777.4219
66 | 16587 | 777.5156
66 | 16580 | 777.1875
66 | 16591 | 777.7031
66 | 16581 | 777.2344
66 | 16593 | 777.7969
66 | 16585 | 777.4219
66 | 16589 | 777.6094
66 | 16591 | 777.7031
66 | 16582 | 777.2812
66 | 16590 | 777.6562
66 | 16579 | 777.1406
66 | 16592 | 777.75
66 | 16585 | 777.4219
66 | 16585 | 777.4219
66 | 16588 | 777.5625
66 | 16579 | 777.1406
66 | 16590 | 777.6562
66 | 16580 | 777.1875
66 | 16592 | 777.75
66 | 16586 | 777.4688
66 | 16588 | 777.5625
66 | 16590 | 777.6562
66 | 16581 | 777.2344
66 | 16592 | 777.75
66 | 16582 | 777.2812
66 | 16590 | 777.6562
66 | 16585 | 777.4219
66 | 16587 | 777.5156
66 | 16587 | 777.5156
66 | 16584 | 777.375
66 | 16587 | 777.5156
66 | 16581 | 777.2344
66 | 16589 | 777.6094
66 | 16578 | 777.0938
66 | 16591 | 777.7031
66 | 16582 | 777.2812
66 | 16591 | 777.7031
66 | 16585 | 777.4219
66 | 16587 | 777.5156
66 | 16590 | 777.6562
66 | 16584 | 777.375
66 | 16590 | 777.6562
66 | 16579 | 777.1406
66 | 16590 | 777.6562
66 | 16580 | 777.1875
66 | 16588 | 777.5625
66 | 16582 | 777.2812
66 | 16587 | 777.5156
66 | 16587 | 777.5156
66 | 16585 | 777.4219
66 | 16588 | 777.5625
66 | 16585 | 777.4219
66 | 16588 | 777.5625
66 | 16584 | 777.375
66 | 16589 | 777.6094
66 | 16583 | 777.3281
66 | 16590 | 777.6562
66 | 16579 | 777.1406
66 | 16590 | 777.6562
66 | 16579 | 777.1406
66 | 16588 | 777.5625
66 | 16581 | 777.2344
66 | 16588 | 777.5625
66 | 16582 | 777.2812
66 | 16587 | 777.5156
66 | 16584 | 777.375
66 | 16587 | 777.5156
66 | 16586 | 777.4688
66 | 16587 | 777.5156
66 | 16588 | 777.5625
66 | 16584 | 777.375
66 | 16588 | 777.5625
66 | 16587 | 777.5156
66 | 16587 | 777.5156
66 | 16586 | 777.4688
66 | 16586 | 777.4688
66 | 16583 | 777.3281
66 | 16586 | 777.4688
66 | 16580 | 777.1875
66 | 16587 | 777.5156
66 | 16584 | 777.375
66 | 16588 | 777.5625
66 | 16583 | 777.3281
66 | 16589 | 777.6094
66 | 16582 | 777.2812
66 | 16588 | 777.5625
66 | 16579 | 777.1406
66 | 16589 | 777.6094
66 | 16579 | 777.1406
66 | 16589 | 777.6094
66 | 16580 | 777.1875
66 | 16586 | 777.4688
66 | 16581 | 777.2344
66 | 16586 | 777.4688
66 | 16581 | 777.2344
66 | 16587 | 777.5156
66 | 16585 | 777.4219
66 | 16586 | 777.4688
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16586 | 777.4688
66 | 16587 | 777.5156
66 | 16585 | 777.4219
66 | 16588 | 777.5625
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16582 | 777.2812
66 | 16585 | 777.4219
66 | 16582 | 777.2812
66 | 16585 | 777.4219
66 | 16582 | 777.2812
66 | 16585 | 777.4219
66 | 16582 | 777.2812
66 | 16584 | 777.375
66 | 16583 | 777.3281
66 | 16586 | 777.4688
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16582 | 777.2812
66 | 16587 | 777.5156
66 | 16580 | 777.1875
66 | 16586 | 777.4688
66 | 16580 | 777.1875
66 | 16586 | 777.4688
66 | 16580 | 777.1875
66 | 16586 | 777.4688
66 | 16580 | 777.1875
66 | 16587 | 777.5156
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16583 | 777.3281
66 | 16588 | 777.5625
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16583 | 777.3281
66 | 16584 | 777.375
66 | 16584 | 777.375
66 | 16583 | 777.3281
66 | 16579 | 777.1406
66 | 16585 | 777.4219
66 | 16579 | 777.1406
66 | 16585 | 777.4219
66 | 16580 | 777.1875
66 | 16587 | 777.5156
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16583 | 777.3281
66 | 16587 | 777.5156
66 | 16582 | 777.2812
66 | 16585 | 777.4219
66 | 16582 | 777.2812
66 | 16584 | 777.375
66 | 16580 | 777.1875
66 | 16585 | 777.4219
66 | 16581 | 777.2344
66 | 16585 | 777.4219
66 | 16578 | 777.0938
66 | 16585 | 777.4219
66 | 16578 | 777.0938
66 | 16586 | 777.4688
66 | 16581 | 777.2344
66 | 16588 | 777.5625
66 | 16579 | 777.1406
66 | 16589 | 777.6094
66 | 16579 | 777.1406
66 | 16587 | 777.5156
66 | 16579 | 777.1406
66 | 16587 | 777.5156
66 | 16579 | 777.1406
66 | 16585 | 777.4219
66 | 16577 | 777.0469
66 | 16585 | 777.4219
66 | 16577 | 777.0469
66 | 16587 | 777.5156
66 | 16577 | 777.0469
66 | 16587 | 777.5156
66 | 16577 | 777.0469
66 | 16585 | 777.4219
66 | 16580 | 777.1875
66 | 16588 | 777.5625
66 | 16579 | 777.1406
66 | 16588 | 777.5625
66 | 16577 | 777.0469
66 | 16587 | 777.5156
66 | 16576 | 777
66 | 16585 | 777.4219
66 | 16575 | 776.9531
66 | 16585 | 777.4219
66 | 16575 | 776.9531
66 | 16584 | 777.375
66 | 16576 | 777
66 | 16587 | 777.5156
66 | 16578 | 777.0938
66 | 16587 | 777.5156
66 | 16577 | 777.0469
66 | 16588 | 777.5625
66 | 16577 | 777.0469
66 | 16587 | 777.5156
66 | 16577 | 777.0469
66 | 16585 | 777.4219
66 | 16576 | 777
66 | 16584 | 777.375
66 | 16574 | 776.9062
66 | 16585 | 777.4219
66 | 16574 | 776.9062
66 | 16585 | 777.4219
66 | 16576 | 777
66 | 16587 | 777.5156
66 | 16577 | 777.0469
66 | 16588 | 777.5625
66 | 16577 | 777.0469
66 | 16588 | 777.5625
66 | 16577 | 777.0469
66 | 16586 | 777.4688
66 | 16575 | 776.9531
66 | 16585 | 777.4219
66 | 16575 | 776.9531
66 | 16585 | 777.4219
66 | 16572 | 776.8125
66 | 16585 | 777.4219
66 | 16574 | 776.9062
66 | 16586 | 777.4688
66 | 16575 | 776.9531
66 | 16586 | 777.4688
66 | 16576 | 777
66 | 16588 | 777.5625
66 | 16576 | 777
66 | 16586 | 777.4688
66 | 16577 | 777.0469
66 | 16583 | 777.3281
66 | 16585 | 777.4219
66 | 16577 | 777.0469
66 | 16585 | 777.4219
66 | 16574 | 776.9062
66 | 16585 | 777.4219
66 | 16573 | 776.8594
66 | 16585 | 777.4219
66 | 16575 | 776.9531
66 | 16585 | 777.4219
66 | 16575 | 776.9531
66 | 16586 | 777.4688
66 | 16576 | 777
66 | 16586 | 777.4688
66 | 16575 | 776.9531
66 | 16586 | 777.4688
66 | 16574 | 776.9062
66 | 16584 | 777.375
66 | 16574 | 776.9062
66 | 16583 | 777.3281
66 | 16573 | 776.8594
66 | 16584 | 777.375
66 | 16573 | 776.8594
66 | 16583 | 777.3281
66 | 16577 | 777.0469
66 | 16580 | 777.1875
66 | 16584 | 777.375
66 | 16577 | 777.0469
66 | 16585 | 777.4219
66 | 16577 | 777.0469
66 | 16585 | 777.4219
66 | 16575 | 776.9531
66 | 16586 | 777.4688
66 | 16573 | 776.8594
66 | 16584 | 777.375
66 | 16573 | 776.8594
66 | 16580 | 777.1875
66 | 16583 | 777.3281
66 | 16574 | 776.9062
66 | 16583 | 777.3281
66 | 16584 | 777.375
66 | 16583 | 777.3281
66 | 16581 | 777.2344
66 | 16574 | 776.9062
66 | 16576 | 777
66 | 16579 | 777.1406
66 | 16575 | 776.9531
66 | 16586 | 777.4688
66 | 16582 | 777.2812
66 | 16586 | 777.4688
66 | 16584 | 777.375
66 | 16582 | 777.2812
66 | 16583 | 777.3281
66 | 16573 | 776.8594
66 | 16582 | 777.2812
66 | 16572 | 776.8125
66 | 16582 | 777.2812
66 | 16578 | 777.0938
66 | 16573 | 776.8594
66 | 16584 | 777.375
66 | 16572 | 776.8125
66 | 16584 | 777.375
66 | 16572 | 776.8125
66 | 16583 | 777.3281
66 | 16578 | 777.0938
66 | 16577 | 777.0469
66 | 16584 | 777.375
66 | 16573 | 776.8594
66 | 16583 | 777.3281
66 | 16572 | 776.8125
66 | 16580 | 777.1875
66 | 16579 | 777.1406
66 | 16574 | 776.9062
66 | 16582 | 777.2812
66 | 16571 | 776.7656
66 | 16582 | 777.2812
66 | 16572 | 776.8125
66 | 16581 | 777.2344
66 | 16580 | 777.1875
66 | 16576 | 777
66 | 16585 | 777.4219
66 | 16573 | 776.8594
66 | 16582 | 777.2812
66 | 16580 | 777.1875
66 | 16574 | 776.9062
66 | 16583 | 777.3281
66 | 16571 | 776.7656
66 | 16580 | 777.1875
66 | 16579 | 777.1406
66 | 16571 | 776.7656
66 | 16582 | 777.2812
66 | 16571 | 776.7656
66 | 16581 | 777.2344
66 | 16576 | 777
66 | 16579 | 777.1406
66 | 16584 | 777.375
66 | 16574 | 776.9062
66 | 16584 | 777.375
66 | 16573 | 776.8594
66 | 16583 | 777.3281
66 | 16584 | 777.375
66 | 16574 | 776.9062
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16581 | 777.2344
66 | 16578 | 777.0938
66 | 16571 | 776.7656
66 | 16582 | 777.2812
66 | 16571 | 776.7656
66 | 16581 | 777.2344
66 | 16581 | 777.2344
66 | 16574 | 776.9062
66 | 16584 | 777.375
66 | 16572 | 776.8125
66 | 16582 | 777.2812
66 | 16580 | 777.1875
66 | 16575 | 776.9531
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16581 | 777.2344
66 | 16573 | 776.8594
66 | 16573 | 776.8594
66 | 16582 | 777.2812
66 | 16569 | 776.6719
66 | 16579 | 777.1406
66 | 16581 | 777.2344
66 | 16573 | 776.8594
66 | 16583 | 777.3281
66 | 16573 | 776.8594
66 | 16581 | 777.2344
66 | 16582 | 777.2812
66 | 16573 | 776.8594
66 | 16583 | 777.3281
66 | 16572 | 776.8125
66 | 16579 | 777.1406
66 | 16580 | 777.1875
66 | 16571 | 776.7656
66 | 16581 | 777.2344
66 | 16571 | 776.7656
66 | 16576 | 777
66 | 16582 | 777.2812
66 | 16572 | 776.8125
66 | 16583 | 777.3281
66 | 16571 | 776.7656
66 | 16581 | 777.2344
66 | 16583 | 777.3281
66 | 16574 | 776.9062
66 | 16583 | 777.3281
66 | 16572 | 776.8125
66 | 16577 | 777.0469
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16581 | 777.2344
66 | 16568 | 776.625
66 | 16575 | 776.9531
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16579 | 777.1406
66 | 16581 | 777.2344
66 | 16571 | 776.7656
66 | 16581 | 777.2344
66 | 16575 | 776.9531
66 | 16576 | 777
66 | 16584 | 777.375
66 | 16573 | 776.8594
66 | 16578 | 777.0938
66 | 16583 | 777.3281
66 | 16570 | 776.7188
66 | 16580 | 777.1875
66 | 16575 | 776.9531
66 | 16572 | 776.8125
66 | 16580 | 777.1875
66 | 16572 | 776.8125
66 | 16577 | 777.0469
66 | 16582 | 777.2812
66 | 16572 | 776.8125
66 | 16580 | 777.1875
66 | 16584 | 777.375
66 | 16573 | 776.8594
66 | 16581 | 777.2344
66 | 16583 | 777.3281
66 | 16571 | 776.7656
66 | 16581 | 777.2344
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16578 | 777.0938
66 | 16580 | 777.1875
66 | 16570 | 776.7188
66 | 16578 | 777.0938
66 | 16582 | 777.2812
66 | 16570 | 776.7188
66 | 16579 | 777.1406
66 | 16582 | 777.2812
66 | 16571 | 776.7656
66 | 16580 | 777.1875
66 | 16583 | 777.3281
66 | 16570 | 776.7188
66 | 16579 | 777.1406
66 | 16581 | 777.2344
66 | 16570 | 776.7188
66 | 16577 | 777.0469
66 | 16580 | 777.1875
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16580 | 777.1875
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16582 | 777.2812
66 | 16571 | 776.7656
66 | 16579 | 777.1406
66 | 16582 | 777.2812
66 | 16572 | 776.8125
66 | 16574 | 776.9062
66 | 16581 | 777.2344
66 | 16576 | 777
66 | 16569 | 776.6719
66 | 16578 | 777.0938
66 | 16575 | 776.9531
66 | 16568 | 776.625
66 | 16578 | 777.0938
66 | 16578 | 777.0938
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16581 | 777.2344
66 | 16571 | 776.7656
66 | 16579 | 777.1406
66 | 16580 | 777.1875
66 | 16570 | 776.7188
66 | 16578 | 777.0938
66 | 16580 | 777.1875
66 | 16569 | 776.6719
66 | 16575 | 776.9531
66 | 16579 | 777.1406
66 | 16568 | 776.625
66 | 16574 | 776.9062
66 | 16579 | 777.1406
66 | 16568 | 776.625
66 | 16574 | 776.9062
66 | 16580 | 777.1875
66 | 16573 | 776.8594
66 | 16571 | 776.7656
66 | 16580 | 777.1875
66 | 16578 | 777.0938
66 | 16570 | 776.7188
66 | 16578 | 777.0938
66 | 16580 | 777.1875
66 | 16568 | 776.625
66 | 16576 | 777
66 | 16579 | 777.1406
66 | 16568 | 776.625
66 | 16573 | 776.8594
66 | 16578 | 777.0938
66 | 16574 | 776.9062
66 | 16570 | 776.7188
66 | 16579 | 777.1406
66 | 16580 | 777.1875
66 | 16570 | 776.7188
66 | 16578 | 777.0938
66 | 16581 | 777.2344
66 | 16570 | 776.7188
66 | 16576 | 777
66 | 16579 | 777.1406
66 | 16568 | 776.625
66 | 16572 | 776.8125
66 | 16578 | 777.0938
66 | 16571 | 776.7656
66 | 16568 | 776.625
66 | 16578 | 777.0938
66 | 16579 | 777.1406
66 | 16568 | 776.625
66 | 16575 | 776.9531
66 | 16580 | 777.1875
66 | 16571 | 776.7656
66 | 16573 | 776.8594
66 | 16580 | 777.1875
66 | 16576 | 777
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16578 | 777.0938
66 | 16568 | 776.625
66 | 16573 | 776.8594
66 | 16577 | 777.0469
66 | 16568 | 776.625
66 | 16570 | 776.7188
66 | 16579 | 777.1406
66 | 16573 | 776.8594
66 | 16570 | 776.7188
66 | 16580 | 777.1875
66 | 16580 | 777.1875
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16581 | 777.2344
66 | 16569 | 776.6719
66 | 16573 | 776.8594
66 | 16579 | 777.1406
66 | 16571 | 776.7656
66 | 16567 | 776.5781
66 | 16578 | 777.0938
66 | 16576 | 777
66 | 16566 | 776.5312
66 | 16575 | 776.9531
66 | 16580 | 777.1875
66 | 16568 | 776.625
66 | 16573 | 776.8594
66 | 16578 | 777.0938
66 | 16578 | 777.0938
66 | 16569 | 776.6719
66 | 16576 | 777
66 | 16579 | 777.1406
66 | 16567 | 776.5781
66 | 16570 | 776.7188
66 | 16577 | 777.0469
66 | 16574 | 776.9062
66 | 16566 | 776.5312
66 | 16576 | 777
66 | 16578 | 777.0938
66 | 16567 | 776.5781
66 | 16574 | 776.9062
66 | 16578 | 777.0938
66 | 16571 | 776.7656
66 | 16570 | 776.7188
66 | 16578 | 777.0938
66 | 16579 | 777.1406
66 | 16567 | 776.5781
66 | 16571 | 776.7656
66 | 16576 | 777
66 | 16576 | 777
66 | 16565 | 776.4844
66 | 16570 | 776.7188
66 | 16577 | 777.0469
66 | 16572 | 776.8125
66 | 16566 | 776.5312
66 | 16574 | 776.9062
66 | 16578 | 777.0938
66 | 16572 | 776.8125
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16578 | 777.0938
66 | 16569 | 776.6719
66 | 16570 | 776.7188
66 | 16577 | 777.0469
66 | 16577 | 777.0469
66 | 16565 | 776.4844
66 | 16571 | 776.7656
66 | 16575 | 776.9531
66 | 16570 | 776.7188
66 | 16565 | 776.4844
66 | 16572 | 776.8125
66 | 16577 | 777.0469
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16576 | 777
66 | 16577 | 777.0469
66 | 16569 | 776.6719
66 | 16569 | 776.6719
66 | 16576 | 777
66 | 16576 | 777
66 | 16567 | 776.5781
66 | 16567 | 776.5781
66 | 16574 | 776.9062
66 | 16575 | 776.9531
66 | 16564 | 776.4375
66 | 16568 | 776.625
66 | 16576 | 777
66 | 16577 | 777.0469
66 | 16566 | 776.5312
66 | 16569 | 776.6719
66 | 16577 | 777.0469
66 | 16577 | 777.0469
66 | 16567 | 776.5781
66 | 16570 | 776.7188
66 | 16576 | 777
66 | 16576 | 777
66 | 16566 | 776.5312
66 | 16568 | 776.625
66 | 16573 | 776.8594
66 | 16576 | 777
66 | 16565 | 776.4844
66 | 16565 | 776.4844
66 | 16573 | 776.8594
66 | 16576 | 777
66 | 16568 | 776.625
66 | 16568 | 776.625
66 | 16575 | 776.9531
66 | 16578 | 777.0938
66 | 16568 | 776.625
66 | 16566 | 776.5312
66 | 16576 | 777
66 | 16576 | 777
66 | 16568 | 776.625
66 | 16565 | 776.4844
66 | 16574 | 776.9062
66 | 16575 | 776.9531
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16575 | 776.9531
66 | 16577 | 777.0469
66 | 16566 | 776.5312
66 | 16570 | 776.7188
66 | 16576 | 777
66 | 16577 | 777.0469
66 | 16567 | 776.5781
66 | 16568 | 776.625
66 | 16576 | 777
66 | 16576 | 777
66 | 16566 | 776.5312
66 | 16563 | 776.3906
66 | 16571 | 776.7656
66 | 16575 | 776.9531
66 | 16569 | 776.6719
66 | 16565 | 776.4844
66 | 16569 | 776.6719
66 | 16575 | 776.9531
66 | 16577 | 777.0469
66 | 16569 | 776.6719
66 | 16568 | 776.625
66 | 16576 | 777
66 | 16578 | 777.0938
66 | 16570 | 776.7188
66 | 16565 | 776.4844
66 | 16572 | 776.8125
66 | 16574 | 776.9062
66 | 16572 | 776.8125
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16574 | 776.9062
66 | 16576 | 777
66 | 16568 | 776.625
66 | 16565 | 776.4844
66 | 16569 | 776.6719
66 | 16576 | 777
66 | 16577 | 777.0469
66 | 16567 | 776.5781
66 | 16568 | 776.625
66 | 16575 | 776.9531
66 | 16575 | 776.9531
66 | 16565 | 776.4844
66 | 16566 | 776.5312
66 | 16573 | 776.8594
66 | 16574 | 776.9062
66 | 16562 | 776.3438
66 | 16566 | 776.5312
66 | 16573 | 776.8594
66 | 16576 | 777
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16574 | 776.9062
66 | 16577 | 777.0469
66 | 16569 | 776.6719
66 | 16564 | 776.4375
66 | 16570 | 776.7188
66 | 16576 | 777
66 | 16574 | 776.9062
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16572 | 776.8125
66 | 16574 | 776.9062
66 | 16565 | 776.4844
66 | 16565 | 776.4844
66 | 16571 | 776.7656
66 | 16575 | 776.9531
66 | 16576 | 777
66 | 16564 | 776.4375
66 | 16567 | 776.5781
66 | 16574 | 776.9062
66 | 16576 | 777
66 | 16569 | 776.6719
66 | 16564 | 776.4375
66 | 16566 | 776.5312
66 | 16573 | 776.8594
66 | 16573 | 776.8594
66 | 16565 | 776.4844
66 | 16563 | 776.3906
66 | 16571 | 776.7656
66 | 16575 | 776.9531
66 | 16573 | 776.8594
66 | 16564 | 776.4375
66 | 16568 | 776.625
66 | 16574 | 776.9062
66 | 16576 | 777
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16571 | 776.7656
66 | 16575 | 776.9531
66 | 16568 | 776.625
66 | 16562 | 776.3438
66 | 16567 | 776.5781
66 | 16572 | 776.8125
66 | 16573 | 776.8594
66 | 16565 | 776.4844
66 | 16564 | 776.4375
66 | 16573 | 776.8594
66 | 16576 | 777
66 | 16569 | 776.6719
66 | 16565 | 776.4844
66 | 16566 | 776.5312
66 | 16573 | 776.8594
66 | 16574 | 776.9062
66 | 16566 | 776.5312
66 | 16562 | 776.3438
66 | 16567 | 776.5781
66 | 16572 | 776.8125
66 | 16573 | 776.8594
66 | 16565 | 776.4844
66 | 16563 | 776.3906
66 | 16569 | 776.6719
66 | 16573 | 776.8594
66 | 16575 | 776.9531
66 | 16565 | 776.4844
66 | 16565 | 776.4844
66 | 16573 | 776.8594
66 | 16576 | 777
66 | 16573 | 776.8594
66 | 16565 | 776.4844
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16572 | 776.8125
66 | 16573 | 776.8594
66 | 16568 | 776.625
66 | 16562 | 776.3438
66 | 16565 | 776.4844
66 | 16571 | 776.7656
66 | 16574 | 776.9062
66 | 16573 | 776.8594
66 | 16564 | 776.4375
66 | 16563 | 776.3906
66 | 16571 | 776.7656
66 | 16573 | 776.8594
66 | 16571 | 776.7656
66 | 16563 | 776.3906
66 | 16562 | 776.3438
66 | 16566 | 776.5312
66 | 16572 | 776.8125
66 | 16572 | 776.8125
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16570 | 776.7188
66 | 16574 | 776.9062
66 | 16573 | 776.8594
66 | 16565 | 776.4844
66 | 16563 | 776.3906
66 | 16571 | 776.7656
66 | 16573 | 776.8594
66 | 16572 | 776.8125
66 | 16565 | 776.4844
66 | 16561 | 776.2969
66 | 16564 | 776.4375
66 | 16570 | 776.7188
66 | 16572 | 776.8125
66 | 16565 | 776.4844
66 | 16560 | 776.25
66 | 16563 | 776.3906
66 | 16570 | 776.7188
66 | 16573 | 776.8594
66 | 16574 | 776.9062
66 | 16564 | 776.4375
66 | 16563 | 776.3906
66 | 16568 | 776.625
66 | 16573 | 776.8594
66 | 16572 | 776.8125
66 | 16564 | 776.4375
66 | 16561 | 776.2969
66 | 16564 | 776.4375
66 | 16570 | 776.7188
66 | 16572 | 776.8125
66 | 16572 | 776.8125
66 | 16565 | 776.4844
66 | 16562 | 776.3438
66 | 16564 | 776.4375
66 | 16572 | 776.8125
66 | 16574 | 776.9062
66 | 16576 | 777
66 | 16567 | 776.5781
66 | 16562 | 776.3438
66 | 16564 | 776.4375
66 | 16571 | 776.7656
66 | 16572 | 776.8125
66 | 16570 | 776.7188
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16564 | 776.4375
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16569 | 776.6719
66 | 16562 | 776.3438
66 | 16564 | 776.4375
66 | 16570 | 776.7188
66 | 16574 | 776.9062
66 | 16570 | 776.7188
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16569 | 776.6719
66 | 16571 | 776.7656
66 | 16572 | 776.8125
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16565 | 776.4844
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16571 | 776.7656
66 | 16561 | 776.2969
66 | 16563 | 776.3906
66 | 16569 | 776.6719
66 | 16572 | 776.8125
66 | 16572 | 776.8125
66 | 16566 | 776.5312
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16566 | 776.5312
66 | 16570 | 776.7188
66 | 16572 | 776.8125
66 | 16569 | 776.6719
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16566 | 776.5312
66 | 16572 | 776.8125
66 | 16573 | 776.8594
66 | 16572 | 776.8125
66 | 16565 | 776.4844
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16565 | 776.4844
66 | 16567 | 776.5781
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16562 | 776.3438
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16564 | 776.4375
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16572 | 776.8125
66 | 16564 | 776.4375
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16566 | 776.5312
66 | 16569 | 776.6719
66 | 16569 | 776.6719
66 | 16569 | 776.6719
66 | 16564 | 776.4375
66 | 16558 | 776.1562
66 | 16561 | 776.2969
66 | 16566 | 776.5312
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16573 | 776.8594
66 | 16564 | 776.4375
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16566 | 776.5312
66 | 16569 | 776.6719
66 | 16570 | 776.7188
66 | 16570 | 776.7188
66 | 16569 | 776.6719
66 | 16564 | 776.4375
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16562 | 776.3438
66 | 16567 | 776.5781
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16571 | 776.7656
66 | 16568 | 776.625
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16567 | 776.5781
66 | 16569 | 776.6719
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16569 | 776.6719
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16557 | 776.1094
66 | 16560 | 776.25
66 | 16562 | 776.3438
66 | 16567 | 776.5781
66 | 16570 | 776.7188
66 | 16570 | 776.7188
66 | 16570 | 776.7188
66 | 16571 | 776.7656
66 | 16562 | 776.3438
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16559 | 776.2031
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16569 | 776.6719
66 | 16569 | 776.6719
66 | 16568 | 776.625
66 | 16562 | 776.3438
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16562 | 776.3438
66 | 16569 | 776.6719
66 | 16567 | 776.5781
66 | 16569 | 776.6719
66 | 16569 | 776.6719
66 | 16567 | 776.5781
66 | 16560 | 776.25
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16569 | 776.6719
66 | 16570 | 776.7188
66 | 16568 | 776.625
66 | 16563 | 776.3906
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16564 | 776.4375
66 | 16567 | 776.5781
66 | 16567 | 776.5781
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16563 | 776.3906
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16561 | 776.2969
66 | 16568 | 776.625
66 | 16568 | 776.625
66 | 16569 | 776.6719
66 | 16570 | 776.7188
66 | 16569 | 776.6719
66 | 16568 | 776.625
66 | 16565 | 776.4844
66 | 16561 | 776.2969
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16560 | 776.25
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16570 | 776.7188
66 | 16568 | 776.625
66 | 16570 | 776.7188
66 | 16569 | 776.6719
66 | 16567 | 776.5781
66 | 16563 | 776.3906
66 | 16557 | 776.1094
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16563 | 776.3906
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16568 | 776.625
66 | 16568 | 776.625
66 | 16569 | 776.6719
66 | 16563 | 776.3906
66 | 16561 | 776.2969
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16562 | 776.3438
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16568 | 776.625
66 | 16566 | 776.5312
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16556 | 776.0625
66 | 16559 | 776.2031
66 | 16564 | 776.4375
66 | 16564 | 776.4375
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16567 | 776.5781
66 | 16565 | 776.4844
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16557 | 776.1094
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16564 | 776.4375
66 | 16565 | 776.4844
66 | 16565 | 776.4844
66 | 16566 | 776.5312
66 | 16567 | 776.5781
66 | 16568 | 776.625
66 | 16567 | 776.5781
66 | 16563 | 776.3906
66 | 16562 | 776.3438
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16555 | 776.0156
66 | 16554 | 775.9688
66 | 16554 | 775.9688
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16559 | 776.2031
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16565 | 776.4844
66 | 16565 | 776.4844
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16565 | 776.4844
66 | 16565 | 776.4844
66 | 16566 | 776.5312
66 | 16564 | 776.4375
66 | 16563 | 776.3906
66 | 16564 | 776.4375
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16566 | 776.5312
66 | 16564 | 776.4375
66 | 16559 | 776.2031
66 | 16562 | 776.3438
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16564 | 776.4375
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16554 | 775.9688
66 | 16555 | 776.0156
66 | 16555 | 776.0156
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16563 | 776.3906
66 | 16563 | 776.3906
66 | 16566 | 776.5312
66 | 16565 | 776.4844
66 | 16567 | 776.5781
66 | 16566 | 776.5312
66 | 16565 | 776.4844
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16562 | 776.3438
66 | 16563 | 776.3906
66 | 16562 | 776.3438
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16556 | 776.0625
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16562 | 776.3438
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16562 | 776.3438
66 | 16563 | 776.3906
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16562 | 776.3438
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16563 | 776.3906
66 | 16564 | 776.4375
66 | 16563 | 776.3906
66 | 16563 | 776.3906
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16562 | 776.3438
66 | 16562 | 776.3438
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16562 | 776.3438
66 | 16560 | 776.25
66 | 16563 | 776.3906
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16561 | 776.2969
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16561 | 776.2969
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16560 | 776.25
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16556 | 776.0625
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16552 | 775.875
66 | 16549 | 775.7344
66 | 16548 | 775.6875
66 | 16549 | 775.7344
66 | 16555 | 776.0156
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16555 | 776.0156
66 | 16557 | 776.1094
66 | 16551 | 775.8281
66 | 16550 | 775.7812
66 | 16548 | 775.6875
66 | 16547 | 775.6406
66 | 16552 | 775.875
66 | 16557 | 776.1094
66 | 16559 | 776.2031
66 | 16560 | 776.25
66 | 16559 | 776.2031
66 | 16559 | 776.2031
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16549 | 775.7344
66 | 16545 | 775.5469
66 | 16544 | 775.5
66 | 16545 | 775.5469
66 | 16553 | 775.9219
66 | 16555 | 776.0156
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16559 | 776.2031
66 | 16557 | 776.1094
66 | 16554 | 775.9688
66 | 16550 | 775.7812
66 | 16547 | 775.6406
66 | 16547 | 775.6406
66 | 16549 | 775.7344
66 | 16554 | 775.9688
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16556 | 776.0625
66 | 16552 | 775.875
66 | 16549 | 775.7344
66 | 16548 | 775.6875
66 | 16549 | 775.7344
66 | 16554 | 775.9688
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16558 | 776.1562
66 | 16557 | 776.1094
66 | 16554 | 775.9688
66 | 16549 | 775.7344
66 | 16547 | 775.6406
66 | 16545 | 775.5469
66 | 16546 | 775.5938
66 | 16553 | 775.9219
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16551 | 775.8281
66 | 16548 | 775.6875
66 | 16545 | 775.5469
66 | 16548 | 775.6875
66 | 16554 | 775.9688
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16556 | 776.0625
66 | 16553 | 775.9219
66 | 16550 | 775.7812
66 | 16545 | 775.5469
66 | 16546 | 775.5938
66 | 16550 | 775.7812
66 | 16555 | 776.0156
66 | 16557 | 776.1094
66 | 16558 | 776.1562
66 | 16556 | 776.0625
66 | 16553 | 775.9219
66 | 16546 | 775.5938
66 | 16546 | 775.5938
66 | 16551 | 775.8281
66 | 16554 | 775.9688
66 | 16554 | 775.9688
66 | 16554 | 775.9688
66 | 16554 | 775.9688
66 | 16554 | 775.9688
66 | 16551 | 775.8281
66 | 16548 | 775.6875
66 | 16546 | 775.5938
66 | 16546 | 775.5938
66 | 16555 | 776.0156
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16557 | 776.1094
66 | 16553 | 775.9219
66 | 16548 | 775.6875
66 | 16544 | 775.5
66 | 16550 | 775.7812
66 | 16555 | 776.0156
66 | 16555 | 776.0156
66 | 16554 | 775.9688
66 | 16551 | 775.8281
66 | 16546 | 775.5938
66 | 16545 | 775.5469
66 | 16546 | 775.5938
66 | 16554 | 775.9688
66 | 16556 | 776.0625
66 | 16557 | 776.1094
66 | 16555 | 776.0156
66 | 16549 | 775.7344
66 | 16546 | 775.5938
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16554 | 775.9688
66 | 16547 | 775.6406
66 | 16543 | 775.4531
66 | 16550 | 775.7812
66 | 16553 | 775.9219
66 | 16555 | 776.0156
66 | 16553 | 775.9219
66 | 16548 | 775.6875
66 | 16546 | 775.5938
66 | 16553 | 775.9219
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16545 | 775.5469
66 | 16554 | 775.9688
66 | 16555 | 776.0156
66 | 16553 | 775.9219
66 | 16549 | 775.7344
66 | 16544 | 775.5
66 | 16543 | 775.4531
66 | 16552 | 775.875
66 | 16554 | 775.9688
66 | 16556 | 776.0625
66 | 16553 | 775.9219
66 | 16548 | 775.6875
66 | 16545 | 775.5469
66 | 16553 | 775.9219
66 | 16557 | 776.1094
66 | 16555 | 776.0156
66 | 16551 | 775.8281
66 | 16546 | 775.5938
66 | 16542 | 775.4062
66 | 16550 | 775.7812
66 | 16553 | 775.9219
66 | 16552 | 775.875
66 | 16553 | 775.9219
66 | 16547 | 775.6406
66 | 16544 | 775.5
66 | 16544 | 775.5
66 | 16553 | 775.9219
66 | 16555 | 776.0156
66 | 16557 | 776.1094
66 | 16550 | 775.7812
66 | 16545 | 775.5469
66 | 16546 | 775.5938
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16551 | 775.8281
66 | 16545 | 775.5469
66 | 16543 | 775.4531
66 | 16554 | 775.9688
66 | 16552 | 775.875
66 | 16553 | 775.9219
66 | 16546 | 775.5938
66 | 16543 | 775.4531
66 | 16555 | 776.0156
66 | 16556 | 776.0625
66 | 16555 | 776.0156
66 | 16547 | 775.6406
66 | 16545 | 775.5469
66 | 16554 | 775.9688
66 | 16554 | 775.9688
66 | 16550 | 775.7812
66 | 16544 | 775.5
66 | 16545 | 775.5469
66 | 16554 | 775.9688
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16542 | 775.4062
66 | 16551 | 775.8281
66 | 16554 | 775.9688
66 | 16552 | 775.875
66 | 16544 | 775.5
66 | 16544 | 775.5
66 | 16556 | 776.0625
66 | 16554 | 775.9688
66 | 16549 | 775.7344
66 | 16543 | 775.4531
66 | 16549 | 775.7344
66 | 16553 | 775.9219
66 | 16550 | 775.7812
66 | 16544 | 775.5
66 | 16543 | 775.4531
66 | 16552 | 775.875
66 | 16552 | 775.875
66 | 16547 | 775.6406
66 | 16544 | 775.5
66 | 16553 | 775.9219
66 | 16554 | 775.9688
66 | 16548 | 775.6875
66 | 16544 | 775.5
66 | 16552 | 775.875
66 | 16554 | 775.9688
66 | 16550 | 775.7812
66 | 16542 | 775.4062
66 | 16542 | 775.4062
66 | 16552 | 775.875
66 | 16550 | 775.7812
66 | 16541 | 775.3594
66 | 16542 | 775.4062
66 | 16551 | 775.8281
66 | 16553 | 775.9219
66 | 16546 | 775.5938
66 | 16542 | 775.4062
66 | 16554 | 775.9688
66 | 16553 | 775.9219
66 | 16546 | 775.5938
66 | 16544 | 775.5
66 | 16553 | 775.9219
66 | 16553 | 775.9219
66 | 16546 | 775.5938
66 | 16541 | 775.3594
66 | 16550 | 775.7812
66 | 16551 | 775.8281
66 | 16545 | 775.5469
66 | 16540 | 775.3125
66 | 16551 | 775.8281
66 | 16552 | 775.875
66 | 16550 | 775.7812
66 | 16542 | 775.4062
66 | 16543 | 775.4531
66 | 16553 | 775.9219
66 | 16549 | 775.7344
66 | 16542 | 775.4062
66 | 16548 | 775.6875
66 | 16554 | 775.9688
66 | 16550 | 775.7812
66 | 16543 | 775.4531
66 | 16545 | 775.5469
66 | 16551 | 775.8281
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16544 | 775.5
66 | 16552 | 775.875
66 | 16551 | 775.8281
66 | 16543 | 775.4531
66 | 16545 | 775.5469
66 | 16554 | 775.9688
66 | 16552 | 775.875
66 | 16544 | 775.5
66 | 16544 | 775.5
66 | 16553 | 775.9219
66 | 16549 | 775.7344
66 | 16542 | 775.4062
66 | 16541 | 775.3594
66 | 16551 | 775.8281
66 | 16550 | 775.7812
66 | 16543 | 775.4531
66 | 16540 | 775.3125
66 | 16548 | 775.6875
66 | 16552 | 775.875
66 | 16549 | 775.7344
66 | 16542 | 775.4062
66 | 16547 | 775.6406
66 | 16553 | 775.9219
66 | 16550 | 775.7812
66 | 16540 | 775.3125
66 | 16547 | 775.6406
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16541 | 775.3594
66 | 16547 | 775.6406
66 | 16552 | 775.875
66 | 16547 | 775.6406
66 | 16541 | 775.3594
66 | 16544 | 775.5
66 | 16552 | 775.875
66 | 16549 | 775.7344
66 | 16543 | 775.4531
66 | 16542 | 775.4062
66 | 16554 | 775.9688
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16542 | 775.4062
66 | 16550 | 775.7812
66 | 16551 | 775.8281
66 | 16546 | 775.5938
66 | 16540 | 775.3125
66 | 16548 | 775.6875
66 | 16551 | 775.8281
66 | 16546 | 775.5938
66 | 16540 | 775.3125
66 | 16546 | 775.5938
66 | 16552 | 775.875
66 | 16549 | 775.7344
66 | 16542 | 775.4062
66 | 16546 | 775.5938
66 | 16553 | 775.9219
66 | 16551 | 775.8281
66 | 16544 | 775.5
66 | 16544 | 775.5
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16540 | 775.3125
66 | 16547 | 775.6406
66 | 16549 | 775.7344
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16544 | 775.5
66 | 16550 | 775.7812
66 | 16544 | 775.5
66 | 16541 | 775.3594
66 | 16551 | 775.8281
66 | 16552 | 775.875
66 | 16545 | 775.5469
66 | 16542 | 775.4062
66 | 16552 | 775.875
66 | 16547 | 775.6406
66 | 16541 | 775.3594
66 | 16552 | 775.875
66 | 16550 | 775.7812
66 | 16543 | 775.4531
66 | 16542 | 775.4062
66 | 16549 | 775.7344
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16543 | 775.4531
66 | 16552 | 775.875
66 | 16549 | 775.7344
66 | 16542 | 775.4062
66 | 16547 | 775.6406
66 | 16553 | 775.9219
66 | 16548 | 775.6875
66 | 16541 | 775.3594
66 | 16552 | 775.875
66 | 16552 | 775.875
66 | 16543 | 775.4531
66 | 16541 | 775.3594
66 | 16550 | 775.7812
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16550 | 775.7812
66 | 16547 | 775.6406
66 | 16540 | 775.3125
66 | 16550 | 775.7812
66 | 16546 | 775.5938
66 | 16540 | 775.3125
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16540 | 775.3125
66 | 16550 | 775.7812
66 | 16551 | 775.8281
66 | 16543 | 775.4531
66 | 16543 | 775.4531
66 | 16550 | 775.7812
66 | 16544 | 775.5
66 | 16539 | 775.2656
66 | 16548 | 775.6875
66 | 16544 | 775.5
66 | 16540 | 775.3125
66 | 16550 | 775.7812
66 | 16547 | 775.6406
66 | 16542 | 775.4062
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16541 | 775.3594
66 | 16551 | 775.8281
66 | 16551 | 775.8281
66 | 16545 | 775.5469
66 | 16541 | 775.3594
66 | 16549 | 775.7344
66 | 16546 | 775.5938
66 | 16540 | 775.3125
66 | 16549 | 775.7344
66 | 16543 | 775.4531
66 | 16539 | 775.2656
66 | 16550 | 775.7812
66 | 16544 | 775.5
66 | 16541 | 775.3594
66 | 16550 | 775.7812
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16551 | 775.8281
66 | 16546 | 775.5938
66 | 16540 | 775.3125
66 | 16550 | 775.7812
66 | 16545 | 775.5469
66 | 16537 | 775.1719
66 | 16548 | 775.6875
66 | 16549 | 775.7344
66 | 16539 | 775.2656
66 | 16542 | 775.4062
66 | 16550 | 775.7812
66 | 16545 | 775.5469
66 | 16540 | 775.3125
66 | 16550 | 775.7812
66 | 16547 | 775.6406
66 | 16542 | 775.4062
66 | 16552 | 775.875
66 | 16548 | 775.6875
66 | 16541 | 775.3594
66 | 16550 | 775.7812
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16548 | 775.6875
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16540 | 775.3125
66 | 16548 | 775.6875
66 | 16543 | 775.4531
66 | 16541 | 775.3594
66 | 16550 | 775.7812
66 | 16542 | 775.4062
66 | 16541 | 775.3594
66 | 16552 | 775.875
66 | 16543 | 775.4531
66 | 16543 | 775.4531
66 | 16552 | 775.875
66 | 16543 | 775.4531
66 | 16542 | 775.4062
66 | 16549 | 775.7344
66 | 16542 | 775.4062
66 | 16540 | 775.3125
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16542 | 775.4062
66 | 16549 | 775.7344
66 | 16543 | 775.4531
66 | 16544 | 775.5
66 | 16551 | 775.8281
66 | 16544 | 775.5
66 | 16543 | 775.4531
66 | 16551 | 775.8281
66 | 16544 | 775.5
66 | 16543 | 775.4531
66 | 16550 | 775.7812
66 | 16542 | 775.4062
66 | 16541 | 775.3594
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16541 | 775.3594
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16541 | 775.3594
66 | 16550 | 775.7812
66 | 16544 | 775.5
66 | 16543 | 775.4531
66 | 16552 | 775.875
66 | 16544 | 775.5
66 | 16543 | 775.4531
66 | 16551 | 775.8281
66 | 16543 | 775.4531
66 | 16544 | 775.5
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16541 | 775.3594
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16541 | 775.3594
66 | 16542 | 775.4062
66 | 16550 | 775.7812
66 | 16542 | 775.4062
66 | 16544 | 775.5
66 | 16550 | 775.7812
66 | 16541 | 775.3594
66 | 16545 | 775.5469
66 | 16550 | 775.7812
66 | 16540 | 775.3125
66 | 16545 | 775.5469
66 | 16549 | 775.7344
66 | 16538 | 775.2188
66 | 16542 | 775.4062
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16545 | 775.5469
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16546 | 775.5938
66 | 16550 | 775.7812
66 | 16541 | 775.3594
66 | 16545 | 775.5469
66 | 16550 | 775.7812
66 | 16541 | 775.3594
66 | 16544 | 775.5
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16546 | 775.5938
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16542 | 775.4062
66 | 16547 | 775.6406
66 | 16539 | 775.2656
66 | 16546 | 775.5938
66 | 16547 | 775.6406
66 | 16540 | 775.3125
66 | 16544 | 775.5
66 | 16550 | 775.7812
66 | 16541 | 775.3594
66 | 16545 | 775.5469
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16543 | 775.4531
66 | 16550 | 775.7812
66 | 16537 | 775.1719
66 | 16542 | 775.4062
66 | 16548 | 775.6875
66 | 16540 | 775.3125
66 | 16542 | 775.4062
66 | 16548 | 775.6875
66 | 16540 | 775.3125
66 | 16545 | 775.5469
66 | 16550 | 775.7812
66 | 16542 | 775.4062
66 | 16544 | 775.5
66 | 16551 | 775.8281
66 | 16542 | 775.4062
66 | 16543 | 775.4531
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16541 | 775.3594
66 | 16547 | 775.6406
66 | 16541 | 775.3594
66 | 16541 | 775.3594
66 | 16548 | 775.6875
66 | 16539 | 775.2656
66 | 16541 | 775.3594
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16546 | 775.5938
66 | 16549 | 775.7344
66 | 16541 | 775.3594
66 | 16545 | 775.5469
66 | 16549 | 775.7344
66 | 16539 | 775.2656
66 | 16544 | 775.5
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16542 | 775.4062
66 | 16547 | 775.6406
66 | 16539 | 775.2656
66 | 16543 | 775.4531
66 | 16547 | 775.6406
66 | 16539 | 775.2656
66 | 16542 | 775.4062
66 | 16549 | 775.7344
66 | 16540 | 775.3125
66 | 16545 | 775.5469
66 | 16548 | 775.6875
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16549 | 775.7344
66 | 16538 | 775.2188
66 | 16546 | 775.5938
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16544 | 775.5
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16544 | 775.5
66 | 16549 | 775.7344
66 | 16539 | 775.2656
66 | 16546 | 775.5938
66 | 16548 | 775.6875
66 | 16540 | 775.3125
66 | 16549 | 775.7344
66 | 16549 | 775.7344
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16544 | 775.5
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16544 | 775.5
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16542 | 775.4062
66 | 16548 | 775.6875
66 | 16538 | 775.2188
66 | 16545 | 775.5469
66 | 16549 | 775.7344
66 | 16539 | 775.2656
66 | 16546 | 775.5938
66 | 16549 | 775.7344
66 | 16538 | 775.2188
66 | 16543 | 775.4531
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16543 | 775.4531
66 | 16547 | 775.6406
66 | 16536 | 775.125
66 | 16543 | 775.4531
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16546 | 775.5938
66 | 16548 | 775.6875
66 | 16538 | 775.2188
66 | 16547 | 775.6406
66 | 16548 | 775.6875
66 | 16539 | 775.2656
66 | 16544 | 775.5
66 | 16548 | 775.6875
66 | 16538 | 775.2188
66 | 16544 | 775.5
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16544 | 775.5
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16544 | 775.5
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16549 | 775.7344
66 | 16538 | 775.2188
66 | 16546 | 775.5938
66 | 16548 | 775.6875
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16545 | 775.5469
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16548 | 775.6875
66 | 16538 | 775.2188
66 | 16547 | 775.6406
66 | 16547 | 775.6406
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16548 | 775.6875
66 | 16538 | 775.2188
66 | 16547 | 775.6406
66 | 16543 | 775.4531
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16544 | 775.5
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16546 | 775.5938
66 | 16538 | 775.2188
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16541 | 775.3594
66 | 16543 | 775.4531
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16540 | 775.3125
66 | 16544 | 775.5
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16540 | 775.3125
66 | 16543 | 775.4531
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16544 | 775.5
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16549 | 775.7344
66 | 16543 | 775.4531
66 | 16540 | 775.3125
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16544 | 775.5
66 | 16537 | 775.1719
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16548 | 775.6875
66 | 16539 | 775.2656
66 | 16548 | 775.6875
66 | 16545 | 775.5469
66 | 16544 | 775.5
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16542 | 775.4062
66 | 16540 | 775.3125
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16547 | 775.6406
66 | 16539 | 775.2656
66 | 16544 | 775.5
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16540 | 775.3125
66 | 16545 | 775.5469
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16544 | 775.5
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16542 | 775.4062
66 | 16545 | 775.5469
66 | 16546 | 775.5938
66 | 16538 | 775.2188
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16548 | 775.6875
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16540 | 775.3125
66 | 16542 | 775.4062
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16545 | 775.5469
66 | 16541 | 775.3594
66 | 16544 | 775.5
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16545 | 775.5469
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16533 | 774.9844
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16547 | 775.6406
66 | 16538 | 775.2188
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16545 | 775.5469
66 | 16540 | 775.3125
66 | 16542 | 775.4062
66 | 16542 | 775.4062
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16533 | 774.9844
66 | 16545 | 775.5469
66 | 16533 | 774.9844
66 | 16545 | 775.5469
66 | 16537 | 775.1719
66 | 16544 | 775.5
66 | 16542 | 775.4062
66 | 16543 | 775.4531
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16548 | 775.6875
66 | 16536 | 775.125
66 | 16547 | 775.6406
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16533 | 774.9844
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16548 | 775.6875
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16539 | 775.2656
66 | 16547 | 775.6406
66 | 16540 | 775.3125
66 | 16544 | 775.5
66 | 16544 | 775.5
66 | 16538 | 775.2188
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16549 | 775.7344
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16544 | 775.5
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16539 | 775.2656
66 | 16543 | 775.4531
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16544 | 775.5
66 | 16538 | 775.2188
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16533 | 774.9844
66 | 16544 | 775.5
66 | 16531 | 774.8906
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16539 | 775.2656
66 | 16542 | 775.4062
66 | 16544 | 775.5
66 | 16536 | 775.125
66 | 16547 | 775.6406
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16546 | 775.5938
66 | 16537 | 775.1719
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16533 | 774.9844
66 | 16544 | 775.5
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16533 | 774.9844
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16531 | 774.8906
66 | 16542 | 775.4062
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16536 | 775.125
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16534 | 775.0312
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16542 | 775.4062
66 | 16532 | 774.9375
66 | 16544 | 775.5
66 | 16531 | 774.8906
66 | 16541 | 775.3594
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16537 | 775.1719
66 | 16545 | 775.5469
66 | 16537 | 775.1719
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16535 | 775.0781
66 | 16546 | 775.5938
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16532 | 774.9375
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16534 | 775.0312
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16545 | 775.5469
66 | 16536 | 775.125
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16545 | 775.5469
66 | 16535 | 775.0781
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16534 | 775.0312
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16538 | 775.2188
66 | 16541 | 775.3594
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16535 | 775.0781
66 | 16543 | 775.4531
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16542 | 775.4062
66 | 16534 | 775.0312
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16532 | 774.9375
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16545 | 775.5469
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16541 | 775.3594
66 | 16531 | 774.8906
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16539 | 775.2656
66 | 16532 | 774.9375
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16534 | 775.0312
66 | 16544 | 775.5
66 | 16534 | 775.0312
66 | 16543 | 775.4531
66 | 16539 | 775.2656
66 | 16542 | 775.4062
66 | 16541 | 775.3594
66 | 16541 | 775.3594
66 | 16537 | 775.1719
66 | 16541 | 775.3594
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16536 | 775.125
66 | 16542 | 775.4062
66 | 16535 | 775.0781
66 | 16542 | 775.4062
66 | 16538 | 775.2188
66 | 16544 | 775.5
66 | 16537 | 775.1719
66 | 16543 | 775.4531
66 | 16535 | 775.0781
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16539 | 775.2656
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16533 | 774.9844
66 | 16541 | 775.3594
66 | 16534 | 775.0312
66 | 16541 | 775.3594
66 | 16541 | 775.3594
66 | 16540 | 775.3125
66 | 16542 | 775.4062
66 | 16541 | 775.3594
66 | 16542 | 775.4062
66 | 16539 | 775.2656
66 | 16540 | 775.3125
66 | 16540 | 775.3125
66 | 16539 | 775.2656
66 | 16539 | 775.2656
66 | 16538 | 775.2188
66 | 16540 | 775.3125
66 | 16537 | 775.1719
66 | 16541 | 775.3594
66 | 16537 | 775.1719
66 | 16542 | 775.4062
66 | 16541 | 775.3594
66 | 16541 | 775.3594
66 | 16542 | 775.4062
66 | 16540 | 775.3125
66 | 16544 | 775.5
66 | 16539 | 775.2656
66 | 16543 | 775.4531
66 | 16536 | 775.125
66 | 16543 | 775.4531
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16531 | 774.8906
66 | 16540 | 775.3125
66 | 16537 | 775.1719
66 | 16539 | 775.2656
66 | 16539 | 775.2656
66 | 16538 | 775.2188
66 | 16541 | 775.3594
66 | 16537 | 775.1719
66 | 16542 | 775.4062
66 | 16536 | 775.125
66 | 16544 | 775.5
66 | 16532 | 774.9375
66 | 16543 | 775.4531
66 | 16531 | 774.8906
66 | 16542 | 775.4062
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16536 | 775.125
66 | 16539 | 775.2656
66 | 16537 | 775.1719
66 | 16540 | 775.3125
66 | 16536 | 775.125
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16539 | 775.2656
66 | 16539 | 775.2656
66 | 16541 | 775.3594
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16534 | 775.0312
66 | 16540 | 775.3125
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16542 | 775.4062
66 | 16539 | 775.2656
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16541 | 775.3594
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16538 | 775.2188
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16540 | 775.3125
66 | 16533 | 774.9844
66 | 16541 | 775.3594
66 | 16539 | 775.2656
66 | 16539 | 775.2656
66 | 16540 | 775.3125
66 | 16538 | 775.2188
66 | 16542 | 775.4062
66 | 16536 | 775.125
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16541 | 775.3594
66 | 16531 | 774.8906
66 | 16537 | 775.1719
66 | 16537 | 775.1719
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16535 | 775.0781
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16533 | 774.9844
66 | 16543 | 775.4531
66 | 16531 | 774.8906
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16534 | 775.0312
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16540 | 775.3125
66 | 16528 | 774.75
66 | 16538 | 775.2188
66 | 16535 | 775.0781
66 | 16539 | 775.2656
66 | 16540 | 775.3125
66 | 16534 | 775.0312
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16542 | 775.4062
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16535 | 775.0781
66 | 16537 | 775.1719
66 | 16539 | 775.2656
66 | 16537 | 775.1719
66 | 16541 | 775.3594
66 | 16534 | 775.0312
66 | 16541 | 775.3594
66 | 16533 | 774.9844
66 | 16541 | 775.3594
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16540 | 775.3125
66 | 16533 | 774.9844
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16534 | 775.0312
66 | 16541 | 775.3594
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16538 | 775.2188
66 | 16538 | 775.2188
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16541 | 775.3594
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16538 | 775.2188
66 | 16533 | 774.9844
66 | 16538 | 775.2188
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16533 | 774.9844
66 | 16541 | 775.3594
66 | 16532 | 774.9375
66 | 16544 | 775.5
66 | 16531 | 774.8906
66 | 16539 | 775.2656
66 | 16538 | 775.2188
66 | 16537 | 775.1719
66 | 16539 | 775.2656
66 | 16532 | 774.9375
66 | 16539 | 775.2656
66 | 16531 | 774.8906
66 | 16539 | 775.2656
66 | 16528 | 774.75
66 | 16538 | 775.2188
66 | 16538 | 775.2188
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16542 | 775.4062
66 | 16531 | 774.8906
66 | 16541 | 775.3594
66 | 16538 | 775.2188
66 | 16537 | 775.1719
66 | 16538 | 775.2188
66 | 16533 | 774.9844
66 | 16538 | 775.2188
66 | 16531 | 774.8906
66 | 16539 | 775.2656
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16534 | 775.0312
66 | 16537 | 775.1719
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16536 | 775.125
66 | 16539 | 775.2656
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16538 | 775.2188
66 | 16530 | 774.8438
66 | 16536 | 775.125
66 | 16537 | 775.1719
66 | 16535 | 775.0781
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16539 | 775.2656
66 | 16540 | 775.3125
66 | 16534 | 775.0312
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16536 | 775.125
66 | 16537 | 775.1719
66 | 16534 | 775.0312
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16538 | 775.2188
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16542 | 775.4062
66 | 16531 | 774.8906
66 | 16539 | 775.2656
66 | 16535 | 775.0781
66 | 16536 | 775.125
66 | 16537 | 775.1719
66 | 16533 | 774.9844
66 | 16537 | 775.1719
66 | 16532 | 774.9375
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16538 | 775.2188
66 | 16539 | 775.2656
66 | 16536 | 775.125
66 | 16541 | 775.3594
66 | 16531 | 774.8906
66 | 16541 | 775.3594
66 | 16531 | 774.8906
66 | 16539 | 775.2656
66 | 16528 | 774.75
66 | 16536 | 775.125
66 | 16536 | 775.125
66 | 16533 | 774.9844
66 | 16537 | 775.1719
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16539 | 775.2656
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16538 | 775.2188
66 | 16534 | 775.0312
66 | 16539 | 775.2656
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16528 | 774.75
66 | 16536 | 775.125
66 | 16537 | 775.1719
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16535 | 775.0781
66 | 16536 | 775.125
66 | 16536 | 775.125
66 | 16538 | 775.2188
66 | 16535 | 775.0781
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16541 | 775.3594
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16538 | 775.2188
66 | 16534 | 775.0312
66 | 16539 | 775.2656
66 | 16530 | 774.8438
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16537 | 775.1719
66 | 16532 | 774.9375
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16537 | 775.1719
66 | 16536 | 775.125
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16528 | 774.75
66 | 16536 | 775.125
66 | 16536 | 775.125
66 | 16534 | 775.0312
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16540 | 775.3125
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16537 | 775.1719
66 | 16534 | 775.0312
66 | 16540 | 775.3125
66 | 16531 | 774.8906
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16537 | 775.1719
66 | 16539 | 775.2656
66 | 16532 | 774.9375
66 | 16538 | 775.2188
66 | 16527 | 774.7031
66 | 16535 | 775.0781
66 | 16536 | 775.125
66 | 16530 | 774.8438
66 | 16539 | 775.2656
66 | 16526 | 774.6562
66 | 16537 | 775.1719
66 | 16529 | 774.7969
66 | 16536 | 775.125
66 | 16539 | 775.2656
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16539 | 775.2656
66 | 16529 | 774.7969
66 | 16536 | 775.125
66 | 16535 | 775.0781
66 | 16532 | 774.9375
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16535 | 775.0781
66 | 16535 | 775.0781
66 | 16538 | 775.2188
66 | 16533 | 774.9844
66 | 16540 | 775.3125
66 | 16532 | 774.9375
66 | 16539 | 775.2656
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16537 | 775.1719
66 | 16529 | 774.7969
66 | 16536 | 775.125
66 | 16539 | 775.2656
66 | 16532 | 774.9375
66 | 16540 | 775.3125
66 | 16530 | 774.8438
66 | 16540 | 775.3125
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16537 | 775.1719
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16538 | 775.2188
66 | 16531 | 774.8906
66 | 16539 | 775.2656
66 | 16530 | 774.8438
66 | 16537 | 775.1719
66 | 16537 | 775.1719
66 | 16532 | 774.9375
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16535 | 775.0781
66 | 16525 | 774.6094
66 | 16534 | 775.0312
66 | 16535 | 775.0781
66 | 16531 | 774.8906
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16528 | 774.75
66 | 16537 | 775.1719
66 | 16535 | 775.0781
66 | 16534 | 775.0312
66 | 16537 | 775.1719
66 | 16528 | 774.75
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16534 | 775.0312
66 | 16527 | 774.7031
66 | 16535 | 775.0781
66 | 16524 | 774.5625
66 | 16535 | 775.0781
66 | 16534 | 775.0312
66 | 16531 | 774.8906
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16538 | 775.2188
66 | 16529 | 774.7969
66 | 16534 | 775.0312
66 | 16537 | 775.1719
66 | 16528 | 774.75
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16536 | 775.125
66 | 16528 | 774.75
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16536 | 775.125
66 | 16530 | 774.8438
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16536 | 775.125
66 | 16532 | 774.9375
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16535 | 775.0781
66 | 16536 | 775.125
66 | 16530 | 774.8438
66 | 16536 | 775.125
66 | 16525 | 774.6094
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16535 | 775.0781
66 | 16529 | 774.7969
66 | 16537 | 775.1719
66 | 16536 | 775.125
66 | 16530 | 774.8438
66 | 16538 | 775.2188
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16534 | 775.0312
66 | 16527 | 774.7031
66 | 16537 | 775.1719
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16535 | 775.0781
66 | 16526 | 774.6562
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16532 | 774.9375
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16528 | 774.75
66 | 16533 | 774.9844
66 | 16536 | 775.125
66 | 16525 | 774.6094
66 | 16536 | 775.125
66 | 16523 | 774.5156
66 | 16531 | 774.8906
66 | 16533 | 774.9844
66 | 16526 | 774.6562
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16534 | 775.0312
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16538 | 775.2188
66 | 16527 | 774.7031
66 | 16534 | 775.0312
66 | 16536 | 775.125
66 | 16525 | 774.6094
66 | 16534 | 775.0312
66 | 16524 | 774.5625
66 | 16531 | 774.8906
66 | 16534 | 775.0312
66 | 16525 | 774.6094
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16531 | 774.8906
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16532 | 774.9375
66 | 16533 | 774.9844
66 | 16524 | 774.5625
66 | 16534 | 775.0312
66 | 16524 | 774.5625
66 | 16533 | 774.9844
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16535 | 775.0781
66 | 16535 | 775.0781
66 | 16526 | 774.6562
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16533 | 774.9844
66 | 16536 | 775.125
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16532 | 774.9375
66 | 16524 | 774.5625
66 | 16534 | 775.0312
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16536 | 775.125
66 | 16525 | 774.6094
66 | 16534 | 775.0312
66 | 16535 | 775.0781
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16536 | 775.125
66 | 16525 | 774.6094
66 | 16533 | 774.9844
66 | 16534 | 775.0312
66 | 16524 | 774.5625
66 | 16531 | 774.8906
66 | 16535 | 775.0781
66 | 16526 | 774.6562
66 | 16531 | 774.8906
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16535 | 775.0781
66 | 16526 | 774.6562
66 | 16534 | 775.0312
66 | 16529 | 774.7969
66 | 16530 | 774.8438
66 | 16534 | 775.0312
66 | 16526 | 774.6562
66 | 16531 | 774.8906
66 | 16534 | 775.0312
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16533 | 774.9844
66 | 16523 | 774.5156
66 | 16531 | 774.8906
66 | 16535 | 775.0781
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16535 | 775.0781
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16536 | 775.125
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16533 | 774.9844
66 | 16525 | 774.6094
66 | 16530 | 774.8438
66 | 16532 | 774.9375
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16533 | 774.9844
66 | 16537 | 775.1719
66 | 16527 | 774.7031
66 | 16531 | 774.8906
66 | 16536 | 775.125
66 | 16526 | 774.6562
66 | 16532 | 774.9375
66 | 16534 | 775.0312
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16534 | 775.0312
66 | 16525 | 774.6094
66 | 16530 | 774.8438
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16529 | 774.7969
66 | 16536 | 775.125
66 | 16528 | 774.75
66 | 16532 | 774.9375
66 | 16537 | 775.1719
66 | 16526 | 774.6562
66 | 16531 | 774.8906
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16530 | 774.8438
66 | 16533 | 774.9844
66 | 16524 | 774.5625
66 | 16528 | 774.75
66 | 16534 | 775.0312
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16534 | 775.0312
66 | 16526 | 774.6562
66 | 16531 | 774.8906
66 | 16535 | 775.0781
66 | 16528 | 774.75
66 | 16530 | 774.8438
66 | 16535 | 775.0781
66 | 16527 | 774.7031
66 | 16527 | 774.7031
66 | 16534 | 775.0312
66 | 16526 | 774.6562
66 | 16527 | 774.7031
66 | 16533 | 774.9844
66 | 16527 | 774.7031
66 | 16528 | 774.75
66 | 16535 | 775.0781
66 | 16529 | 774.7969
66 | 16530 | 774.8438
66 | 16534 | 775.0312
66 | 16527 | 774.7031
66 | 16529 | 774.7969
66 | 16535 | 775.0781
66 | 16526 | 774.6562
66 | 16529 | 774.7969
66 | 16533 | 774.9844
66 | 16523 | 774.5156
66 | 16527 | 774.7031
66 | 16531 | 774.8906
66 | 16523 | 774.5156
66 | 16528 | 774.75
66 | 16533 | 774.9844
66 | 16523 | 774.5156
66 | 16529 | 774.7969
66 | 16535 | 775.0781
66 | 16528 | 774.75
66 | 16531 | 774.8906
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16529 | 774.7969
66 | 16534 | 775.0312
66 | 16527 | 774.7031
66 | 16528 | 774.75
66 | 16532 | 774.9375
66 | 16522 | 774.4688
66 | 16528 | 774.75
66 | 16534 | 775.0312
66 | 16525 | 774.6094
66 | 16528 | 774.75
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16527 | 774.7031
66 | 16536 | 775.125
66 | 16529 | 774.7969
66 | 16530 | 774.8438
66 | 16534 | 775.0312
66 | 16529 | 774.7969
66 | 16527 | 774.7031
66 | 16533 | 774.9844
66 | 16525 | 774.6094
66 | 16528 | 774.75
66 | 16533 | 774.9844
66 | 16526 | 774.6562
66 | 16527 | 774.7031
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16527 | 774.7031
66 | 16535 | 775.0781
66 | 16528 | 774.75
66 | 16530 | 774.8438
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16527 | 774.7031
66 | 16534 | 775.0312
66 | 16525 | 774.6094
66 | 16526 | 774.6562
66 | 16532 | 774.9375
66 | 16528 | 774.75
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16524 | 774.5625
66 | 16528 | 774.75
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16530 | 774.8438
66 | 16536 | 775.125
66 | 16527 | 774.7031
66 | 16528 | 774.75
66 | 16533 | 774.9844
66 | 16525 | 774.6094
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16525 | 774.6094
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16525 | 774.6094
66 | 16528 | 774.75
66 | 16533 | 774.9844
66 | 16526 | 774.6562
66 | 16528 | 774.75
66 | 16535 | 775.0781
66 | 16528 | 774.75
66 | 16529 | 774.7969
66 | 16533 | 774.9844
66 | 16525 | 774.6094
66 | 16527 | 774.7031
66 | 16533 | 774.9844
66 | 16526 | 774.6562
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16525 | 774.6094
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16527 | 774.7031
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16526 | 774.6562
66 | 16527 | 774.7031
66 | 16534 | 775.0312
66 | 16528 | 774.75
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16525 | 774.6094
66 | 16526 | 774.6562
66 | 16531 | 774.8906
66 | 16526 | 774.6562
66 | 16526 | 774.6562
66 | 16532 | 774.9375
66 | 16524 | 774.5625
66 | 16524 | 774.5625
66 | 16532 | 774.9375
66 | 16526 | 774.6562
66 | 16528 | 774.75
66 | 16534 | 775.0312
66 | 16526 | 774.6562
66 | 16526 | 774.6562
66 | 16532 | 774.9375
66 | 16525 | 774.6094
66 | 16527 | 774.7031
66 | 16531 | 774.8906
66 | 16523 | 774.5156
66 | 16524 | 774.5625
66 | 16531 | 774.8906
66 | 16523 | 774.5156
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16523 | 774.5156
66 | 16526 | 774.6562
66 | 16533 | 774.9844
66 | 16524 | 774.5625
66 | 16528 | 774.75
66 | 16533 | 774.9844
66 | 16525 | 774.6094
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16523 | 774.5156
66 | 16525 | 774.6094
66 | 16530 | 774.8438
66 | 16525 | 774.6094
66 | 16524 | 774.5625
66 | 16532 | 774.9375
66 | 16524 | 774.5625
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16523 | 774.5156
66 | 16527 | 774.7031
66 | 16533 | 774.9844
66 | 16525 | 774.6094
66 | 16527 | 774.7031
66 | 16531 | 774.8906
66 | 16522 | 774.4688
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16522 | 774.4688
66 | 16525 | 774.6094
66 | 16529 | 774.7969
66 | 16524 | 774.5625
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16525 | 774.6094
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16524 | 774.5625
66 | 16527 | 774.7031
66 | 16532 | 774.9375
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16521 | 774.4219
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16525 | 774.6094
66 | 16523 | 774.5156
66 | 16531 | 774.8906
66 | 16528 | 774.75
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16524 | 774.5625
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16521 | 774.4219
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16522 | 774.4688
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16524 | 774.5625
66 | 16525 | 774.6094
66 | 16530 | 774.8438
66 | 16519 | 774.3281
66 | 16524 | 774.5625
66 | 16528 | 774.75
66 | 16520 | 774.375
66 | 16524 | 774.5625
66 | 16529 | 774.7969
66 | 16520 | 774.375
66 | 16524 | 774.5625
66 | 16529 | 774.7969
66 | 16522 | 774.4688
66 | 16525 | 774.6094
66 | 16532 | 774.9375
66 | 16524 | 774.5625
66 | 16526 | 774.6562
66 | 16530 | 774.8438
66 | 16523 | 774.5156
66 | 16524 | 774.5625
66 | 16530 | 774.8438
66 | 16521 | 774.4219
66 | 16524 | 774.5625
66 | 16529 | 774.7969
66 | 16522 | 774.4688
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16520 | 774.375
66 | 16525 | 774.6094
66 | 16530 | 774.8438
66 | 16523 | 774.5156
66 | 16524 | 774.5625
66 | 16529 | 774.7969
66 | 16520 | 774.375
66 | 16523 | 774.5156
66 | 16529 | 774.7969
66 | 16522 | 774.4688
66 | 16522 | 774.4688
66 | 16527 | 774.7031
66 | 16520 | 774.375
66 | 16522 | 774.4688
66 | 16529 | 774.7969
66 | 16523 | 774.5156
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16523 | 774.5156
66 | 16525 | 774.6094
66 | 16531 | 774.8906
66 | 16523 | 774.5156
66 | 16524 | 774.5625
66 | 16528 | 774.75
66 | 16522 | 774.4688
66 | 16520 | 774.375
66 | 16528 | 774.75
66 | 16521 | 774.4219
66 | 16522 | 774.4688
66 | 16529 | 774.7969
66 | 16522 | 774.4688
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16524 | 774.5625
66 | 16531 | 774.8906
66 | 16522 | 774.4688
66 | 16523 | 774.5156
66 | 16529 | 774.7969
66 | 16524 | 774.5625
66 | 16522 | 774.4688
66 | 16528 | 774.75
66 | 16523 | 774.5156
66 | 16521 | 774.4219
66 | 16528 | 774.75
66 | 16524 | 774.5625
66 | 16525 | 774.6094
66 | 16528 | 774.75
66 | 16526 | 774.6562
66 | 16521 | 774.4219
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16522 | 774.4688
66 | 16528 | 774.75
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16527 | 774.7031
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16528 | 774.75
66 | 16525 | 774.6094
66 | 16519 | 774.3281
66 | 16529 | 774.7969
66 | 16524 | 774.5625
66 | 16523 | 774.5156
66 | 16529 | 774.7969
66 | 16523 | 774.5156
66 | 16521 | 774.4219
66 | 16529 | 774.7969
66 | 16523 | 774.5156
66 | 16521 | 774.4219
66 | 16527 | 774.7031
66 | 16523 | 774.5156
66 | 16519 | 774.3281
66 | 16528 | 774.75
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16529 | 774.7969
66 | 16524 | 774.5625
66 | 16523 | 774.5156
66 | 16530 | 774.8438
66 | 16527 | 774.7031
66 | 16523 | 774.5156
66 | 16528 | 774.75
66 | 16522 | 774.4688
66 | 16520 | 774.375
66 | 16528 | 774.75
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16527 | 774.7031
66 | 16524 | 774.5625
66 | 16519 | 774.3281
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16529 | 774.7969
66 | 16525 | 774.6094
66 | 16522 | 774.4688
66 | 16529 | 774.7969
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16527 | 774.7031
66 | 16521 | 774.4219
66 | 16520 | 774.375
66 | 16526 | 774.6562
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16528 | 774.75
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16529 | 774.7969
66 | 16525 | 774.6094
66 | 16522 | 774.4688
66 | 16529 | 774.7969
66 | 16524 | 774.5625
66 | 16518 | 774.2812
66 | 16527 | 774.7031
66 | 16525 | 774.6094
66 | 16519 | 774.3281
66 | 16525 | 774.6094
66 | 16524 | 774.5625
66 | 16518 | 774.2812
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16520 | 774.375
66 | 16527 | 774.7031
66 | 16526 | 774.6562
66 | 16519 | 774.3281
66 | 16527 | 774.7031
66 | 16526 | 774.6562
66 | 16521 | 774.4219
66 | 16528 | 774.75
66 | 16524 | 774.5625
66 | 16517 | 774.2344
66 | 16527 | 774.7031
66 | 16522 | 774.4688
66 | 16519 | 774.3281
66 | 16527 | 774.7031
66 | 16521 | 774.4219
66 | 16520 | 774.375
66 | 16530 | 774.8438
66 | 16524 | 774.5625
66 | 16520 | 774.375
66 | 16530 | 774.8438
66 | 16525 | 774.6094
66 | 16521 | 774.4219
66 | 16529 | 774.7969
66 | 16523 | 774.5156
66 | 16521 | 774.4219
66 | 16526 | 774.6562
66 | 16523 | 774.5156
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16518 | 774.2812
66 | 16525 | 774.6094
66 | 16524 | 774.5625
66 | 16519 | 774.3281
66 | 16528 | 774.75
66 | 16526 | 774.6562
66 | 16520 | 774.375
66 | 16529 | 774.7969
66 | 16525 | 774.6094
66 | 16517 | 774.2344
66 | 16523 | 774.5156
66 | 16527 | 774.7031
66 | 16522 | 774.4688
66 | 16518 | 774.2812
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16520 | 774.375
66 | 16526 | 774.6562
66 | 16521 | 774.4219
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16525 | 774.6094
66 | 16520 | 774.375
66 | 16528 | 774.75
66 | 16522 | 774.4688
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16521 | 774.4219
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16520 | 774.375
66 | 16518 | 774.2812
66 | 16527 | 774.7031
66 | 16522 | 774.4688
66 | 16519 | 774.3281
66 | 16527 | 774.7031
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16527 | 774.7031
66 | 16525 | 774.6094
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16522 | 774.4688
66 | 16516 | 774.1875
66 | 16525 | 774.6094
66 | 16522 | 774.4688
66 | 16519 | 774.3281
66 | 16526 | 774.6562
66 | 16522 | 774.4688
66 | 16516 | 774.1875
66 | 16525 | 774.6094
66 | 16526 | 774.6562
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16527 | 774.7031
66 | 16522 | 774.4688
66 | 16520 | 774.375
66 | 16525 | 774.6094
66 | 16522 | 774.4688
66 | 16516 | 774.1875
66 | 16523 | 774.5156
66 | 16523 | 774.5156
66 | 16516 | 774.1875
66 | 16524 | 774.5625
66 | 16520 | 774.375
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16518 | 774.2812
66 | 16526 | 774.6562
66 | 16523 | 774.5156
66 | 16517 | 774.2344
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16516 | 774.1875
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16514 | 774.0938
66 | 16524 | 774.5625
66 | 16523 | 774.5156
66 | 16514 | 774.0938
66 | 16520 | 774.375
66 | 16527 | 774.7031
66 | 16525 | 774.6094
66 | 16517 | 774.2344
66 | 16522 | 774.4688
66 | 16528 | 774.75
66 | 16522 | 774.4688
66 | 16518 | 774.2812
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16514 | 774.0938
66 | 16522 | 774.4688
66 | 16524 | 774.5625
66 | 16520 | 774.375
66 | 16515 | 774.1406
66 | 16526 | 774.6562
66 | 16522 | 774.4688
66 | 16514 | 774.0938
66 | 16519 | 774.3281
66 | 16526 | 774.6562
66 | 16524 | 774.5625
66 | 16516 | 774.1875
66 | 16523 | 774.5156
66 | 16525 | 774.6094
66 | 16519 | 774.3281
66 | 16514 | 774.0938
66 | 16521 | 774.4219
66 | 16522 | 774.4688
66 | 16515 | 774.1406
66 | 16516 | 774.1875
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16515 | 774.1406
66 | 16525 | 774.6094
66 | 16524 | 774.5625
66 | 16516 | 774.1875
66 | 16519 | 774.3281
66 | 16526 | 774.6562
66 | 16521 | 774.4219
66 | 16516 | 774.1875
66 | 16523 | 774.5156
66 | 16524 | 774.5625
66 | 16521 | 774.4219
66 | 16514 | 774.0938
66 | 16521 | 774.4219
66 | 16523 | 774.5156
66 | 16522 | 774.4688
66 | 16514 | 774.0938
66 | 16522 | 774.4688
66 | 16526 | 774.6562
66 | 16522 | 774.4688
66 | 16515 | 774.1406
66 | 16524 | 774.5625
66 | 16525 | 774.6094
66 | 16521 | 774.4219
66 | 16514 | 774.0938
66 | 16517 | 774.2344
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16511 | 773.9531
66 | 16516 | 774.1875
66 | 16522 | 774.4688
66 | 16520 | 774.375
66 | 16514 | 774.0938
66 | 16520 | 774.375
66 | 16525 | 774.6094
66 | 16523 | 774.5156
66 | 16514 | 774.0938
66 | 16518 | 774.2812
66 | 16525 | 774.6094
66 | 16520 | 774.375
66 | 16513 | 774.0469
66 | 16520 | 774.375
66 | 16523 | 774.5156
66 | 16518 | 774.2812
66 | 16512 | 774
66 | 16521 | 774.4219
66 | 16523 | 774.5156
66 | 16518 | 774.2812
66 | 16515 | 774.1406
66 | 16523 | 774.5156
66 | 16522 | 774.4688
66 | 16514 | 774.0938
66 | 16522 | 774.4688
66 | 16525 | 774.6094
66 | 16521 | 774.4219
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16521 | 774.4219
66 | 16521 | 774.4219
66 | 16512 | 774
66 | 16517 | 774.2344
66 | 16524 | 774.5625
66 | 16522 | 774.4688
66 | 16514 | 774.0938
66 | 16519 | 774.3281
66 | 16524 | 774.5625
66 | 16522 | 774.4688
66 | 16515 | 774.1406
66 | 16516 | 774.1875
66 | 16522 | 774.4688
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16514 | 774.0938
66 | 16521 | 774.4219
66 | 16523 | 774.5156
66 | 16521 | 774.4219
66 | 16513 | 774.0469
66 | 16515 | 774.1406
66 | 16522 | 774.4688
66 | 16523 | 774.5156
66 | 16515 | 774.1406
66 | 16517 | 774.2344
66 | 16523 | 774.5156
66 | 16525 | 774.6094
66 | 16521 | 774.4219
66 | 16513 | 774.0469
66 | 16515 | 774.1406
66 | 16521 | 774.4219
66 | 16522 | 774.4688
66 | 16518 | 774.2812
66 | 16512 | 774
66 | 16515 | 774.1406
66 | 16521 | 774.4219
66 | 16521 | 774.4219
66 | 16513 | 774.0469
66 | 16517 | 774.2344
66 | 16522 | 774.4688
66 | 16522 | 774.4688
66 | 16513 | 774.0469
66 | 16517 | 774.2344
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16520 | 774.375
66 | 16522 | 774.4688
66 | 16510 | 773.9062
66 | 16514 | 774.0938
66 | 16520 | 774.375
66 | 16521 | 774.4219
66 | 16513 | 774.0469
66 | 16514 | 774.0938
66 | 16521 | 774.4219
66 | 16521 | 774.4219
66 | 16514 | 774.0938
66 | 16514 | 774.0938
66 | 16522 | 774.4688
66 | 16522 | 774.4688
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16520 | 774.375
66 | 16520 | 774.375
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16522 | 774.4688
66 | 16522 | 774.4688
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16522 | 774.4688
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16513 | 774.0469
66 | 16516 | 774.1875
66 | 16521 | 774.4219
66 | 16519 | 774.3281
66 | 16511 | 773.9531
66 | 16513 | 774.0469
66 | 16518 | 774.2812
66 | 16520 | 774.375
66 | 16516 | 774.1875
66 | 16512 | 774
66 | 16520 | 774.375
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16513 | 774.0469
66 | 16516 | 774.1875
66 | 16521 | 774.4219
66 | 16522 | 774.4688
66 | 16512 | 774
66 | 16516 | 774.1875
66 | 16522 | 774.4688
66 | 16521 | 774.4219
66 | 16516 | 774.1875
66 | 16509 | 773.8594
66 | 16518 | 774.2812
66 | 16521 | 774.4219
66 | 16518 | 774.2812
66 | 16511 | 773.9531
66 | 16514 | 774.0938
66 | 16520 | 774.375
66 | 16522 | 774.4688
66 | 16518 | 774.2812
66 | 16513 | 774.0469
66 | 16520 | 774.375
66 | 16523 | 774.5156
66 | 16520 | 774.375
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16521 | 774.4219
66 | 16519 | 774.3281
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16522 | 774.4688
66 | 16519 | 774.3281
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16521 | 774.4219
66 | 16522 | 774.4688
66 | 16519 | 774.3281
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16519 | 774.3281
66 | 16517 | 774.2344
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16519 | 774.3281
66 | 16521 | 774.4219
66 | 16518 | 774.2812
66 | 16512 | 774
66 | 16515 | 774.1406
66 | 16520 | 774.375
66 | 16522 | 774.4688
66 | 16519 | 774.3281
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16518 | 774.2812
66 | 16520 | 774.375
66 | 16516 | 774.1875
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16518 | 774.2812
66 | 16521 | 774.4219
66 | 16519 | 774.3281
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16520 | 774.375
66 | 16520 | 774.375
66 | 16520 | 774.375
66 | 16511 | 773.9531
66 | 16513 | 774.0469
66 | 16518 | 774.2812
66 | 16521 | 774.4219
66 | 16518 | 774.2812
66 | 16508 | 773.8125
66 | 16510 | 773.9062
66 | 16516 | 774.1875
66 | 16518 | 774.2812
66 | 16519 | 774.3281
66 | 16513 | 774.0469
66 | 16510 | 773.9062
66 | 16519 | 774.3281
66 | 16520 | 774.375
66 | 16520 | 774.375
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16519 | 774.3281
66 | 16519 | 774.3281
66 | 16517 | 774.2344
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16516 | 774.1875
66 | 16518 | 774.2812
66 | 16519 | 774.3281
66 | 16512 | 774
66 | 16509 | 773.8594
66 | 16518 | 774.2812
66 | 16518 | 774.2812
66 | 16519 | 774.3281
66 | 16515 | 774.1406
66 | 16512 | 774
66 | 16517 | 774.2344
66 | 16520 | 774.375
66 | 16518 | 774.2812
66 | 16512 | 774
66 | 16509 | 773.8594
66 | 16516 | 774.1875
66 | 16517 | 774.2344
66 | 16518 | 774.2812
66 | 16514 | 774.0938
66 | 16508 | 773.8125
66 | 16516 | 774.1875
66 | 16520 | 774.375
66 | 16519 | 774.3281
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16520 | 774.375
66 | 16517 | 774.2344
66 | 16511 | 773.9531
66 | 16510 | 773.9062
66 | 16515 | 774.1406
66 | 16519 | 774.3281
66 | 16517 | 774.2344
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16516 | 774.1875
66 | 16520 | 774.375
66 | 16518 | 774.2812
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16518 | 774.2812
66 | 16521 | 774.4219
66 | 16519 | 774.3281
66 | 16510 | 773.9062
66 | 16511 | 773.9531
66 | 16518 | 774.2812
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16515 | 774.1406
66 | 16518 | 774.2812
66 | 16517 | 774.2344
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16519 | 774.3281
66 | 16520 | 774.375
66 | 16519 | 774.3281
66 | 16515 | 774.1406
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16516 | 774.1875
66 | 16517 | 774.2344
66 | 16518 | 774.2812
66 | 16516 | 774.1875
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16515 | 774.1406
66 | 16517 | 774.2344
66 | 16519 | 774.3281
66 | 16519 | 774.3281
66 | 16512 | 774
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16516 | 774.1875
66 | 16518 | 774.2812
66 | 16519 | 774.3281
66 | 16516 | 774.1875
66 | 16514 | 774.0938
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16513 | 774.0469
66 | 16515 | 774.1406
66 | 16517 | 774.2344
66 | 16517 | 774.2344
66 | 16515 | 774.1406
66 | 16511 | 773.9531
66 | 16510 | 773.9062
66 | 16513 | 774.0469
66 | 16517 | 774.2344
66 | 16518 | 774.2812
66 | 16520 | 774.375
66 | 16518 | 774.2812
66 | 16511 | 773.9531
66 | 16507 | 773.7656
66 | 16509 | 773.8594
66 | 16514 | 774.0938
66 | 16516 | 774.1875
66 | 16517 | 774.2344
66 | 16518 | 774.2812
66 | 16516 | 774.1875
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16511 | 773.9531
66 | 16515 | 774.1406
66 | 16519 | 774.3281
66 | 16520 | 774.375
66 | 16517 | 774.2344
66 | 16512 | 774
66 | 16508 | 773.8125
66 | 16509 | 773.8594
66 | 16514 | 774.0938
66 | 16515 | 774.1406
66 | 16516 | 774.1875
66 | 16515 | 774.1406
66 | 16509 | 773.8594
66 | 16507 | 773.7656
66 | 16509 | 773.8594
66 | 16517 | 774.2344
66 | 16518 | 774.2812
66 | 16519 | 774.3281
66 | 16517 | 774.2344
66 | 16513 | 774.0469
66 | 16508 | 773.8125
66 | 16509 | 773.8594
66 | 16515 | 774.1406
66 | 16517 | 774.2344
66 | 16516 | 774.1875
66 | 16515 | 774.1406
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16508 | 773.8125
66 | 16513 | 774.0469
66 | 16516 | 774.1875
66 | 16518 | 774.2812
66 | 16517 | 774.2344
66 | 16517 | 774.2344
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16514 | 774.0938
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16508 | 773.8125
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16514 | 774.0938
66 | 16516 | 774.1875
66 | 16518 | 774.2812
66 | 16517 | 774.2344
66 | 16515 | 774.1406
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16513 | 774.0469
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16513 | 774.0469
66 | 16509 | 773.8594
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16512 | 774
66 | 16516 | 774.1875
66 | 16517 | 774.2344
66 | 16518 | 774.2812
66 | 16515 | 774.1406
66 | 16512 | 774
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16510 | 773.9062
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16514 | 774.0938
66 | 16515 | 774.1406
66 | 16513 | 774.0469
66 | 16508 | 773.8125
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16509 | 773.8594
66 | 16515 | 774.1406
66 | 16515 | 774.1406
66 | 16517 | 774.2344
66 | 16516 | 774.1875
66 | 16511 | 773.9531
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16512 | 774
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16512 | 774
66 | 16507 | 773.7656
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16515 | 774.1406
66 | 16516 | 774.1875
66 | 16515 | 774.1406
66 | 16509 | 773.8594
66 | 16504 | 773.625
66 | 16506 | 773.7188
66 | 16510 | 773.9062
66 | 16513 | 774.0469
66 | 16512 | 774
66 | 16512 | 774
66 | 16513 | 774.0469
66 | 16508 | 773.8125
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16511 | 773.9531
66 | 16515 | 774.1406
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16514 | 774.0938
66 | 16507 | 773.7656
66 | 16504 | 773.625
66 | 16507 | 773.7656
66 | 16511 | 773.9531
66 | 16512 | 774
66 | 16513 | 774.0469
66 | 16515 | 774.1406
66 | 16514 | 774.0938
66 | 16509 | 773.8594
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16513 | 774.0469
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16515 | 774.1406
66 | 16513 | 774.0469
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16511 | 773.9531
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16514 | 774.0938
66 | 16511 | 773.9531
66 | 16508 | 773.8125
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16516 | 774.1875
66 | 16514 | 774.0938
66 | 16508 | 773.8125
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16514 | 774.0938
66 | 16514 | 774.0938
66 | 16514 | 774.0938
66 | 16514 | 774.0938
66 | 16510 | 773.9062
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16510 | 773.9062
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16512 | 774
66 | 16512 | 774
66 | 16513 | 774.0469
66 | 16511 | 773.9531
66 | 16503 | 773.5781
66 | 16508 | 773.8125
66 | 16505 | 773.6719
66 | 16513 | 774.0469
66 | 16504 | 773.625
66 | 16509 | 773.8594
66 | 16514 | 774.0938
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16508 | 773.8125
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16509 | 773.8594
66 | 16512 | 774
66 | 16512 | 774
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16515 | 774.1406
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16510 | 773.9062
66 | 16505 | 773.6719
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16508 | 773.8125
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16512 | 774
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16515 | 774.1406
66 | 16515 | 774.1406
66 | 16514 | 774.0938
66 | 16511 | 773.9531
66 | 16505 | 773.6719
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16514 | 774.0938
66 | 16513 | 774.0469
66 | 16511 | 773.9531
66 | 16512 | 774
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16511 | 773.9531
66 | 16512 | 774
66 | 16512 | 774
66 | 16512 | 774
66 | 16514 | 774.0938
66 | 16515 | 774.1406
66 | 16513 | 774.0469
66 | 16514 | 774.0938
66 | 16511 | 773.9531
66 | 16510 | 773.9062
66 | 16507 | 773.7656
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16506 | 773.7188
66 | 16509 | 773.8594
66 | 16509 | 773.8594
66 | 16507 | 773.7656
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16512 | 774
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16513 | 774.0469
66 | 16511 | 773.9531
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16504 | 773.625
66 | 16508 | 773.8125
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16511 | 773.9531
66 | 16512 | 774
66 | 16512 | 774
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16511 | 773.9531
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16508 | 773.8125
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16506 | 773.7188
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16511 | 773.9531
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16508 | 773.8125
66 | 16504 | 773.625
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16504 | 773.625
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16509 | 773.8594
66 | 16508 | 773.8125
66 | 16509 | 773.8594
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16504 | 773.625
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16509 | 773.8594
66 | 16509 | 773.8594
66 | 16509 | 773.8594
66 | 16510 | 773.9062
66 | 16510 | 773.9062
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16509 | 773.8594
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16504 | 773.625
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16508 | 773.8125
66 | 16508 | 773.8125
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16506 | 773.7188
66 | 16504 | 773.625
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16508 | 773.8125
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16508 | 773.8125
66 | 16506 | 773.7188
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16505 | 773.6719
66 | 16507 | 773.7656
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16506 | 773.7188
66 | 16506 | 773.7188
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16504 | 773.625
66 | 16502 | 773.5312
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16496 | 773.25
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16495 | 773.2031
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16495 | 773.2031
66 | 16496 | 773.25
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16493 | 773.1094
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16492 | 773.0625
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16497 | 773.2969
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16505 | 773.6719
66 | 16504 | 773.625
66 | 16504 | 773.625
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16502 | 773.5312
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16504 | 773.625
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16503 | 773.5781
66 | 16503 | 773.5781
66 | 16502 | 773.5312
66 | 16502 | 773.5312
66 | 16504 | 773.625
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16494 | 773.1562
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16489 | 772.9219
66 | 16490 | 772.9688
66 | 16491 | 773.0156
66 | 16491 | 773.0156
66 | 16491 | 773.0156
66 | 16493 | 773.1094
66 | 16494 | 773.1562
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16502 | 773.5312
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16492 | 773.0625
66 | 16492 | 773.0625
66 | 16490 | 772.9688
66 | 16490 | 772.9688
66 | 16489 | 772.9219
66 | 16490 | 772.9688
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16494 | 773.1562
66 | 16498 | 773.3438
66 | 16501 | 773.4844
66 | 16501 | 773.4844
66 | 16502 | 773.5312
66 | 16501 | 773.4844
66 | 16500 | 773.4375
66 | 16501 | 773.4844
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16495 | 773.2031
66 | 16491 | 773.0156
66 | 16489 | 772.9219
66 | 16490 | 772.9688
66 | 16489 | 772.9219
66 | 16489 | 772.9219
66 | 16488 | 772.875
66 | 16490 | 772.9688
66 | 16495 | 773.2031
66 | 16497 | 773.2969
66 | 16500 | 773.4375
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16500 | 773.4375
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16496 | 773.25
66 | 16492 | 773.0625
66 | 16493 | 773.1094
66 | 16490 | 772.9688
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16488 | 772.875
66 | 16489 | 772.9219
66 | 16493 | 773.1094
66 | 16496 | 773.25
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16495 | 773.2031
66 | 16491 | 773.0156
66 | 16490 | 772.9688
66 | 16488 | 772.875
66 | 16496 | 773.25
66 | 16498 | 773.3438
66 | 16500 | 773.4375
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16493 | 773.1094
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16486 | 772.7812
66 | 16487 | 772.8281
66 | 16491 | 773.0156
66 | 16496 | 773.25
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16489 | 772.9219
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16487 | 772.8281
66 | 16486 | 772.7812
66 | 16487 | 772.8281
66 | 16491 | 773.0156
66 | 16494 | 773.1562
66 | 16496 | 773.25
66 | 16496 | 773.25
66 | 16499 | 773.3906
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16498 | 773.3438
66 | 16495 | 773.2031
66 | 16493 | 773.1094
66 | 16488 | 772.875
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16489 | 772.9219
66 | 16491 | 773.0156
66 | 16495 | 773.2031
66 | 16499 | 773.3906
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16494 | 773.1562
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16489 | 772.9219
66 | 16492 | 773.0625
66 | 16497 | 773.2969
66 | 16499 | 773.3906
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16494 | 773.1562
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16487 | 772.8281
66 | 16490 | 772.9688
66 | 16493 | 773.1094
66 | 16498 | 773.3438
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16498 | 773.3438
66 | 16496 | 773.25
66 | 16495 | 773.2031
66 | 16489 | 772.9219
66 | 16486 | 772.7812
66 | 16486 | 772.7812
66 | 16492 | 773.0625
66 | 16494 | 773.1562
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16493 | 773.1094
66 | 16488 | 772.875
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16492 | 773.0625
66 | 16497 | 773.2969
66 | 16499 | 773.3906
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16489 | 772.9219
66 | 16485 | 772.7344
66 | 16490 | 772.9688
66 | 16494 | 773.1562
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16490 | 772.9688
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16490 | 772.9688
66 | 16496 | 773.25
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16497 | 773.2969
66 | 16493 | 773.1094
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16485 | 772.7344
66 | 16488 | 772.875
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16492 | 773.0625
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16489 | 772.9219
66 | 16495 | 773.2031
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16493 | 773.1094
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16487 | 772.8281
66 | 16494 | 773.1562
66 | 16496 | 773.25
66 | 16494 | 773.1562
66 | 16490 | 772.9688
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16490 | 772.9688
66 | 16495 | 773.2031
66 | 16497 | 773.2969
66 | 16495 | 773.2031
66 | 16489 | 772.9219
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16487 | 772.8281
66 | 16491 | 773.0156
66 | 16492 | 773.0625
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16484 | 772.6875
66 | 16485 | 772.7344
66 | 16491 | 773.0156
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16491 | 773.0156
66 | 16487 | 772.8281
66 | 16485 | 772.7344
66 | 16490 | 772.9688
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16488 | 772.875
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16491 | 773.0156
66 | 16494 | 773.1562
66 | 16495 | 773.2031
66 | 16493 | 773.1094
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16491 | 773.0156
66 | 16496 | 773.25
66 | 16495 | 773.2031
66 | 16492 | 773.0625
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16490 | 772.9688
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16492 | 773.0625
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16485 | 772.7344
66 | 16491 | 773.0156
66 | 16496 | 773.25
66 | 16495 | 773.2031
66 | 16495 | 773.2031
66 | 16491 | 773.0156
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16486 | 772.7812
66 | 16491 | 773.0156
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16485 | 772.7344
66 | 16486 | 772.7812
66 | 16493 | 773.1094
66 | 16497 | 773.2969
66 | 16496 | 773.25
66 | 16493 | 773.1094
66 | 16487 | 772.8281
66 | 16485 | 772.7344
66 | 16490 | 772.9688
66 | 16494 | 773.1562
66 | 16493 | 773.1094
66 | 16491 | 773.0156
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16485 | 772.7344
66 | 16492 | 773.0625
66 | 16494 | 773.1562
66 | 16494 | 773.1562
66 | 16489 | 772.9219
66 | 16486 | 772.7812
66 | 16487 | 772.8281
66 | 16494 | 773.1562
66 | 16495 | 773.2031
66 | 16494 | 773.1562
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16487 | 772.8281
66 | 16492 | 773.0625
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16486 | 772.7812
66 | 16492 | 773.0625
66 | 16493 | 773.1094
66 | 16494 | 773.1562
66 | 16490 | 772.9688
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16484 | 772.6875
66 | 16490 | 772.9688
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16481 | 772.5469
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16488 | 772.875
66 | 16492 | 773.0625
66 | 16493 | 773.1094
66 | 16491 | 773.0156
66 | 16487 | 772.8281
66 | 16482 | 772.5938
66 | 16482 | 772.5938
66 | 16488 | 772.875
66 | 16490 | 772.9688
66 | 16492 | 773.0625
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16482 | 772.5938
66 | 16487 | 772.8281
66 | 16492 | 773.0625
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16485 | 772.7344
66 | 16483 | 772.6406
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16491 | 773.0156
66 | 16487 | 772.8281
66 | 16481 | 772.5469
66 | 16483 | 772.6406
66 | 16491 | 773.0156
66 | 16494 | 773.1562
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16487 | 772.8281
66 | 16492 | 773.0625
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16485 | 772.7344
66 | 16483 | 772.6406
66 | 16489 | 772.9219
66 | 16491 | 773.0156
66 | 16490 | 772.9688
66 | 16486 | 772.7812
66 | 16481 | 772.5469
66 | 16485 | 772.7344
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16483 | 772.6406
66 | 16488 | 772.875
66 | 16493 | 773.1094
66 | 16493 | 773.1094
66 | 16490 | 772.9688
66 | 16483 | 772.6406
66 | 16483 | 772.6406
66 | 16491 | 773.0156
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16482 | 772.5938
66 | 16481 | 772.5469
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16483 | 772.6406
66 | 16489 | 772.9219
66 | 16493 | 773.1094
66 | 16492 | 773.0625
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16487 | 772.8281
66 | 16492 | 773.0625
66 | 16490 | 772.9688
66 | 16484 | 772.6875
66 | 16481 | 772.5469
66 | 16486 | 772.7812
66 | 16490 | 772.9688
66 | 16490 | 772.9688
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16484 | 772.6875
66 | 16491 | 773.0156
66 | 16493 | 773.1094
66 | 16490 | 772.9688
66 | 16483 | 772.6406
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16490 | 772.9688
66 | 16486 | 772.7812
66 | 16481 | 772.5469
66 | 16486 | 772.7812
66 | 16490 | 772.9688
66 | 16490 | 772.9688
66 | 16484 | 772.6875
66 | 16480 | 772.5
66 | 16484 | 772.6875
66 | 16492 | 773.0625
66 | 16492 | 773.0625
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16485 | 772.7344
66 | 16491 | 773.0156
66 | 16491 | 773.0156
66 | 16489 | 772.9219
66 | 16483 | 772.6406
66 | 16481 | 772.5469
66 | 16488 | 772.875
66 | 16490 | 772.9688
66 | 16488 | 772.875
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16487 | 772.8281
66 | 16490 | 772.9688
66 | 16491 | 773.0156
66 | 16486 | 772.7812
66 | 16480 | 772.5
66 | 16484 | 772.6875
66 | 16492 | 773.0625
66 | 16491 | 773.0156
66 | 16484 | 772.6875
66 | 16480 | 772.5
66 | 16483 | 772.6406
66 | 16488 | 772.875
66 | 16488 | 772.875
66 | 16484 | 772.6875
66 | 16480 | 772.5
66 | 16485 | 772.7344
66 | 16491 | 773.0156
66 | 16487 | 772.8281
66 | 16482 | 772.5938
66 | 16480 | 772.5
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16481 | 772.5469
66 | 16479 | 772.4531
66 | 16488 | 772.875
66 | 16489 | 772.9219
66 | 16485 | 772.7344
66 | 16479 | 772.4531
66 | 16481 | 772.5469
66 | 16489 | 772.9219
66 | 16490 | 772.9688
66 | 16484 | 772.6875
66 | 16480 | 772.5
66 | 16484 | 772.6875
66 | 16490 | 772.9688
66 | 16489 | 772.9219
66 | 16485 | 772.7344
66 | 16480 | 772.5
66 | 16488 | 772.875
66 | 16489 | 772.9219
66 | 16485 | 772.7344
66 | 16480 | 772.5
66 | 16484 | 772.6875
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16480 | 772.5
66 | 16478 | 772.4062
66 | 16488 | 772.875
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16480 | 772.5
66 | 16486 | 772.7812
66 | 16489 | 772.9219
66 | 16488 | 772.875
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16481 | 772.5469
66 | 16476 | 772.3125
66 | 16487 | 772.8281
66 | 16488 | 772.875
66 | 16484 | 772.6875
66 | 16479 | 772.4531
66 | 16488 | 772.875
66 | 16490 | 772.9688
66 | 16487 | 772.8281
66 | 16480 | 772.5
66 | 16488 | 772.875
66 | 16491 | 773.0156
66 | 16488 | 772.875
66 | 16480 | 772.5
66 | 16478 | 772.4062
66 | 16487 | 772.8281
66 | 16486 | 772.7812
66 | 16481 | 772.5469
66 | 16477 | 772.3594
66 | 16482 | 772.5938
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16479 | 772.4531
66 | 16488 | 772.875
66 | 16490 | 772.9688
66 | 16485 | 772.7344
66 | 16480 | 772.5
66 | 16486 | 772.7812
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16486 | 772.7812
66 | 16488 | 772.875
66 | 16480 | 772.5
66 | 16478 | 772.4062
66 | 16487 | 772.8281
66 | 16489 | 772.9219
66 | 16486 | 772.7812
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16488 | 772.875
66 | 16490 | 772.9688
66 | 16488 | 772.875
66 | 16480 | 772.5
66 | 16478 | 772.4062
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16482 | 772.5938
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16487 | 772.8281
66 | 16486 | 772.7812
66 | 16482 | 772.5938
66 | 16477 | 772.3594
66 | 16486 | 772.7812
66 | 16490 | 772.9688
66 | 16486 | 772.7812
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16486 | 772.7812
66 | 16485 | 772.7344
66 | 16478 | 772.4062
66 | 16480 | 772.5
66 | 16487 | 772.8281
66 | 16488 | 772.875
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16486 | 772.7812
66 | 16491 | 773.0156
66 | 16486 | 772.7812
66 | 16480 | 772.5
66 | 16483 | 772.6406
66 | 16489 | 772.9219
66 | 16485 | 772.7344
66 | 16481 | 772.5469
66 | 16476 | 772.3125
66 | 16485 | 772.7344
66 | 16486 | 772.7812
66 | 16481 | 772.5469
66 | 16478 | 772.4062
66 | 16485 | 772.7344
66 | 16489 | 772.9219
66 | 16487 | 772.8281
66 | 16480 | 772.5
66 | 16481 | 772.5469
66 | 16488 | 772.875
66 | 16487 | 772.8281
66 | 16481 | 772.5469
66 | 16477 | 772.3594
66 | 16484 | 772.6875
66 | 16487 | 772.8281
66 | 16484 | 772.6875
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16477 | 772.3594
66 | 16479 | 772.4531
66 | 16487 | 772.8281
66 | 16487 | 772.8281
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16487 | 772.8281
66 | 16488 | 772.875
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16484 | 772.6875
66 | 16487 | 772.8281
66 | 16481 | 772.5469
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16486 | 772.7812
66 | 16482 | 772.5938
66 | 16477 | 772.3594
66 | 16486 | 772.7812
66 | 16488 | 772.875
66 | 16484 | 772.6875
66 | 16480 | 772.5
66 | 16484 | 772.6875
66 | 16489 | 772.9219
66 | 16485 | 772.7344
66 | 16477 | 772.3594
66 | 16482 | 772.5938
66 | 16485 | 772.7344
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16480 | 772.5
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16483 | 772.6406
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16478 | 772.4062
66 | 16485 | 772.7344
66 | 16488 | 772.875
66 | 16486 | 772.7812
66 | 16477 | 772.3594
66 | 16484 | 772.6875
66 | 16486 | 772.7812
66 | 16483 | 772.6406
66 | 16475 | 772.2656
66 | 16481 | 772.5469
66 | 16485 | 772.7344
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16487 | 772.8281
66 | 16485 | 772.7344
66 | 16478 | 772.4062
66 | 16477 | 772.3594
66 | 16486 | 772.7812
66 | 16486 | 772.7812
66 | 16480 | 772.5
66 | 16476 | 772.3125
66 | 16484 | 772.6875
66 | 16485 | 772.7344
66 | 16479 | 772.4531
66 | 16474 | 772.2188
66 | 16484 | 772.6875
66 | 16485 | 772.7344
66 | 16478 | 772.4062
66 | 16475 | 772.2656
66 | 16483 | 772.6406
66 | 16485 | 772.7344
66 | 16480 | 772.5
66 | 16475 | 772.2656
66 | 16486 | 772.7812
66 | 16484 | 772.6875
66 | 16477 | 772.3594
66 | 16482 | 772.5938
66 | 16486 | 772.7812
66 | 16482 | 772.5938
66 | 16475 | 772.2656
66 | 16483 | 772.6406
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16474 | 772.2188
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16477 | 772.3594
66 | 16482 | 772.5938
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16477 | 772.3594
66 | 16486 | 772.7812
66 | 16484 | 772.6875
66 | 16478 | 772.4062
66 | 16479 | 772.4531
66 | 16485 | 772.7344
66 | 16482 | 772.5938
66 | 16474 | 772.2188
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16477 | 772.3594
66 | 16474 | 772.2188
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16477 | 772.3594
66 | 16482 | 772.5938
66 | 16487 | 772.8281
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16484 | 772.6875
66 | 16484 | 772.6875
66 | 16478 | 772.4062
66 | 16474 | 772.2188
66 | 16483 | 772.6406
66 | 16482 | 772.5938
66 | 16475 | 772.2656
66 | 16473 | 772.1719
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16475 | 772.2656
66 | 16477 | 772.3594
66 | 16486 | 772.7812
66 | 16484 | 772.6875
66 | 16476 | 772.3125
66 | 16479 | 772.4531
66 | 16485 | 772.7344
66 | 16484 | 772.6875
66 | 16475 | 772.2656
66 | 16483 | 772.6406
66 | 16483 | 772.6406
66 | 16477 | 772.3594
66 | 16473 | 772.1719
66 | 16482 | 772.5938
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16479 | 772.4531
66 | 16485 | 772.7344
66 | 16486 | 772.7812
66 | 16477 | 772.3594
66 | 16483 | 772.6406
66 | 16486 | 772.7812
66 | 16478 | 772.4062
66 | 16476 | 772.3125
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16474 | 772.2188
66 | 16477 | 772.3594
66 | 16484 | 772.6875
66 | 16482 | 772.5938
66 | 16472 | 772.125
66 | 16480 | 772.5
66 | 16483 | 772.6406
66 | 16474 | 772.2188
66 | 16479 | 772.4531
66 | 16485 | 772.7344
66 | 16482 | 772.5938
66 | 16475 | 772.2656
66 | 16483 | 772.6406
66 | 16484 | 772.6875
66 | 16475 | 772.2656
66 | 16481 | 772.5469
66 | 16482 | 772.5938
66 | 16478 | 772.4062
66 | 16473 | 772.1719
66 | 16482 | 772.5938
66 | 16479 | 772.4531
66 | 16472 | 772.125
66 | 16482 | 772.5938
66 | 16482 | 772.5938
66 | 16474 | 772.2188
66 | 16479 | 772.4531
66 | 16485 | 772.7344
66 | 16481 | 772.5469
66 | 16475 | 772.2656
66 | 16484 | 772.6875
66 | 16483 | 772.6406
66 | 16477 | 772.3594
66 | 16475 | 772.2656
66 | 16484 | 772.6875
66 | 16481 | 772.5469
66 | 16473 | 772.1719
66 | 16480 | 772.5
66 | 16481 | 772.5469
66 | 16473 | 772.1719
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16474 | 772.2188
66 | 16478 | 772.4062
66 | 16485 | 772.7344
66 | 16482 | 772.5938
66 | 16475 | 772.2656
66 | 16483 | 772.6406
66 | 16483 | 772.6406
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16482 | 772.5938
66 | 16474 | 772.2188
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16473 | 772.1719
66 | 16481 | 772.5469
66 | 16482 | 772.5938
66 | 16474 | 772.2188
66 | 16482 | 772.5938
66 | 16484 | 772.6875
66 | 16476 | 772.3125
66 | 16481 | 772.5469
66 | 16482 | 772.5938
66 | 16473 | 772.1719
66 | 16476 | 772.3125
66 | 16482 | 772.5938
66 | 16477 | 772.3594
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16472 | 772.125
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16474 | 772.2188
66 | 16483 | 772.6406
66 | 16482 | 772.5938
66 | 16474 | 772.2188
66 | 16482 | 772.5938
66 | 16481 | 772.5469
66 | 16472 | 772.125
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16472 | 772.125
66 | 16476 | 772.3125
66 | 16480 | 772.5
66 | 16473 | 772.1719
66 | 16474 | 772.2188
66 | 16482 | 772.5938
66 | 16480 | 772.5
66 | 16474 | 772.2188
66 | 16483 | 772.6406
66 | 16482 | 772.5938
66 | 16473 | 772.1719
66 | 16483 | 772.6406
66 | 16481 | 772.5469
66 | 16473 | 772.1719
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16479 | 772.4531
66 | 16470 | 772.0312
66 | 16481 | 772.5469
66 | 16479 | 772.4531
66 | 16472 | 772.125
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16474 | 772.2188
66 | 16484 | 772.6875
66 | 16482 | 772.5938
66 | 16474 | 772.2188
66 | 16483 | 772.6406
66 | 16480 | 772.5
66 | 16472 | 772.125
66 | 16482 | 772.5938
66 | 16479 | 772.4531
66 | 16470 | 772.0312
66 | 16481 | 772.5469
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16472 | 772.125
66 | 16483 | 772.6406
66 | 16480 | 772.5
66 | 16473 | 772.1719
66 | 16483 | 772.6406
66 | 16481 | 772.5469
66 | 16473 | 772.1719
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16472 | 772.125
66 | 16481 | 772.5469
66 | 16477 | 772.3594
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16476 | 772.3125
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16479 | 772.4531
66 | 16473 | 772.1719
66 | 16482 | 772.5938
66 | 16481 | 772.5469
66 | 16473 | 772.1719
66 | 16482 | 772.5938
66 | 16481 | 772.5469
66 | 16471 | 772.0781
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16473 | 772.1719
66 | 16480 | 772.5
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16472 | 772.125
66 | 16483 | 772.6406
66 | 16481 | 772.5469
66 | 16472 | 772.125
66 | 16481 | 772.5469
66 | 16481 | 772.5469
66 | 16472 | 772.125
66 | 16481 | 772.5469
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16478 | 772.4062
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16471 | 772.0781
66 | 16482 | 772.5938
66 | 16479 | 772.4531
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16477 | 772.3594
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16475 | 772.2656
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16481 | 772.5469
66 | 16480 | 772.5
66 | 16471 | 772.0781
66 | 16482 | 772.5938
66 | 16479 | 772.4531
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16478 | 772.4062
66 | 16468 | 771.9375
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16468 | 771.9375
66 | 16479 | 772.4531
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16476 | 772.3125
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16479 | 772.4531
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16467 | 771.8906
66 | 16479 | 772.4531
66 | 16475 | 772.2656
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16479 | 772.4531
66 | 16471 | 772.0781
66 | 16481 | 772.5469
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16475 | 772.2656
66 | 16466 | 771.8438
66 | 16478 | 772.4062
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16477 | 772.3594
66 | 16476 | 772.3125
66 | 16470 | 772.0312
66 | 16479 | 772.4531
66 | 16479 | 772.4531
66 | 16470 | 772.0312
66 | 16480 | 772.5
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16475 | 772.2656
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16474 | 772.2188
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16475 | 772.2656
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16474 | 772.2188
66 | 16469 | 771.9844
66 | 16480 | 772.5
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16474 | 772.2188
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16474 | 772.2188
66 | 16467 | 771.8906
66 | 16478 | 772.4062
66 | 16473 | 772.1719
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16475 | 772.2656
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16479 | 772.4531
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16474 | 772.2188
66 | 16467 | 771.8906
66 | 16476 | 772.3125
66 | 16473 | 772.1719
66 | 16466 | 771.8438
66 | 16476 | 772.3125
66 | 16474 | 772.2188
66 | 16466 | 771.8438
66 | 16477 | 772.3594
66 | 16476 | 772.3125
66 | 16468 | 771.9375
66 | 16479 | 772.4531
66 | 16474 | 772.2188
66 | 16471 | 772.0781
66 | 16480 | 772.5
66 | 16474 | 772.2188
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16474 | 772.2188
66 | 16467 | 771.8906
66 | 16476 | 772.3125
66 | 16472 | 772.125
66 | 16467 | 771.8906
66 | 16477 | 772.3594
66 | 16473 | 772.1719
66 | 16467 | 771.8906
66 | 16479 | 772.4531
66 | 16472 | 772.125
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16475 | 772.2656
66 | 16467 | 771.8906
66 | 16477 | 772.3594
66 | 16471 | 772.0781
66 | 16468 | 771.9375
66 | 16476 | 772.3125
66 | 16470 | 772.0312
66 | 16468 | 771.9375
66 | 16475 | 772.2656
66 | 16471 | 772.0781
66 | 16467 | 771.8906
66 | 16478 | 772.4062
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16479 | 772.4531
66 | 16473 | 772.1719
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16467 | 771.8906
66 | 16475 | 772.2656
66 | 16469 | 771.9844
66 | 16469 | 771.9844
66 | 16476 | 772.3125
66 | 16468 | 771.9375
66 | 16470 | 772.0312
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16472 | 772.125
66 | 16478 | 772.4062
66 | 16471 | 772.0781
66 | 16469 | 771.9844
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16470 | 772.0312
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16467 | 771.8906
66 | 16475 | 772.2656
66 | 16469 | 771.9844
66 | 16467 | 771.8906
66 | 16476 | 772.3125
66 | 16470 | 772.0312
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16471 | 772.0781
66 | 16468 | 771.9375
66 | 16477 | 772.3594
66 | 16471 | 772.0781
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16471 | 772.0781
66 | 16468 | 771.9375
66 | 16477 | 772.3594
66 | 16469 | 771.9844
66 | 16468 | 771.9375
66 | 16475 | 772.2656
66 | 16468 | 771.9375
66 | 16469 | 771.9844
66 | 16476 | 772.3125
66 | 16466 | 771.8438
66 | 16477 | 772.3594
66 | 16474 | 772.2188
66 | 16468 | 771.9375
66 | 16478 | 772.4062
66 | 16470 | 772.0312
66 | 16470 | 772.0312
66 | 16477 | 772.3594
66 | 16470 | 772.0312
66 | 16469 | 771.9844
66 | 16476 | 772.3125
66 | 16468 | 771.9375
66 | 16473 | 772.1719
66 | 16474 | 772.2188
66 | 16465 | 771.7969
66 | 16475 | 772.2656
66 | 16472 | 772.125
66 | 16465 | 771.7969
66 | 16477 | 772.3594
66 | 16474 | 772.2188
66 | 16468 | 771.9375
66 | 16476 | 772.3125
66 | 16473 | 772.1719
66 | 16467 | 771.8906
66 | 16476 | 772.3125
66 | 16470 | 772.0312
66 | 16466 | 771.8438
66 | 16475 | 772.2656
66 | 16467 | 771.8906
66 | 16465 | 771.7969
66 | 16475 | 772.2656
66 | 16467 | 771.8906
66 | 16473 | 772.1719
66 | 16474 | 772.2188
66 | 16465 | 771.7969
66 | 16476 | 772.3125
66 | 16470 | 772.0312
66 | 16469 | 771.9844
66 | 16478 | 772.4062
66 | 16469 | 771.9844
66 | 16475 | 772.2656
66 | 16475 | 772.2656
66 | 16467 | 771.8906
66 | 16477 | 772.3594
66 | 16467 | 771.8906
66 | 16473 | 772.1719
66 | 16471 | 772.0781
66 | 16465 | 771.7969
66 | 16474 | 772.2188
66 | 16467 | 771.8906
66 | 16474 | 772.2188
66 | 16473 | 772.1719
66 | 16467 | 771.8906
66 | 16477 | 772.3594
66 | 16468 | 771.9375
66 | 16472 | 772.125
66 | 16475 | 772.2656
66 | 16466 | 771.8438
66 | 16476 | 772.3125
66 | 16471 | 772.0781
66 | 16467 | 771.8906
66 | 16475 | 772.2656
66 | 16466 | 771.8438
66 | 16471 | 772.0781
66 | 16471 | 772.0781
66 | 16463 | 771.7031
66 | 16475 | 772.2656
66 | 16469 | 771.9844
66 | 16471 | 772.0781
66 | 16475 | 772.2656
66 | 16466 | 771.8438
66 | 16475 | 772.2656
66 | 16473 | 772.1719
66 | 16466 | 771.8438
66 | 16476 | 772.3125
66 | 16468 | 771.9375
66 | 16472 | 772.125
66 | 16473 | 772.1719
66 | 16462 | 771.6562
66 | 16474 | 772.2188
66 | 16467 | 771.8906
66 | 16465 | 771.7969
66 | 16473 | 772.1719
66 | 16464 | 771.75
66 | 16476 | 772.3125
66 | 16472 | 772.125
66 | 16467 | 771.8906
66 | 16476 | 772.3125
66 | 16467 | 771.8906
66 | 16474 | 772.2188
66 | 16473 | 772.1719
66 | 16464 | 771.75
66 | 16476 | 772.3125
66 | 16468 | 771.9375
66 | 16471 | 772.0781
66 | 16472 | 772.125
66 | 16463 | 771.7031
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16466 | 771.8438
66 | 16474 | 772.2188
66 | 16467 | 771.8906
66 | 16473 | 772.1719
66 | 16472 | 772.125
66 | 16464 | 771.75
66 | 16476 | 772.3125
66 | 16469 | 771.9844
66 | 16470 | 772.0312
66 | 16473 | 772.1719
66 | 16464 | 771.75
66 | 16475 | 772.2656
66 | 16468 | 771.9375
66 | 16467 | 771.8906
66 | 16474 | 772.2188
66 | 16463 | 771.7031
66 | 16474 | 772.2188
66 | 16468 | 771.9375
66 | 16465 | 771.7969
66 | 16474 | 772.2188
66 | 16466 | 771.8438
66 | 16472 | 772.125
66 | 16474 | 772.2188
66 | 16465 | 771.7969
66 | 16474 | 772.2188
66 | 16468 | 771.9375
66 | 16470 | 772.0312
66 | 16472 | 772.125
66 | 16463 | 771.7031
66 | 16472 | 772.125
66 | 16466 | 771.8438
66 | 16468 | 771.9375
66 | 16470 | 772.0312
66 | 16462 | 771.6562
66 | 16473 | 772.1719
66 | 16467 | 771.8906
66 | 16469 | 771.9844
66 | 16472 | 772.125
66 | 16465 | 771.7969
66 | 16475 | 772.2656
66 | 16470 | 772.0312
66 | 16468 | 771.9375
66 | 16475 | 772.2656
66 | 16464 | 771.75
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16473 | 772.1719
66 | 16462 | 771.6562
66 | 16471 | 772.0781
66 | 16469 | 771.9844
66 | 16463 | 771.7031
66 | 16472 | 772.125
66 | 16464 | 771.75
66 | 16474 | 772.2188
66 | 16469 | 771.9844
66 | 16466 | 771.8438
66 | 16475 | 772.2656
66 | 16465 | 771.7969
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16465 | 771.7969
66 | 16472 | 772.125
66 | 16461 | 771.6094
66 | 16471 | 772.0781
66 | 16467 | 771.8906
66 | 16463 | 771.7031
66 | 16473 | 772.1719
66 | 16465 | 771.7969
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16467 | 771.8906
66 | 16475 | 772.2656
66 | 16465 | 771.7969
66 | 16475 | 772.2656
66 | 16468 | 771.9375
66 | 16466 | 771.8438
66 | 16473 | 772.1719
66 | 16464 | 771.75
66 | 16470 | 772.0312
66 | 16469 | 771.9844
66 | 16462 | 771.6562
66 | 16471 | 772.0781
66 | 16465 | 771.7969
66 | 16469 | 771.9844
66 | 16469 | 771.9844
66 | 16462 | 771.6562
66 | 16473 | 772.1719
66 | 16469 | 771.9844
66 | 16468 | 771.9375
66 | 16474 | 772.2188
66 | 16464 | 771.75
66 | 16473 | 772.1719
66 | 16470 | 772.0312
66 | 16462 | 771.6562
66 | 16471 | 772.0781
66 | 16465 | 771.7969
66 | 16469 | 771.9844
66 | 16468 | 771.9375
66 | 16460 | 771.5625
66 | 16471 | 772.0781
66 | 16464 | 771.75
66 | 16470 | 772.0312
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16473 | 772.1719
66 | 16463 | 771.7031
66 | 16473 | 772.1719
66 | 16468 | 771.9375
66 | 16468 | 771.9375
66 | 16472 | 772.125
66 | 16462 | 771.6562
66 | 16471 | 772.0781
66 | 16463 | 771.7031
66 | 16469 | 771.9844
66 | 16467 | 771.8906
66 | 16460 | 771.5625
66 | 16470 | 772.0312
66 | 16462 | 771.6562
66 | 16470 | 772.0312
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16472 | 772.125
66 | 16462 | 771.6562
66 | 16472 | 772.125
66 | 16463 | 771.7031
66 | 16472 | 772.125
66 | 16464 | 771.75
66 | 16467 | 771.8906
66 | 16467 | 771.8906
66 | 16463 | 771.7031
66 | 16470 | 772.0312
66 | 16461 | 771.6094
66 | 16472 | 772.125
66 | 16464 | 771.75
66 | 16470 | 772.0312
66 | 16467 | 771.8906
66 | 16467 | 771.8906
66 | 16470 | 772.0312
66 | 16465 | 771.7969
66 | 16473 | 772.1719
66 | 16464 | 771.75
66 | 16473 | 772.1719
66 | 16464 | 771.75
66 | 16471 | 772.0781
66 | 16464 | 771.75
66 | 16467 | 771.8906
66 | 16467 | 771.8906
66 | 16462 | 771.6562
66 | 16470 | 772.0312
66 | 16461 | 771.6094
66 | 16471 | 772.0781
66 | 16465 | 771.7969
66 | 16470 | 772.0312
66 | 16469 | 771.9844
66 | 16464 | 771.75
66 | 16473 | 772.1719
66 | 16463 | 771.7031
66 | 16472 | 772.125
66 | 16465 | 771.7969
66 | 16465 | 771.7969
66 | 16470 | 772.0312
66 | 16460 | 771.5625
66 | 16471 | 772.0781
66 | 16462 | 771.6562
66 | 16469 | 771.9844
66 | 16466 | 771.8438
66 | 16462 | 771.6562
66 | 16469 | 771.9844
66 | 16460 | 771.5625
66 | 16470 | 772.0312
66 | 16464 | 771.75
66 | 16468 | 771.9375
66 | 16471 | 772.0781
66 | 16462 | 771.6562
66 | 16472 | 772.125
66 | 16463 | 771.7031
66 | 16471 | 772.0781
66 | 16466 | 771.8438
66 | 16464 | 771.75
66 | 16469 | 771.9844
66 | 16460 | 771.5625
66 | 16469 | 771.9844
66 | 16461 | 771.6094
66 | 16468 | 771.9375
66 | 16467 | 771.8906
66 | 16461 | 771.6094
66 | 16471 | 772.0781
66 | 16462 | 771.6562
66 | 16471 | 772.0781
66 | 16467 | 771.8906
66 | 16465 | 771.7969
66 | 16470 | 772.0312
66 | 16461 | 771.6094
66 | 16472 | 772.125
66 | 16464 | 771.75
66 | 16464 | 771.75
66 | 16468 | 771.9375
66 | 16460 | 771.5625
66 | 16469 | 771.9844
66 | 16462 | 771.6562
66 | 16464 | 771.75
66 | 16468 | 771.9375
66 | 16461 | 771.6094
66 | 16471 | 772.0781
66 | 16461 | 771.6094
66 | 16468 | 771.9375
66 | 16467 | 771.8906
66 | 16461 | 771.6094
66 | 16471 | 772.0781
66 | 16463 | 771.7031
66 | 16466 | 771.8438
66 | 16466 | 771.8438
66 | 16460 | 771.5625
66 | 16470 | 772.0312
66 | 16460 | 771.5625
66 | 16467 | 771.8906
66 | 16465 | 771.7969
66 | 16462 | 771.6562
66 | 16467 | 771.8906
66 | 16460 | 771.5625
66 | 16469 | 771.9844
66 | 16461 | 771.6094
66 | 16470 | 772.0312
66 | 16463 | 771.7031
66 | 16467 | 771.8906
66 | 16468 | 771.9375
66 | 16462 | 771.6562
66 | 16469 | 771.9844
66 | 16459 | 771.5156
66 | 16468 | 771.9375
66 | 16459 | 771.5156
66 | 16468 | 771.9375
66 | 16462 | 771.6562
66 | 16463 | 771.7031
66 | 16465 | 771.7969
66 | 16459 | 771.5156
66 | 16469 | 771.9844
66 | 16459 | 771.5156
66 | 16470 | 772.0312
66 | 16465 | 771.7969
66 | 16468 | 771.9375
66 | 16469 | 771.9844
66 | 16461 | 771.6094
66 | 16469 | 771.9844
66 | 16459 | 771.5156
66 | 16469 | 771.9844
66 | 16460 | 771.5625
66 | 16467 | 771.8906
66 | 16465 | 771.7969
66 | 16460 | 771.5625
66 | 16467 | 771.8906
66 | 16458 | 771.4688
66 | 16469 | 771.9844
66 | 16460 | 771.5625
66 | 16469 | 771.9844
66 | 16462 | 771.6562
66 | 16466 | 771.8438
66 | 16467 | 771.8906
66 | 16462 | 771.6562
66 | 16471 | 772.0781
66 | 16459 | 771.5156
66 | 16469 | 771.9844
66 | 16463 | 771.7031
66 | 16466 | 771.8438
66 | 16466 | 771.8438
66 | 16459 | 771.5156
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16468 | 771.9375
66 | 16462 | 771.6562
66 | 16466 | 771.8438
66 | 16467 | 771.8906
66 | 16462 | 771.6562
66 | 16470 | 772.0312
66 | 16459 | 771.5156
66 | 16470 | 772.0312
66 | 16461 | 771.6094
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16462 | 771.6562
66 | 16466 | 771.8438
66 | 16458 | 771.4688
66 | 16469 | 771.9844
66 | 16457 | 771.4219
66 | 16469 | 771.9844
66 | 16459 | 771.5156
66 | 16467 | 771.8906
66 | 16463 | 771.7031
66 | 16465 | 771.7969
66 | 16466 | 771.8438
66 | 16462 | 771.6562
66 | 16468 | 771.9375
66 | 16459 | 771.5156
66 | 16470 | 772.0312
66 | 16460 | 771.5625
66 | 16467 | 771.8906
66 | 16464 | 771.75
66 | 16461 | 771.6094
66 | 16467 | 771.8906
66 | 16457 | 771.4219
66 | 16467 | 771.8906
66 | 16457 | 771.4219
66 | 16467 | 771.8906
66 | 16460 | 771.5625
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16465 | 771.7969
66 | 16467 | 771.8906
66 | 16459 | 771.5156
66 | 16470 | 772.0312
66 | 16460 | 771.5625
66 | 16469 | 771.9844
66 | 16458 | 771.4688
66 | 16469 | 771.9844
66 | 16461 | 771.6094
66 | 16466 | 771.8438
66 | 16464 | 771.75
66 | 16458 | 771.4688
66 | 16465 | 771.7969
66 | 16457 | 771.4219
66 | 16468 | 771.9375
66 | 16458 | 771.4688
66 | 16469 | 771.9844
66 | 16461 | 771.6094
66 | 16468 | 771.9375
66 | 16464 | 771.75
66 | 16465 | 771.7969
66 | 16468 | 771.9375
66 | 16460 | 771.5625
66 | 16468 | 771.9375
66 | 16458 | 771.4688
66 | 16469 | 771.9844
66 | 16459 | 771.5156
66 | 16466 | 771.8438
66 | 16462 | 771.6562
66 | 16461 | 771.6094
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16467 | 771.8906
66 | 16459 | 771.5156
66 | 16470 | 772.0312
66 | 16464 | 771.75
66 | 16465 | 771.7969
66 | 16468 | 771.9375
66 | 16459 | 771.5156
66 | 16469 | 771.9844
66 | 16458 | 771.4688
66 | 16468 | 771.9375
66 | 16459 | 771.5156
66 | 16467 | 771.8906
66 | 16462 | 771.6562
66 | 16460 | 771.5625
66 | 16466 | 771.8438
66 | 16456 | 771.375
66 | 16467 | 771.8906
66 | 16458 | 771.4688
66 | 16467 | 771.8906
66 | 16465 | 771.7969
66 | 16463 | 771.7031
66 | 16470 | 772.0312
66 | 16459 | 771.5156
66 | 16469 | 771.9844
66 | 16462 | 771.6562
66 | 16466 | 771.8438
66 | 16463 | 771.7031
66 | 16460 | 771.5625
66 | 16467 | 771.8906
66 | 16457 | 771.4219
66 | 16465 | 771.7969
66 | 16456 | 771.375
66 | 16466 | 771.8438
66 | 16461 | 771.6094
66 | 16461 | 771.6094
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16468 | 771.9375
66 | 16458 | 771.4688
66 | 16467 | 771.8906
66 | 16461 | 771.6094
66 | 16464 | 771.75
66 | 16465 | 771.7969
66 | 16458 | 771.4688
66 | 16467 | 771.8906
66 | 16456 | 771.375
66 | 16466 | 771.8438
66 | 16460 | 771.5625
66 | 16459 | 771.5156
66 | 16464 | 771.75
66 | 16456 | 771.375
66 | 16466 | 771.8438
66 | 16459 | 771.5156
66 | 16465 | 771.7969
66 | 16464 | 771.75
66 | 16461 | 771.6094
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16468 | 771.9375
66 | 16456 | 771.375
66 | 16466 | 771.8438
66 | 16461 | 771.6094
66 | 16464 | 771.75
66 | 16462 | 771.6562
66 | 16455 | 771.3281
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16466 | 771.8438
66 | 16459 | 771.5156
66 | 16466 | 771.8438
66 | 16461 | 771.6094
66 | 16463 | 771.7031
66 | 16464 | 771.75
66 | 16460 | 771.5625
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16465 | 771.7969
66 | 16457 | 771.4219
66 | 16463 | 771.7031
66 | 16461 | 771.6094
66 | 16459 | 771.5156
66 | 16464 | 771.75
66 | 16456 | 771.375
66 | 16466 | 771.8438
66 | 16458 | 771.4688
66 | 16465 | 771.7969
66 | 16463 | 771.7031
66 | 16460 | 771.5625
66 | 16465 | 771.7969
66 | 16457 | 771.4219
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16468 | 771.9375
66 | 16458 | 771.4688
66 | 16466 | 771.8438
66 | 16456 | 771.375
66 | 16464 | 771.75
66 | 16459 | 771.5156
66 | 16463 | 771.7031
66 | 16463 | 771.7031
66 | 16459 | 771.5156
66 | 16465 | 771.7969
66 | 16458 | 771.4688
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16468 | 771.9375
66 | 16456 | 771.375
66 | 16467 | 771.8906
66 | 16457 | 771.4219
66 | 16465 | 771.7969
66 | 16459 | 771.5156
66 | 16464 | 771.75
66 | 16460 | 771.5625
66 | 16461 | 771.6094
66 | 16462 | 771.6562
66 | 16457 | 771.4219
66 | 16464 | 771.75
66 | 16455 | 771.3281
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16467 | 771.8906
66 | 16457 | 771.4219
66 | 16468 | 771.9375
66 | 16457 | 771.4219
66 | 16466 | 771.8438
66 | 16457 | 771.4219
66 | 16465 | 771.7969
66 | 16457 | 771.4219
66 | 16463 | 771.7031
66 | 16459 | 771.5156
66 | 16460 | 771.5625
66 | 16460 | 771.5625
66 | 16459 | 771.5156
66 | 16463 | 771.7031
66 | 16458 | 771.4688
66 | 16464 | 771.75
66 | 16457 | 771.4219
66 | 16467 | 771.8906
66 | 16456 | 771.375
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16465 | 771.7969
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16455 | 771.3281
66 | 16463 | 771.7031
66 | 16457 | 771.4219
66 | 16463 | 771.7031
66 | 16460 | 771.5625
66 | 16459 | 771.5156
66 | 16459 | 771.5156
66 | 16462 | 771.6562
66 | 16463 | 771.7031
66 | 16460 | 771.5625
66 | 16462 | 771.6562
66 | 16460 | 771.5625
66 | 16461 | 771.6094
66 | 16460 | 771.5625
66 | 16462 | 771.6562
66 | 16456 | 771.375
66 | 16462 | 771.6562
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16453 | 771.2344
66 | 16463 | 771.7031
66 | 16454 | 771.2812
66 | 16464 | 771.75
66 | 16454 | 771.2812
66 | 16465 | 771.7969
66 | 16455 | 771.3281
66 | 16465 | 771.7969
66 | 16455 | 771.3281
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16464 | 771.75
66 | 16455 | 771.3281
66 | 16462 | 771.6562
66 | 16454 | 771.2812
66 | 16462 | 771.6562
66 | 16454 | 771.2812
66 | 16461 | 771.6094
66 | 16457 | 771.4219
66 | 16463 | 771.7031
66 | 16458 | 771.4688
66 | 16463 | 771.7031
66 | 16460 | 771.5625
66 | 16460 | 771.5625
66 | 16461 | 771.6094
66 | 16459 | 771.5156
66 | 16462 | 771.6562
66 | 16458 | 771.4688
66 | 16463 | 771.7031
66 | 16456 | 771.375
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16463 | 771.7031
66 | 16452 | 771.1875
66 | 16465 | 771.7969
66 | 16453 | 771.2344
66 | 16464 | 771.75
66 | 16455 | 771.3281
66 | 16465 | 771.7969
66 | 16457 | 771.4219
66 | 16464 | 771.75
66 | 16459 | 771.5156
66 | 16462 | 771.6562
66 | 16461 | 771.6094
66 | 16459 | 771.5156
66 | 16460 | 771.5625
66 | 16456 | 771.375
66 | 16461 | 771.6094
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16453 | 771.2344
66 | 16463 | 771.7031
66 | 16453 | 771.2344
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16464 | 771.75
66 | 16455 | 771.3281
66 | 16466 | 771.8438
66 | 16455 | 771.3281
66 | 16464 | 771.75
66 | 16454 | 771.2812
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16461 | 771.6094
66 | 16451 | 771.1406
66 | 16462 | 771.6562
66 | 16453 | 771.2344
66 | 16463 | 771.7031
66 | 16454 | 771.2812
66 | 16465 | 771.7969
66 | 16454 | 771.2812
66 | 16464 | 771.75
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16453 | 771.2344
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16453 | 771.2344
66 | 16464 | 771.75
66 | 16453 | 771.2344
66 | 16463 | 771.7031
66 | 16455 | 771.3281
66 | 16464 | 771.75
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16453 | 771.2344
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16461 | 771.6094
66 | 16450 | 771.0938
66 | 16461 | 771.6094
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16452 | 771.1875
66 | 16464 | 771.75
66 | 16454 | 771.2812
66 | 16464 | 771.75
66 | 16455 | 771.3281
66 | 16464 | 771.75
66 | 16455 | 771.3281
66 | 16463 | 771.7031
66 | 16454 | 771.2812
66 | 16462 | 771.6562
66 | 16453 | 771.2344
66 | 16461 | 771.6094
66 | 16455 | 771.3281
66 | 16457 | 771.4219
66 | 16458 | 771.4688
66 | 16456 | 771.375
66 | 16458 | 771.4688
66 | 16456 | 771.375
66 | 16462 | 771.6562
66 | 16456 | 771.375
66 | 16460 | 771.5625
66 | 16456 | 771.375
66 | 16460 | 771.5625
66 | 16456 | 771.375
66 | 16461 | 771.6094
66 | 16454 | 771.2812
66 | 16460 | 771.5625
66 | 16450 | 771.0938
66 | 16459 | 771.5156
66 | 16451 | 771.1406
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16453 | 771.2344
66 | 16463 | 771.7031
66 | 16454 | 771.2812
66 | 16463 | 771.7031
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16450 | 771.0938
66 | 16461 | 771.6094
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16453 | 771.2344
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16463 | 771.7031
66 | 16451 | 771.1406
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16461 | 771.6094
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16455 | 771.3281
66 | 16459 | 771.5156
66 | 16458 | 771.4688
66 | 16457 | 771.4219
66 | 16459 | 771.5156
66 | 16459 | 771.5156
66 | 16460 | 771.5625
66 | 16459 | 771.5156
66 | 16458 | 771.4688
66 | 16459 | 771.5156
66 | 16456 | 771.375
66 | 16456 | 771.375
66 | 16457 | 771.4219
66 | 16454 | 771.2812
66 | 16459 | 771.5156
66 | 16457 | 771.4219
66 | 16457 | 771.4219
66 | 16458 | 771.4688
66 | 16455 | 771.3281
66 | 16460 | 771.5625
66 | 16454 | 771.2812
66 | 16462 | 771.6562
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16449 | 771.0469
66 | 16460 | 771.5625
66 | 16449 | 771.0469
66 | 16460 | 771.5625
66 | 16449 | 771.0469
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16460 | 771.5625
66 | 16450 | 771.0938
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16452 | 771.1875
66 | 16461 | 771.6094
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16459 | 771.5156
66 | 16449 | 771.0469
66 | 16459 | 771.5156
66 | 16448 | 771
66 | 16457 | 771.4219
66 | 16449 | 771.0469
66 | 16461 | 771.6094
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16451 | 771.1406
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16460 | 771.5625
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16449 | 771.0469
66 | 16459 | 771.5156
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16460 | 771.5625
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16451 | 771.1406
66 | 16461 | 771.6094
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16450 | 771.0938
66 | 16459 | 771.5156
66 | 16448 | 771
66 | 16459 | 771.5156
66 | 16448 | 771
66 | 16458 | 771.4688
66 | 16449 | 771.0469
66 | 16459 | 771.5156
66 | 16450 | 771.0938
66 | 16460 | 771.5625
66 | 16452 | 771.1875
66 | 16462 | 771.6562
66 | 16453 | 771.2344
66 | 16459 | 771.5156
66 | 16454 | 771.2812
66 | 16457 | 771.4219
66 | 16456 | 771.375
66 | 16452 | 771.1875
66 | 16456 | 771.375
66 | 16451 | 771.1406
66 | 16457 | 771.4219
66 | 16451 | 771.1406
66 | 16457 | 771.4219
66 | 16454 | 771.2812
66 | 16459 | 771.5156
66 | 16452 | 771.1875
66 | 16458 | 771.4688
66 | 16453 | 771.2344
66 | 16458 | 771.4688
66 | 16452 | 771.1875
66 | 16458 | 771.4688
66 | 16453 | 771.2344
66 | 16457 | 771.4219
66 | 16454 | 771.2812
66 | 16455 | 771.3281
66 | 16451 | 771.1406
66 | 16456 | 771.375
66 | 16447 | 770.9531
66 | 16458 | 771.4688
66 | 16449 | 771.0469
66 | 16459 | 771.5156
66 | 16449 | 771.0469
66 | 16461 | 771.6094
66 | 16453 | 771.2344
66 | 16459 | 771.5156
66 | 16452 | 771.1875
66 | 16458 | 771.4688
66 | 16452 | 771.1875
66 | 16458 | 771.4688
66 | 16452 | 771.1875
66 | 16455 | 771.3281
66 | 16451 | 771.1406
66 | 16456 | 771.375
66 | 16452 | 771.1875
66 | 16454 | 771.2812
66 | 16455 | 771.3281
66 | 16454 | 771.2812
66 | 16456 | 771.375
66 | 16452 | 771.1875
66 | 16459 | 771.5156
66 | 16454 | 771.2812
66 | 16457 | 771.4219
66 | 16450 | 771.0938
66 | 16457 | 771.4219
66 | 16448 | 771
66 | 16457 | 771.4219
66 | 16448 | 771
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16449 | 771.0469
66 | 16456 | 771.375
66 | 16451 | 771.1406
66 | 16457 | 771.4219
66 | 16452 | 771.1875
66 | 16457 | 771.4219
66 | 16452 | 771.1875
66 | 16458 | 771.4688
66 | 16452 | 771.1875
66 | 16456 | 771.375
66 | 16452 | 771.1875
66 | 16453 | 771.2344
66 | 16455 | 771.3281
66 | 16452 | 771.1875
66 | 16455 | 771.3281
66 | 16451 | 771.1406
66 | 16457 | 771.4219
66 | 16449 | 771.0469
66 | 16458 | 771.4688
66 | 16449 | 771.0469
66 | 16458 | 771.4688
66 | 16449 | 771.0469
66 | 16458 | 771.4688
66 | 16450 | 771.0938
66 | 16458 | 771.4688
66 | 16453 | 771.2344
66 | 16455 | 771.3281
66 | 16455 | 771.3281
66 | 16450 | 771.0938
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16447 | 770.9531
66 | 16458 | 771.4688
66 | 16449 | 771.0469
66 | 16458 | 771.4688
66 | 16451 | 771.1406
66 | 16457 | 771.4219
66 | 16454 | 771.2812
66 | 16456 | 771.375
66 | 16456 | 771.375
66 | 16452 | 771.1875
66 | 16456 | 771.375
66 | 16450 | 771.0938
66 | 16456 | 771.375
66 | 16447 | 770.9531
66 | 16457 | 771.4219
66 | 16445 | 770.8594
66 | 16457 | 771.4219
66 | 16447 | 770.9531
66 | 16457 | 771.4219
66 | 16448 | 771
66 | 16458 | 771.4688
66 | 16450 | 771.0938
66 | 16456 | 771.375
66 | 16453 | 771.2344
66 | 16453 | 771.2344
66 | 16455 | 771.3281
66 | 16451 | 771.1406
66 | 16456 | 771.375
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16445 | 770.8594
66 | 16457 | 771.4219
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16450 | 771.0938
66 | 16454 | 771.2812
66 | 16455 | 771.3281
66 | 16455 | 771.3281
66 | 16457 | 771.4219
66 | 16453 | 771.2344
66 | 16457 | 771.4219
66 | 16448 | 771
66 | 16457 | 771.4219
66 | 16446 | 770.9062
66 | 16455 | 771.3281
66 | 16446 | 770.9062
66 | 16456 | 771.375
66 | 16447 | 770.9531
66 | 16454 | 771.2812
66 | 16449 | 771.0469
66 | 16453 | 771.2344
66 | 16454 | 771.2812
66 | 16452 | 771.1875
66 | 16456 | 771.375
66 | 16451 | 771.1406
66 | 16457 | 771.4219
66 | 16451 | 771.1406
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16446 | 770.9062
66 | 16456 | 771.375
66 | 16445 | 770.8594
66 | 16456 | 771.375
66 | 16446 | 770.9062
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16452 | 771.1875
66 | 16455 | 771.3281
66 | 16454 | 771.2812
66 | 16455 | 771.3281
66 | 16455 | 771.3281
66 | 16453 | 771.2344
66 | 16453 | 771.2344
66 | 16452 | 771.1875
66 | 16452 | 771.1875
66 | 16450 | 771.0938
66 | 16451 | 771.1406
66 | 16450 | 771.0938
66 | 16452 | 771.1875
66 | 16450 | 771.0938
66 | 16453 | 771.2344
66 | 16451 | 771.1406
66 | 16455 | 771.3281
66 | 16451 | 771.1406
66 | 16456 | 771.375
66 | 16451 | 771.1406
66 | 16456 | 771.375
66 | 16451 | 771.1406
66 | 16455 | 771.3281
66 | 16450 | 771.0938
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16455 | 771.3281
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16449 | 771.0469
66 | 16456 | 771.375
66 | 16450 | 771.0938
66 | 16456 | 771.375
66 | 16450 | 771.0938
66 | 16453 | 771.2344
66 | 16453 | 771.2344
66 | 16449 | 771.0469
66 | 16453 | 771.2344
66 | 16449 | 771.0469
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16454 | 771.2812
66 | 16448 | 771
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16456 | 771.375
66 | 16448 | 771
66 | 16457 | 771.4219
66 | 16448 | 771
66 | 16457 | 771.4219
66 | 16446 | 770.9062
66 | 16455 | 771.3281
66 | 16445 | 770.8594
66 | 16453 | 771.2344
66 | 16445 | 770.8594
66 | 16454 | 771.2812
66 | 16446 | 770.9062
66 | 16452 | 771.1875
66 | 16448 | 771
66 | 16454 | 771.2812
66 | 16452 | 771.1875
66 | 16453 | 771.2344
66 | 16453 | 771.2344
66 | 16453 | 771.2344
66 | 16452 | 771.1875
66 | 16452 | 771.1875
66 | 16454 | 771.2812
66 | 16450 | 771.0938
66 | 16455 | 771.3281
66 | 16448 | 771
66 | 16453 | 771.2344
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16445 | 770.8594
66 | 16455 | 771.3281
66 | 16445 | 770.8594
66 | 16456 | 771.375
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16450 | 771.0938
66 | 16453 | 771.2344
66 | 16452 | 771.1875
66 | 16450 | 771.0938
66 | 16454 | 771.2812
66 | 16448 | 771
66 | 16453 | 771.2344
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16444 | 770.8125
66 | 16455 | 771.3281
66 | 16445 | 770.8594
66 | 16455 | 771.3281
66 | 16445 | 770.8594
66 | 16455 | 771.3281
66 | 16447 | 770.9531
66 | 16454 | 771.2812
66 | 16450 | 771.0938
66 | 16453 | 771.2344
66 | 16452 | 771.1875
66 | 16450 | 771.0938
66 | 16452 | 771.1875
66 | 16447 | 770.9531
66 | 16453 | 771.2344
66 | 16445 | 770.8594
66 | 16455 | 771.3281
66 | 16444 | 770.8125
66 | 16454 | 771.2812
66 | 16446 | 770.9062
66 | 16453 | 771.2344
66 | 16450 | 771.0938
66 | 16451 | 771.1406
66 | 16452 | 771.1875
66 | 16449 | 771.0469
66 | 16454 | 771.2812
66 | 16447 | 770.9531
66 | 16453 | 771.2344
66 | 16445 | 770.8594
66 | 16453 | 771.2344
66 | 16444 | 770.8125
66 | 16452 | 771.1875
66 | 16443 | 770.7656
66 | 16453 | 771.2344
66 | 16449 | 771.0469
66 | 16451 | 771.1406
66 | 16451 | 771.1406
66 | 16451 | 771.1406
66 | 16453 | 771.2344
66 | 16450 | 771.0938
66 | 16451 | 771.1406
66 | 16448 | 771
66 | 16452 | 771.1875
66 | 16446 | 770.9062
66 | 16453 | 771.2344
66 | 16444 | 770.8125
66 | 16453 | 771.2344
66 | 16442 | 770.7188
66 | 16453 | 771.2344
66 | 16443 | 770.7656
66 | 16451 | 771.1406
66 | 16450 | 771.0938
66 | 16449 | 771.0469
66 | 16453 | 771.2344
66 | 16448 | 771
66 | 16454 | 771.2812
66 | 16447 | 770.9531
66 | 16455 | 771.3281
66 | 16444 | 770.8125
66 | 16452 | 771.1875
66 | 16445 | 770.8594
66 | 16450 | 771.0938
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16451 | 771.1406
66 | 16444 | 770.8125
66 | 16453 | 771.2344
66 | 16443 | 770.7656
66 | 16451 | 771.1406
66 | 16445 | 770.8594
66 | 16453 | 771.2344
66 | 16447 | 770.9531
66 | 16450 | 771.0938
66 | 16451 | 771.1406
66 | 16448 | 771
66 | 16451 | 771.1406
66 | 16444 | 770.8125
66 | 16453 | 771.2344
66 | 16442 | 770.7188
66 | 16451 | 771.1406
66 | 16443 | 770.7656
66 | 16449 | 771.0469
66 | 16448 | 771
66 | 16448 | 771
66 | 16451 | 771.1406
66 | 16448 | 771
66 | 16453 | 771.2344
66 | 16445 | 770.8594
66 | 16454 | 771.2812
66 | 16443 | 770.7656
66 | 16452 | 771.1875
66 | 16443 | 770.7656
66 | 16449 | 771.0469
66 | 16446 | 770.9062
66 | 16446 | 770.9062
66 | 16450 | 771.0938
66 | 16443 | 770.7656
66 | 16452 | 771.1875
66 | 16442 | 770.7188
66 | 16452 | 771.1875
66 | 16447 | 770.9531
66 | 16450 | 771.0938
66 | 16452 | 771.1875
66 | 16449 | 771.0469
66 | 16453 | 771.2344
66 | 16446 | 770.9062
66 | 16452 | 771.1875
66 | 16443 | 770.7656
66 | 16452 | 771.1875
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16444 | 770.8125
66 | 16447 | 770.9531
66 | 16449 | 771.0469
66 | 16447 | 770.9531
66 | 16453 | 771.2344
66 | 16444 | 770.8125
66 | 16454 | 771.2812
66 | 16443 | 770.7656
66 | 16452 | 771.1875
66 | 16448 | 771
66 | 16448 | 771
66 | 16450 | 771.0938
66 | 16442 | 770.7188
66 | 16451 | 771.1406
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16450 | 771.0938
66 | 16446 | 770.9062
66 | 16452 | 771.1875
66 | 16443 | 770.7656
66 | 16453 | 771.2344
66 | 16446 | 770.9062
66 | 16452 | 771.1875
66 | 16448 | 771
66 | 16448 | 771
66 | 16450 | 771.0938
66 | 16445 | 770.8594
66 | 16452 | 771.1875
66 | 16441 | 770.6719
66 | 16451 | 771.1406
66 | 16443 | 770.7656
66 | 16448 | 771
66 | 16449 | 771.0469
66 | 16446 | 770.9062
66 | 16451 | 771.1406
66 | 16444 | 770.8125
66 | 16454 | 771.2812
66 | 16443 | 770.7656
66 | 16452 | 771.1875
66 | 16448 | 771
66 | 16449 | 771.0469
66 | 16449 | 771.0469
66 | 16444 | 770.8125
66 | 16451 | 771.1406
66 | 16441 | 770.6719
66 | 16448 | 771
66 | 16447 | 770.9531
66 | 16445 | 770.8594
66 | 16450 | 771.0938
66 | 16443 | 770.7656
66 | 16453 | 771.2344
66 | 16442 | 770.7188
66 | 16452 | 771.1875
66 | 16445 | 770.8594
66 | 16449 | 771.0469
66 | 16451 | 771.1406
66 | 16445 | 770.8594
66 | 16452 | 771.1875
66 | 16440 | 770.625
66 | 16449 | 771.0469
66 | 16439 | 770.5781
66 | 16449 | 771.0469
66 | 16442 | 770.7188
66 | 16448 | 771
66 | 16447 | 770.9531
66 | 16445 | 770.8594
66 | 16450 | 771.0938
66 | 16443 | 770.7656
66 | 16453 | 771.2344
66 | 16442 | 770.7188
66 | 16451 | 771.1406
66 | 16444 | 770.8125
66 | 16450 | 771.0938
66 | 16449 | 771.0469
66 | 16446 | 770.9062
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16439 | 770.5781
66 | 16449 | 771.0469
66 | 16443 | 770.7656
66 | 16445 | 770.8594
66 | 16448 | 771
66 | 16444 | 770.8125
66 | 16453 | 771.2344
66 | 16442 | 770.7188
66 | 16452 | 771.1875
66 | 16443 | 770.7656
66 | 16451 | 771.1406
66 | 16446 | 770.9062
66 | 16448 | 771
66 | 16448 | 771
66 | 16443 | 770.7656
66 | 16450 | 771.0938
66 | 16442 | 770.7188
66 | 16450 | 771.0938
66 | 16440 | 770.625
66 | 16451 | 771.1406
66 | 16440 | 770.625
66 | 16450 | 771.0938
66 | 16445 | 770.8594
66 | 16450 | 771.0938
66 | 16449 | 771.0469
66 | 16446 | 770.9062
66 | 16452 | 771.1875
66 | 16443 | 770.7656
66 | 16452 | 771.1875
66 | 16440 | 770.625
66 | 16450 | 771.0938
66 | 16439 | 770.5781
66 | 16449 | 771.0469
66 | 16443 | 770.7656
66 | 16447 | 770.9531
66 | 16445 | 770.8594
66 | 16447 | 770.9531
66 | 16448 | 771
66 | 16445 | 770.8594
66 | 16451 | 771.1406
66 | 16442 | 770.7188
66 | 16451 | 771.1406
66 | 16442 | 770.7188
66 | 16449 | 771.0469
66 | 16444 | 770.8125
66 | 16448 | 771
66 | 16447 | 770.9531
66 | 16442 | 770.7188
66 | 16448 | 771
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16452 | 771.1875
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16445 | 770.8594
66 | 16448 | 771
66 | 16450 | 771.0938
66 | 16444 | 770.8125
66 | 16451 | 771.1406
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16449 | 771.0469
66 | 16440 | 770.625
66 | 16446 | 770.9062
66 | 16444 | 770.8125
66 | 16445 | 770.8594
66 | 16447 | 770.9531
66 | 16444 | 770.8125
66 | 16450 | 771.0938
66 | 16442 | 770.7188
66 | 16451 | 771.1406
66 | 16441 | 770.6719
66 | 16451 | 771.1406
66 | 16441 | 770.6719
66 | 16449 | 771.0469
66 | 16441 | 770.6719
66 | 16438 | 770.5312
66 | 16439 | 770.5781
66 | 16445 | 770.8594
66 | 16439 | 770.5781
66 | 16452 | 771.1875
66 | 16440 | 770.625
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16447 | 770.9531
66 | 16444 | 770.8125
66 | 16442 | 770.7188
66 | 16449 | 771.0469
66 | 16440 | 770.625
66 | 16448 | 771
66 | 16440 | 770.625
66 | 16449 | 771.0469
66 | 16447 | 770.9531
66 | 16446 | 770.9062
66 | 16451 | 771.1406
66 | 16443 | 770.7656
66 | 16451 | 771.1406
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16447 | 770.9531
66 | 16444 | 770.8125
66 | 16443 | 770.7656
66 | 16448 | 771
66 | 16439 | 770.5781
66 | 16449 | 771.0469
66 | 16439 | 770.5781
66 | 16447 | 770.9531
66 | 16442 | 770.7188
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16445 | 770.8594
66 | 16450 | 771.0938
66 | 16442 | 770.7188
66 | 16450 | 771.0938
66 | 16440 | 770.625
66 | 16449 | 771.0469
66 | 16439 | 770.5781
66 | 16446 | 770.9062
66 | 16443 | 770.7656
66 | 16442 | 770.7188
66 | 16447 | 770.9531
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16440 | 770.625
66 | 16450 | 771.0938
66 | 16441 | 770.6719
66 | 16448 | 771
66 | 16447 | 770.9531
66 | 16446 | 770.9062
66 | 16448 | 771
66 | 16441 | 770.6719
66 | 16450 | 771.0938
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16438 | 770.5312
66 | 16446 | 770.9062
66 | 16444 | 770.8125
66 | 16442 | 770.7188
66 | 16447 | 770.9531
66 | 16441 | 770.6719
66 | 16449 | 771.0469
66 | 16439 | 770.5781
66 | 16448 | 771
66 | 16446 | 770.9062
66 | 16446 | 770.9062
66 | 16448 | 771
66 | 16441 | 770.6719
66 | 16448 | 771
66 | 16438 | 770.5312
66 | 16445 | 770.8594
66 | 16442 | 770.7188
66 | 16442 | 770.7188
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16448 | 771
66 | 16439 | 770.5781
66 | 16448 | 771
66 | 16445 | 770.8594
66 | 16444 | 770.8125
66 | 16448 | 771
66 | 16440 | 770.625
66 | 16450 | 771.0938
66 | 16440 | 770.625
66 | 16446 | 770.9062
66 | 16445 | 770.8594
66 | 16441 | 770.6719
66 | 16447 | 770.9531
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16438 | 770.5312
66 | 16446 | 770.9062
66 | 16446 | 770.9062
66 | 16443 | 770.7656
66 | 16450 | 771.0938
66 | 16439 | 770.5781
66 | 16449 | 771.0469
66 | 16447 | 770.9531
66 | 16443 | 770.7656
66 | 16449 | 771.0469
66 | 16440 | 770.625
66 | 16448 | 771
66 | 16441 | 770.6719
66 | 16444 | 770.8125
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16447 | 770.9531
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16444 | 770.8125
66 | 16445 | 770.8594
66 | 16449 | 771.0469
66 | 16440 | 770.625
66 | 16449 | 771.0469
66 | 16443 | 770.7656
66 | 16444 | 770.8125
66 | 16446 | 770.9062
66 | 16441 | 770.6719
66 | 16448 | 771
66 | 16436 | 770.4375
66 | 16446 | 770.9062
66 | 16442 | 770.7188
66 | 16443 | 770.7656
66 | 16446 | 770.9062
66 | 16440 | 770.625
66 | 16449 | 771.0469
66 | 16440 | 770.625
66 | 16447 | 770.9531
66 | 16447 | 770.9531
66 | 16443 | 770.7656
66 | 16448 | 771
66 | 16440 | 770.625
66 | 16447 | 770.9531
66 | 16441 | 770.6719
66 | 16443 | 770.7656
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16446 | 770.9062
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16445 | 770.8594
66 | 16441 | 770.6719
66 | 16449 | 771.0469
66 | 16439 | 770.5781
66 | 16446 | 770.9062
66 | 16443 | 770.7656
66 | 16444 | 770.8125
66 | 16444 | 770.8125
66 | 16441 | 770.6719
66 | 16447 | 770.9531
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16440 | 770.625
66 | 16442 | 770.7188
66 | 16444 | 770.8125
66 | 16438 | 770.5312
66 | 16445 | 770.8594
66 | 16437 | 770.4844
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16446 | 770.9062
66 | 16444 | 770.8125
66 | 16441 | 770.6719
66 | 16448 | 771
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16436 | 770.4375
66 | 16445 | 770.8594
66 | 16442 | 770.7188
66 | 16440 | 770.625
66 | 16443 | 770.7656
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16436 | 770.4375
66 | 16446 | 770.9062
66 | 16443 | 770.7656
66 | 16441 | 770.6719
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16448 | 771
66 | 16438 | 770.5312
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16443 | 770.7656
66 | 16443 | 770.7656
66 | 16439 | 770.5781
66 | 16445 | 770.8594
66 | 16436 | 770.4375
66 | 16445 | 770.8594
66 | 16440 | 770.625
66 | 16443 | 770.7656
66 | 16446 | 770.9062
66 | 16440 | 770.625
66 | 16450 | 771.0938
66 | 16438 | 770.5312
66 | 16444 | 770.8125
66 | 16443 | 770.7656
66 | 16439 | 770.5781
66 | 16446 | 770.9062
66 | 16435 | 770.3906
66 | 16443 | 770.7656
66 | 16442 | 770.7188
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16436 | 770.4375
66 | 16446 | 770.9062
66 | 16442 | 770.7188
66 | 16441 | 770.6719
66 | 16447 | 770.9531
66 | 16440 | 770.625
66 | 16447 | 770.9531
66 | 16438 | 770.5312
66 | 16444 | 770.8125
66 | 16443 | 770.7656
66 | 16438 | 770.5312
66 | 16445 | 770.8594
66 | 16435 | 770.3906
66 | 16443 | 770.7656
66 | 16441 | 770.6719
66 | 16440 | 770.625
66 | 16446 | 770.9062
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16436 | 770.4375
66 | 16446 | 770.9062
66 | 16440 | 770.625
66 | 16442 | 770.7188
66 | 16444 | 770.8125
66 | 16439 | 770.5781
66 | 16444 | 770.8125
66 | 16435 | 770.3906
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16443 | 770.7656
66 | 16441 | 770.6719
66 | 16439 | 770.5781
66 | 16446 | 770.9062
66 | 16438 | 770.5312
66 | 16446 | 770.9062
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16444 | 770.8125
66 | 16440 | 770.625
66 | 16446 | 770.9062
66 | 16435 | 770.3906
66 | 16444 | 770.8125
66 | 16442 | 770.7188
66 | 16439 | 770.5781
66 | 16444 | 770.8125
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16435 | 770.3906
66 | 16443 | 770.7656
66 | 16438 | 770.5312
66 | 16441 | 770.6719
66 | 16446 | 770.9062
66 | 16438 | 770.5312
66 | 16447 | 770.9531
66 | 16435 | 770.3906
66 | 16444 | 770.8125
66 | 16438 | 770.5312
66 | 16441 | 770.6719
66 | 16441 | 770.6719
66 | 16438 | 770.5312
66 | 16444 | 770.8125
66 | 16437 | 770.4844
66 | 16446 | 770.9062
66 | 16434 | 770.3438
66 | 16446 | 770.9062
66 | 16439 | 770.5781
66 | 16444 | 770.8125
66 | 16445 | 770.8594
66 | 16439 | 770.5781
66 | 16445 | 770.8594
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16434 | 770.3438
66 | 16443 | 770.7656
66 | 16440 | 770.625
66 | 16438 | 770.5312
66 | 16443 | 770.7656
66 | 16435 | 770.3906
66 | 16445 | 770.8594
66 | 16435 | 770.3906
66 | 16445 | 770.8594
66 | 16439 | 770.5781
66 | 16442 | 770.7188
66 | 16443 | 770.7656
66 | 16439 | 770.5781
66 | 16445 | 770.8594
66 | 16437 | 770.4844
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16442 | 770.7188
66 | 16438 | 770.5312
66 | 16439 | 770.5781
66 | 16443 | 770.7656
66 | 16437 | 770.4844
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16446 | 770.9062
66 | 16437 | 770.4844
66 | 16445 | 770.8594
66 | 16441 | 770.6719
66 | 16439 | 770.5781
66 | 16443 | 770.7656
66 | 16436 | 770.4375
66 | 16444 | 770.8125
66 | 16433 | 770.2969
66 | 16442 | 770.7188
66 | 16435 | 770.3906
66 | 16439 | 770.5781
66 | 16441 | 770.6719
66 | 16437 | 770.4844
66 | 16444 | 770.8125
66 | 16435 | 770.3906
66 | 16444 | 770.8125
66 | 16435 | 770.3906
66 | 16444 | 770.8125
66 | 16442 | 770.7188
66 | 16440 | 770.625
66 | 16444 | 770.8125
66 | 16437 | 770.4844
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16442 | 770.7188
66 | 16437 | 770.4844
66 | 16441 | 770.6719
66 | 16441 | 770.6719
66 | 16439 | 770.5781
66 | 16445 | 770.8594
66 | 16435 | 770.3906
66 | 16445 | 770.8594
66 | 16437 | 770.4844
66 | 16440 | 770.625
66 | 16444 | 770.8125
66 | 16435 | 770.3906
66 | 16445 | 770.8594
66 | 16434 | 770.3438
66 | 16443 | 770.7656
66 | 16438 | 770.5312
66 | 16440 | 770.625
66 | 16442 | 770.7188
66 | 16436 | 770.4375
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16444 | 770.8125
66 | 16436 | 770.4375
66 | 16443 | 770.7656
66 | 16441 | 770.6719
66 | 16439 | 770.5781
66 | 16445 | 770.8594
66 | 16436 | 770.4375
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16441 | 770.6719
66 | 16435 | 770.3906
66 | 16440 | 770.625
66 | 16438 | 770.5312
66 | 16437 | 770.4844
66 | 16443 | 770.7656
66 | 16436 | 770.4375
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16445 | 770.8594
66 | 16437 | 770.4844
66 | 16443 | 770.7656
66 | 16440 | 770.625
66 | 16438 | 770.5312
66 | 16443 | 770.7656
66 | 16433 | 770.2969
66 | 16444 | 770.8125
66 | 16433 | 770.2969
66 | 16441 | 770.6719
66 | 16437 | 770.4844
66 | 16438 | 770.5312
66 | 16440 | 770.625
66 | 16438 | 770.5312
66 | 16443 | 770.7656
66 | 16437 | 770.4844
66 | 16444 | 770.8125
66 | 16436 | 770.4375
66 | 16446 | 770.9062
66 | 16434 | 770.3438
66 | 16443 | 770.7656
66 | 16439 | 770.5781
66 | 16438 | 770.5312
66 | 16441 | 770.6719
66 | 16435 | 770.3906
66 | 16444 | 770.8125
66 | 16432 | 770.25
66 | 16442 | 770.7188
66 | 16437 | 770.4844
66 | 16438 | 770.5312
66 | 16442 | 770.7188
66 | 16435 | 770.3906
66 | 16445 | 770.8594
66 | 16434 | 770.3438
66 | 16444 | 770.8125
66 | 16438 | 770.5312
66 | 16439 | 770.5781
66 | 16442 | 770.7188
66 | 16435 | 770.3906
66 | 16442 | 770.7188
66 | 16432 | 770.25
66 | 16442 | 770.7188
66 | 16433 | 770.2969
66 | 16441 | 770.6719
66 | 16438 | 770.5312
66 | 16437 | 770.4844
66 | 16441 | 770.6719
66 | 16436 | 770.4375
66 | 16444 | 770.8125
66 | 16435 | 770.3906
66 | 16443 | 770.7656
66 | 16433 | 770.2969
66 | 16440 | 770.625
66 | 16439 | 770.5781
66 | 16436 | 770.4375
66 | 16440 | 770.625
66 | 16433 | 770.2969
66 | 16442 | 770.7188
66 | 16431 | 770.2031
66 | 16440 | 770.625
66 | 16436 | 770.4375
66 | 16439 | 770.5781
66 | 16442 | 770.7188
66 | 16435 | 770.3906
66 | 16443 | 770.7656
66 | 16433 | 770.2969
66 | 16442 | 770.7188
66 | 16436 | 770.4375
66 | 16439 | 770.5781
66 | 16440 | 770.625
66 | 16434 | 770.3438
66 | 16442 | 770.7188
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16434 | 770.3438
66 | 16438 | 770.5312
66 | 16440 | 770.625
66 | 16436 | 770.4375
66 | 16444 | 770.8125
66 | 16434 | 770.3438
66 | 16443 | 770.7656
66 | 16437 | 770.4844
66 | 16441 | 770.6719
66 | 16440 | 770.625
66 | 16436 | 770.4375
66 | 16442 | 770.7188
66 | 16431 | 770.2031
66 | 16441 | 770.6719
66 | 16435 | 770.3906
66 | 16438 | 770.5312
66 | 16438 | 770.5312
66 | 16434 | 770.3438
66 | 16442 | 770.7188
66 | 16432 | 770.25
66 | 16443 | 770.7656
66 | 16433 | 770.2969
66 | 16442 | 770.7188
66 | 16433 | 770.2969
66 | 16441 | 770.6719
66 | 16440 | 770.625
66 | 16437 | 770.4844
66 | 16440 | 770.625
66 | 16433 | 770.2969
66 | 16442 | 770.7188
66 | 16430 | 770.1562
66 | 16438 | 770.5312
66 | 16436 | 770.4375
66 | 16435 | 770.3906
66 | 16441 | 770.6719
66 | 16434 | 770.3438
66 | 16442 | 770.7188
66 | 16432 | 770.25
66 | 16440 | 770.625
66 | 16438 | 770.5312
66 | 16437 | 770.4844
66 | 16443 | 770.7656
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16433 | 770.2969
66 | 16436 | 770.4375
66 | 16438 | 770.5312
66 | 16432 | 770.25
66 | 16440 | 770.625
66 | 16432 | 770.25
66 | 16439 | 770.5781
66 | 16438 | 770.5312
66 | 16436 | 770.4375
66 | 16442 | 770.7188
66 | 16433 | 770.2969
66 | 16442 | 770.7188
66 | 16438 | 770.5312
66 | 16437 | 770.4844
66 | 16441 | 770.6719
66 | 16433 | 770.2969
66 | 16440 | 770.625
66 | 16431 | 770.2031
66 | 16439 | 770.5781
66 | 16435 | 770.3906
66 | 16435 | 770.3906
66 | 16439 | 770.5781
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16432 | 770.25
66 | 16440 | 770.625
66 | 16439 | 770.5781
66 | 16438 | 770.5312
66 | 16440 | 770.625
66 | 16435 | 770.3906
66 | 16442 | 770.7188
66 | 16432 | 770.25
66 | 16440 | 770.625
66 | 16434 | 770.3438
66 | 16434 | 770.3438
66 | 16438 | 770.5312
66 | 16430 | 770.1562
66 | 16440 | 770.625
66 | 16433 | 770.2969
66 | 16437 | 770.4844
66 | 16441 | 770.6719
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16434 | 770.3438
66 | 16438 | 770.5312
66 | 16439 | 770.5781
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16432 | 770.25
66 | 16437 | 770.4844
66 | 16437 | 770.4844
66 | 16433 | 770.2969
66 | 16441 | 770.6719
66 | 16429 | 770.1094
66 | 16439 | 770.5781
66 | 16438 | 770.5312
66 | 16435 | 770.3906
66 | 16441 | 770.6719
66 | 16432 | 770.25
66 | 16440 | 770.625
66 | 16438 | 770.5312
66 | 16436 | 770.4375
66 | 16441 | 770.6719
66 | 16429 | 770.1094
66 | 16439 | 770.5781
66 | 16432 | 770.25
66 | 16434 | 770.3438
66 | 16438 | 770.5312
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16429 | 770.1094
66 | 16438 | 770.5312
66 | 16437 | 770.4844
66 | 16436 | 770.4375
66 | 16441 | 770.6719
66 | 16431 | 770.2031
66 | 16439 | 770.5781
66 | 16437 | 770.4844
66 | 16435 | 770.3906
66 | 16439 | 770.5781
66 | 16429 | 770.1094
66 | 16438 | 770.5312
66 | 16433 | 770.2969
66 | 16433 | 770.2969
66 | 16438 | 770.5312
66 | 16430 | 770.1562
66 | 16439 | 770.5781
66 | 16433 | 770.2969
66 | 16437 | 770.4844
66 | 16438 | 770.5312
66 | 16435 | 770.3906
66 | 16441 | 770.6719
66 | 16430 | 770.1562
66 | 16439 | 770.5781
66 | 16437 | 770.4844
66 | 16434 | 770.3438
66 | 16440 | 770.625
66 | 16431 | 770.2031
66 | 16438 | 770.5312
66 | 16431 | 770.2031
66 | 16435 | 770.3906
66 | 16436 | 770.4375
66 | 16432 | 770.25
66 | 16440 | 770.625
66 | 16431 | 770.2031
66 | 16441 | 770.6719
66 | 16436 | 770.4375
66 | 16437 | 770.4844
66 | 16441 | 770.6719
66 | 16430 | 770.1562
66 | 16440 | 770.625
66 | 16431 | 770.2031
66 | 16438 | 770.5312
66 | 16435 | 770.3906
66 | 16432 | 770.25
66 | 16439 | 770.5781
66 | 16428 | 770.0625
66 | 16438 | 770.5312
66 | 16433 | 770.2969
66 | 16435 | 770.3906
66 | 16439 | 770.5781
66 | 16433 | 770.2969
66 | 16440 | 770.625
66 | 16431 | 770.2031
66 | 16439 | 770.5781
66 | 16436 | 770.4375
66 | 16434 | 770.3438
66 | 16440 | 770.625
66 | 16430 | 770.1562
66 | 16438 | 770.5312
66 | 16431 | 770.2031
66 | 16437 | 770.4844
66 | 16436 | 770.4375
66 | 16431 | 770.2031
66 | 16439 | 770.5781
66 | 16427 | 770.0156
66 | 16438 | 770.5312
66 | 16432 | 770.25
66 | 16438 | 770.5312
66 | 16437 | 770.4844
66 | 16434 | 770.3438
66 | 16441 | 770.6719
66 | 16430 | 770.1562
66 | 16440 | 770.625
66 | 16431 | 770.2031
66 | 16437 | 770.4844
66 | 16436 | 770.4375
66 | 16431 | 770.2031
66 | 16436 | 770.4375
66 | 16427 | 770.0156
66 | 16438 | 770.5312
66 | 16431 | 770.2031
66 | 16434 | 770.3438
66 | 16438 | 770.5312
66 | 16432 | 770.25
66 | 16441 | 770.6719
66 | 16431 | 770.2031
66 | 16440 | 770.625
66 | 16430 | 770.1562
66 | 16438 | 770.5312
66 | 16435 | 770.3906
66 | 16434 | 770.3438
66 | 16437 | 770.4844
66 | 16429 | 770.1094
66 | 16439 | 770.5781
66 | 16427 | 770.0156
66 | 16436 | 770.4375
66 | 16432 | 770.25
66 | 16433 | 770.2969
66 | 16437 | 770.4844
66 | 16432 | 770.25
66 | 16439 | 770.5781
66 | 16429 | 770.1094
66 | 16440 | 770.625
66 | 16435 | 770.3906
66 | 16436 | 770.4375
66 | 16440 | 770.625
66 | 16433 | 770.2969
66 | 16439 | 770.5781
66 | 16430 | 770.1562
66 | 16437 | 770.4844
66 | 16430 | 770.1562
66 | 16434 | 770.3438
66 | 16436 | 770.4375
66 | 16430 | 770.1562
66 | 16438 | 770.5312
66 | 16428 | 770.0625
66 | 16439 | 770.5781
66 | 16434 | 770.3438
66 | 16436 | 770.4375
66 | 16440 | 770.625
66 | 16432 | 770.25
66 | 16439 | 770.5781
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16437 | 770.4844
66 | 16433 | 770.2969
66 | 16436 | 770.4375
66 | 16427 | 770.0156
66 | 16437 | 770.4844
66 | 16431 | 770.2031
66 | 16432 | 770.25
66 | 16438 | 770.5312
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16434 | 770.3438
66 | 16432 | 770.25
66 | 16439 | 770.5781
66 | 16430 | 770.1562
66 | 16439 | 770.5781
66 | 16431 | 770.2031
66 | 16434 | 770.3438
66 | 16436 | 770.4375
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16435 | 770.3906
66 | 16430 | 770.1562
66 | 16439 | 770.5781
66 | 16428 | 770.0625
66 | 16439 | 770.5781
66 | 16436 | 770.4375
66 | 16434 | 770.3438
66 | 16440 | 770.625
66 | 16429 | 770.1094
66 | 16438 | 770.5312
66 | 16435 | 770.3906
66 | 16432 | 770.25
66 | 16436 | 770.4375
66 | 16427 | 770.0156
66 | 16437 | 770.4844
66 | 16430 | 770.1562
66 | 16432 | 770.25
66 | 16437 | 770.4844
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16429 | 770.1094
66 | 16436 | 770.4375
66 | 16435 | 770.3906
66 | 16432 | 770.25
66 | 16438 | 770.5312
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16433 | 770.2969
66 | 16431 | 770.2031
66 | 16436 | 770.4375
66 | 16428 | 770.0625
66 | 16436 | 770.4375
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16434 | 770.3438
66 | 16430 | 770.1562
66 | 16435 | 770.3906
66 | 16429 | 770.1094
66 | 16437 | 770.4844
66 | 16428 | 770.0625
66 | 16436 | 770.4375
66 | 16436 | 770.4375
66 | 16431 | 770.2031
66 | 16437 | 770.4844
66 | 16427 | 770.0156
66 | 16436 | 770.4375
66 | 16431 | 770.2031
66 | 16430 | 770.1562
66 | 16435 | 770.3906
66 | 16426 | 769.9688
66 | 16435 | 770.3906
66 | 16428 | 770.0625
66 | 16435 | 770.3906
66 | 16435 | 770.3906
66 | 16431 | 770.2031
66 | 16437 | 770.4844
66 | 16427 | 770.0156
66 | 16436 | 770.4375
66 | 16432 | 770.25
66 | 16432 | 770.25
66 | 16436 | 770.4375
66 | 16428 | 770.0625
66 | 16436 | 770.4375
66 | 16431 | 770.2031
66 | 16431 | 770.2031
66 | 16434 | 770.3438
66 | 16425 | 769.9219
66 | 16434 | 770.3438
66 | 16429 | 770.1094
66 | 16432 | 770.25
66 | 16435 | 770.3906
66 | 16429 | 770.1094
66 | 16438 | 770.5312
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16434 | 770.3438
66 | 16430 | 770.1562
66 | 16437 | 770.4844
66 | 16426 | 769.9688
66 | 16434 | 770.3438
66 | 16433 | 770.2969
66 | 16428 | 770.0625
66 | 16435 | 770.3906
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16434 | 770.3438
66 | 16429 | 770.1094
66 | 16439 | 770.5781
66 | 16428 | 770.0625
66 | 16435 | 770.3906
66 | 16436 | 770.4375
66 | 16429 | 770.1094
66 | 16438 | 770.5312
66 | 16428 | 770.0625
66 | 16432 | 770.25
66 | 16435 | 770.3906
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16431 | 770.2031
66 | 16430 | 770.1562
66 | 16437 | 770.4844
66 | 16427 | 770.0156
66 | 16435 | 770.3906
66 | 16430 | 770.1562
66 | 16432 | 770.25
66 | 16437 | 770.4844
66 | 16428 | 770.0625
66 | 16438 | 770.5312
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16432 | 770.25
66 | 16428 | 770.0625
66 | 16435 | 770.3906
66 | 16424 | 769.875
66 | 16432 | 770.25
66 | 16431 | 770.2031
66 | 16431 | 770.2031
66 | 16435 | 770.3906
66 | 16426 | 769.9688
66 | 16434 | 770.3438
66 | 16431 | 770.2031
66 | 16431 | 770.2031
66 | 16436 | 770.4375
66 | 16427 | 770.0156
66 | 16436 | 770.4375
66 | 16430 | 770.1562
66 | 16430 | 770.1562
66 | 16434 | 770.3438
66 | 16425 | 769.9219
66 | 16434 | 770.3438
66 | 16428 | 770.0625
66 | 16429 | 770.1094
66 | 16433 | 770.2969
66 | 16427 | 770.0156
66 | 16435 | 770.3906
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16434 | 770.3438
66 | 16429 | 770.1094
66 | 16437 | 770.4844
66 | 16426 | 769.9688
66 | 16434 | 770.3438
66 | 16433 | 770.2969
66 | 16426 | 769.9688
66 | 16434 | 770.3438
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16434 | 770.3438
66 | 16425 | 769.9219
66 | 16434 | 770.3438
66 | 16432 | 770.25
66 | 16427 | 770.0156
66 | 16437 | 770.4844
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16435 | 770.3906
66 | 16429 | 770.1094
66 | 16437 | 770.4844
66 | 16424 | 769.875
66 | 16432 | 770.25
66 | 16432 | 770.25
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16433 | 770.2969
66 | 16427 | 770.0156
66 | 16435 | 770.3906
66 | 16428 | 770.0625
66 | 16433 | 770.2969
66 | 16435 | 770.3906
66 | 16428 | 770.0625
66 | 16437 | 770.4844
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16433 | 770.2969
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16424 | 769.875
66 | 16432 | 770.25
66 | 16428 | 770.0625
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16425 | 769.9219
66 | 16436 | 770.4375
66 | 16429 | 770.1094
66 | 16433 | 770.2969
66 | 16436 | 770.4375
66 | 16427 | 770.0156
66 | 16436 | 770.4375
66 | 16432 | 770.25
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16424 | 769.875
66 | 16433 | 770.2969
66 | 16430 | 770.1562
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16430 | 770.1562
66 | 16429 | 770.1094
66 | 16437 | 770.4844
66 | 16426 | 769.9688
66 | 16434 | 770.3438
66 | 16434 | 770.3438
66 | 16427 | 770.0156
66 | 16435 | 770.3906
66 | 16428 | 770.0625
66 | 16428 | 770.0625
66 | 16433 | 770.2969
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16431 | 770.2031
66 | 16428 | 770.0625
66 | 16434 | 770.3438
66 | 16424 | 769.875
66 | 16432 | 770.25
66 | 16431 | 770.2031
66 | 16426 | 769.9688
66 | 16434 | 770.3438
66 | 16425 | 769.9219
66 | 16433 | 770.2969
66 | 16434 | 770.3438
66 | 16427 | 770.0156
66 | 16435 | 770.3906
66 | 16425 | 769.9219
66 | 16427 | 770.0156
66 | 16430 | 770.1562
66 | 16423 | 769.8281
66 | 16433 | 770.2969
66 | 16425 | 769.9219
66 | 16428 | 770.0625
66 | 16432 | 770.25
66 | 16424 | 769.875
66 | 16433 | 770.2969
66 | 16430 | 770.1562
66 | 16428 | 770.0625
66 | 16433 | 770.2969
66 | 16423 | 769.8281
66 | 16433 | 770.2969
66 | 16430 | 770.1562
66 | 16427 | 770.0156
66 | 16432 | 770.25
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16423 | 769.8281
66 | 16430 | 770.1562
66 | 16431 | 770.2031
66 | 16427 | 770.0156
66 | 16433 | 770.2969
66 | 16424 | 769.875
66 | 16432 | 770.25
66 | 16430 | 770.1562
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16430 | 770.1562
66 | 16430 | 770.1562
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16427 | 770.0156
66 | 16429 | 770.1094
66 | 16434 | 770.3438
66 | 16425 | 769.9219
66 | 16433 | 770.2969
66 | 16433 | 770.2969
66 | 16427 | 770.0156
66 | 16435 | 770.3906
66 | 16423 | 769.8281
66 | 16430 | 770.1562
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16431 | 770.2031
66 | 16424 | 769.875
66 | 16435 | 770.3906
66 | 16424 | 769.875
66 | 16430 | 770.1562
66 | 16431 | 770.2031
66 | 16427 | 770.0156
66 | 16434 | 770.3438
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16432 | 770.25
66 | 16421 | 769.7344
66 | 16429 | 770.1094
66 | 16427 | 770.0156
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16430 | 770.1562
66 | 16428 | 770.0625
66 | 16435 | 770.3906
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16428 | 770.0625
66 | 16424 | 769.875
66 | 16429 | 770.1094
66 | 16420 | 769.6875
66 | 16430 | 770.1562
66 | 16427 | 770.0156
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16432 | 770.25
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16435 | 770.3906
66 | 16422 | 769.7812
66 | 16432 | 770.25
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16430 | 770.1562
66 | 16431 | 770.2031
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16425 | 769.9219
66 | 16429 | 770.1094
66 | 16432 | 770.25
66 | 16425 | 769.9219
66 | 16433 | 770.2969
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16425 | 769.9219
66 | 16430 | 770.1562
66 | 16421 | 769.7344
66 | 16429 | 770.1094
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16428 | 770.0625
66 | 16427 | 770.0156
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16425 | 769.9219
66 | 16425 | 769.9219
66 | 16429 | 770.1094
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16421 | 769.7344
66 | 16431 | 770.2031
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16434 | 770.3438
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16428 | 770.0625
66 | 16427 | 770.0156
66 | 16431 | 770.2031
66 | 16423 | 769.8281
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16428 | 770.0625
66 | 16429 | 770.1094
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16421 | 769.7344
66 | 16432 | 770.25
66 | 16430 | 770.1562
66 | 16426 | 769.9688
66 | 16432 | 770.25
66 | 16422 | 769.7812
66 | 16431 | 770.2031
66 | 16431 | 770.2031
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16427 | 770.0156
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16421 | 769.7344
66 | 16429 | 770.1094
66 | 16429 | 770.1094
66 | 16425 | 769.9219
66 | 16433 | 770.2969
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16430 | 770.1562
66 | 16426 | 769.9688
66 | 16433 | 770.2969
66 | 16422 | 769.7812
66 | 16430 | 770.1562
66 | 16425 | 769.9219
66 | 16425 | 769.9219
66 | 16429 | 770.1094
66 | 16421 | 769.7344
66 | 16431 | 770.2031
66 | 16424 | 769.875
66 | 16425 | 769.9219
66 | 16430 | 770.1562
66 | 16422 | 769.7812
66 | 16433 | 770.2969
66 | 16425 | 769.9219
66 | 16429 | 770.1094
66 | 16433 | 770.2969
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16424 | 769.875
66 | 16424 | 769.875
66 | 16429 | 770.1094
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16424 | 769.875
66 | 16423 | 769.8281
66 | 16430 | 770.1562
66 | 16419 | 769.6406
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16428 | 770.0625
66 | 16432 | 770.25
66 | 16423 | 769.8281
66 | 16429 | 770.1094
66 | 16426 | 769.9688
66 | 16425 | 769.9219
66 | 16430 | 770.1562
66 | 16420 | 769.6875
66 | 16426 | 769.9688
66 | 16426 | 769.9688
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16419 | 769.6406
66 | 16430 | 770.1562
66 | 16428 | 770.0625
66 | 16425 | 769.9219
66 | 16433 | 770.2969
66 | 16422 | 769.7812
66 | 16428 | 770.0625
66 | 16430 | 770.1562
66 | 16421 | 769.7344
66 | 16428 | 770.0625
66 | 16423 | 769.8281
66 | 16425 | 769.9219
66 | 16428 | 770.0625
66 | 16418 | 769.5938
66 | 16429 | 770.1094
66 | 16424 | 769.875
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16422 | 769.7812
66 | 16431 | 770.2031
66 | 16422 | 769.7812
66 | 16430 | 770.1562
66 | 16429 | 770.1094
66 | 16423 | 769.8281
66 | 16430 | 770.1562
66 | 16423 | 769.8281
66 | 16426 | 769.9688
66 | 16427 | 770.0156
66 | 16420 | 769.6875
66 | 16430 | 770.1562
66 | 16423 | 769.8281
66 | 16425 | 769.9219
66 | 16429 | 770.1094
66 | 16421 | 769.7344
66 | 16427 | 770.0156
66 | 16427 | 770.0156
66 | 16426 | 769.9688
66 | 16431 | 770.2031
66 | 16420 | 769.6875
66 | 16427 | 770.0156
66 | 16427 | 770.0156
66 | 16422 | 769.7812
66 | 16428 | 770.0625
66 | 16419 | 769.6406
66 | 16427 | 770.0156
66 | 16422 | 769.7812
66 | 16424 | 769.875
66 | 16428 | 770.0625
66 | 16421 | 769.7344
66 | 16430 | 770.1562
66 | 16422 | 769.7812
66 | 16430 | 770.1562
66 | 16430 | 770.1562
66 | 16422 | 769.7812
66 | 16430 | 770.1562
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16424 | 769.875
66 | 16422 | 769.7812
66 | 16428 | 770.0625
66 | 16418 | 769.5938
66 | 16427 | 770.0156
66 | 16426 | 769.9688
66 | 16424 | 769.875
66 | 16428 | 770.0625
66 | 16421 | 769.7344
66 | 16429 | 770.1094
66 | 16427 | 770.0156
66 | 16424 | 769.875
66 | 16431 | 770.2031
66 | 16420 | 769.6875
66 | 16429 | 770.1094
66 | 16428 | 770.0625
66 | 16421 | 769.7344
66 | 16429 | 770.1094
66 | 16419 | 769.6406
66 | 16425 | 769.9219
66 | 16427 | 770.0156
66 | 16420 | 769.6875
66 | 16429 | 770.1094
66 | 16421 | 769.7344
66 | 16425 | 769.9219
66 | 16428 | 770.0625
66 | 16423 | 769.8281
66 | 16431 | 770.2031
66 | 16420 | 769.6875
66 | 16427 | 770.0156
66 | 16424 | 769.875
66 | 16422 | 769.7812
66 | 16427 | 770.0156
66 | 16418 | 769.5938
66 | 16427 | 770.0156
66 | 16423 | 769.8281
66 | 16423 | 769.8281
66 | 16429 | 770.1094
66 | 16418 | 769.5938
66 | 16428 | 770.0625
66 | 16424 | 769.875
66 | 16425 | 769.9219
66 | 16430 | 770.1562
66 | 16421 | 769.7344
66 | 16430 | 770.1562
66 | 16422 | 769.7812
66 | 16426 | 769.9688
66 | 16427 | 770.0156
66 | 16418 | 769.5938
66 | 16427 | 770.0156
66 | 16419 | 769.6406
66 | 16425 | 769.9219
66 | 16427 | 770.0156
66 | 16420 | 769.6875
66 | 16428 | 770.0625
66 | 16423 | 769.8281
66 | 16426 | 769.9688
66 | 16428 | 770.0625
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16420 | 769.6875
66 | 16426 | 769.9688
66 | 16427 | 770.0156
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16418 | 769.5938
66 | 16425 | 769.9219
66 | 16423 | 769.8281
66 | 16420 | 769.6875
66 | 16428 | 770.0625
66 | 16418 | 769.5938
66 | 16427 | 770.0156
66 | 16425 | 769.9219
66 | 16421 | 769.7344
66 | 16429 | 770.1094
66 | 16419 | 769.6406
66 | 16426 | 769.9688
66 | 16426 | 769.9688
66 | 16418 | 769.5938
66 | 16427 | 770.0156
66 | 16421 | 769.7344
66 | 16423 | 769.8281
66 | 16427 | 770.0156
66 | 16417 | 769.5469
66 | 16427 | 770.0156
66 | 16425 | 769.9219
66 | 16421 | 769.7344
66 | 16430 | 770.1562
66 | 16418 | 769.5938
66 | 16428 | 770.0625
66 | 16427 | 770.0156
66 | 16421 | 769.7344
66 | 16430 | 770.1562
66 | 16418 | 769.5938
66 | 16425 | 769.9219
66 | 16426 | 769.9688
66 | 16419 | 769.6406
66 | 16427 | 770.0156
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16426 | 769.9688
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16428 | 770.0625
66 | 16422 | 769.7812
66 | 16429 | 770.1094
66 | 16419 | 769.6406
66 | 16427 | 770.0156
66 | 16423 | 769.8281
66 | 16420 | 769.6875
66 | 16428 | 770.0625
66 | 16416 | 769.5
66 | 16425 | 769.9219
66 | 16424 | 769.875
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16417 | 769.5469
66 | 16425 | 769.9219
66 | 16426 | 769.9688
66 | 16420 | 769.6875
66 | 16429 | 770.1094
66 | 16420 | 769.6875
66 | 16426 | 769.9688
66 | 16424 | 769.875
66 | 16419 | 769.6406
66 | 16427 | 770.0156
66 | 16419 | 769.6406
66 | 16422 | 769.7812
66 | 16426 | 769.9688
66 | 16416 | 769.5
66 | 16426 | 769.9688
66 | 16423 | 769.8281
66 | 16421 | 769.7344
66 | 16427 | 770.0156
66 | 16417 | 769.5469
66 | 16426 | 769.9688
66 | 16424 | 769.875
66 | 16422 | 769.7812
66 | 16428 | 770.0625
66 | 16418 | 769.5938
66 | 16426 | 769.9688
66 | 16426 | 769.9688
66 | 16418 | 769.5938
66 | 16426 | 769.9688
66 | 16419 | 769.6406
66 | 16420 | 769.6875
66 | 16426 | 769.9688
66 | 16417 | 769.5469
66 | 16425 | 769.9219
66 | 16424 | 769.875
66 | 16420 | 769.6875
66 | 16430 | 770.1562
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16428 | 770.0625
66 | 16418 | 769.5938
66 | 16425 | 769.9219
66 | 16425 | 769.9219
66 | 16418 | 769.5938
66 | 16426 | 769.9688
66 | 16417 | 769.5469
66 | 16422 | 769.7812
66 | 16424 | 769.875
66 | 16417 | 769.5469
66 | 16425 | 769.9219
66 | 16418 | 769.5938
66 | 16426 | 769.9688
66 | 16426 | 769.9688
66 | 16421 | 769.7344
66 | 16427 | 770.0156
66 | 16418 | 769.5938
66 | 16425 | 769.9219
66 | 16426 | 769.9688
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16415 | 769.4531
66 | 16424 | 769.875
66 | 16422 | 769.7812
66 | 16419 | 769.6406
66 | 16426 | 769.9688
66 | 16415 | 769.4531
66 | 16424 | 769.875
66 | 16423 | 769.8281
66 | 16420 | 769.6875
66 | 16427 | 770.0156
66 | 16417 | 769.5469
66 | 16427 | 770.0156
66 | 16424 | 769.875
66 | 16421 | 769.7344
66 | 16426 | 769.9688
66 | 16416 | 769.5
66 | 16423 | 769.8281
66 | 16422 | 769.7812
66 | 16418 | 769.5938
66 | 16425 | 769.9219
66 | 16418 | 769.5938
66 | 16420 | 769.6875
66 | 16425 | 769.9219
66 | 16416 | 769.5
66 | 16425 | 769.9219
66 | 16424 | 769.875
66 | 16419 | 769.6406
66 | 16428 | 770.0625
66 | 16421 | 769.7344
66 | 16421 | 769.7344
66 | 16427 | 770.0156
66 | 16418 | 769.5938
66 | 16426 | 769.9688
66 | 16418 | 769.5938
66 | 16422 | 769.7812
66 | 16423 | 769.8281
66 | 16417 | 769.5469
66 | 16424 | 769.875
66 | 16416 | 769.5
66 | 16424 | 769.875
66 | 16423 | 769.8281
66 | 16419 | 769.6406
66 | 16426 | 769.9688
66 | 16417 | 769.5469
66 | 16426 | 769.9688
66 | 16425 | 769.9219
66 | 16417 | 769.5469
66 | 16427 | 770.0156
66 | 16419 | 769.6406
66 | 16423 | 769.8281
66 | 16423 | 769.8281
66 | 16417 | 769.5469
66 | 16424 | 769.875
66 | 16417 | 769.5469
66 | 16421 | 769.7344
66 | 16426 | 769.9688
66 | 16417 | 769.5469
66 | 16425 | 769.9219
66 | 16424 | 769.875
66 | 16421 | 769.7344
66 | 16427 | 770.0156
66 | 16418 | 769.5938
66 | 16425 | 769.9219
66 | 16420 | 769.6875
66 | 16422 | 769.7812
66 | 16425 | 769.9219
66 | 16417 | 769.5469
66 | 16423 | 769.8281
66 | 16416 | 769.5
66 | 16422 | 769.7812
66 | 16423 | 769.8281
66 | 16415 | 769.4531
66 | 16426 | 769.9688
66 | 16418 | 769.5938
66 | 16421 | 769.7344
66 | 16425 | 769.9219
66 | 16417 | 769.5469
66 | 16425 | 769.9219
66 | 16422 | 769.7812
66 | 16419 | 769.6406
66 | 16425 | 769.9219
66 | 16415 | 769.4531
66 | 16423 | 769.8281
66 | 16423 | 769.8281
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16417 | 769.5469
66 | 16417 | 769.5469
66 | 16424 | 769.875
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16423 | 769.8281
66 | 16417 | 769.5469
66 | 16426 | 769.9688
66 | 16421 | 769.7344
66 | 16419 | 769.6406
66 | 16425 | 769.9219
66 | 16415 | 769.4531
66 | 16423 | 769.8281
66 | 16421 | 769.7344
66 | 16415 | 769.4531
66 | 16424 | 769.875
66 | 16418 | 769.5938
66 | 16418 | 769.5938
66 | 16423 | 769.8281
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16423 | 769.8281
66 | 16416 | 769.5
66 | 16426 | 769.9688
66 | 16417 | 769.5469
66 | 16421 | 769.7344
66 | 16424 | 769.875
66 | 16415 | 769.4531
66 | 16422 | 769.7812
66 | 16418 | 769.5938
66 | 16418 | 769.5938
66 | 16423 | 769.8281
66 | 16412 | 769.3125
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16417 | 769.5469
66 | 16425 | 769.9219
66 | 16416 | 769.5
66 | 16422 | 769.7812
66 | 16425 | 769.9219
66 | 16415 | 769.4531
66 | 16425 | 769.9219
66 | 16423 | 769.8281
66 | 16417 | 769.5469
66 | 16426 | 769.9688
66 | 16414 | 769.4062
66 | 16422 | 769.7812
66 | 16421 | 769.7344
66 | 16413 | 769.3594
66 | 16422 | 769.7812
66 | 16415 | 769.4531
66 | 16420 | 769.6875
66 | 16422 | 769.7812
66 | 16414 | 769.4062
66 | 16426 | 769.9688
66 | 16417 | 769.5469
66 | 16422 | 769.7812
66 | 16425 | 769.9219
66 | 16416 | 769.5
66 | 16425 | 769.9219
66 | 16418 | 769.5938
66 | 16417 | 769.5469
66 | 16424 | 769.875
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16421 | 769.7344
66 | 16413 | 769.3594
66 | 16423 | 769.8281
66 | 16415 | 769.4531
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16416 | 769.5
66 | 16425 | 769.9219
66 | 16416 | 769.5
66 | 16422 | 769.7812
66 | 16423 | 769.8281
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16419 | 769.6406
66 | 16416 | 769.5
66 | 16424 | 769.875
66 | 16413 | 769.3594
66 | 16420 | 769.6875
66 | 16422 | 769.7812
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16416 | 769.5
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16416 | 769.5
66 | 16424 | 769.875
66 | 16415 | 769.4531
66 | 16422 | 769.7812
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16422 | 769.7812
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16414 | 769.4062
66 | 16420 | 769.6875
66 | 16423 | 769.8281
66 | 16415 | 769.4531
66 | 16425 | 769.9219
66 | 16421 | 769.7344
66 | 16418 | 769.5938
66 | 16423 | 769.8281
66 | 16415 | 769.4531
66 | 16419 | 769.6406
66 | 16423 | 769.8281
66 | 16412 | 769.3125
66 | 16418 | 769.5938
66 | 16420 | 769.6875
66 | 16414 | 769.4062
66 | 16422 | 769.7812
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16422 | 769.7812
66 | 16416 | 769.5
66 | 16423 | 769.8281
66 | 16415 | 769.4531
66 | 16421 | 769.7344
66 | 16422 | 769.7812
66 | 16414 | 769.4062
66 | 16421 | 769.7344
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16421 | 769.7344
66 | 16412 | 769.3125
66 | 16419 | 769.6406
66 | 16422 | 769.7812
66 | 16412 | 769.3125
66 | 16421 | 769.7344
66 | 16419 | 769.6406
66 | 16419 | 769.6406
66 | 16424 | 769.875
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16418 | 769.5938
66 | 16416 | 769.5
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16420 | 769.6875
66 | 16417 | 769.5469
66 | 16415 | 769.4531
66 | 16421 | 769.7344
66 | 16413 | 769.3594
66 | 16420 | 769.6875
66 | 16418 | 769.5938
66 | 16415 | 769.4531
66 | 16423 | 769.8281
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16422 | 769.7812
66 | 16414 | 769.4062
66 | 16418 | 769.5938
66 | 16420 | 769.6875
66 | 16414 | 769.4062
66 | 16420 | 769.6875
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16413 | 769.3594
66 | 16419 | 769.6406
66 | 16422 | 769.7812
66 | 16414 | 769.4062
66 | 16422 | 769.7812
66 | 16420 | 769.6875
66 | 16416 | 769.5
66 | 16424 | 769.875
66 | 16412 | 769.3125
66 | 16419 | 769.6406
66 | 16418 | 769.5938
66 | 16413 | 769.3594
66 | 16422 | 769.7812
66 | 16415 | 769.4531
66 | 16420 | 769.6875
66 | 16420 | 769.6875
66 | 16414 | 769.4062
66 | 16422 | 769.7812
66 | 16418 | 769.5938
66 | 16416 | 769.5
66 | 16424 | 769.875
66 | 16414 | 769.4062
66 | 16420 | 769.6875
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16420 | 769.6875
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16425 | 769.9219
66 | 16415 | 769.4531
66 | 16419 | 769.6406
66 | 16423 | 769.8281
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16418 | 769.5938
66 | 16413 | 769.3594
66 | 16420 | 769.6875
66 | 16411 | 769.2656
66 | 16419 | 769.6406
66 | 16421 | 769.7344
66 | 16411 | 769.2656
66 | 16419 | 769.6406
66 | 16420 | 769.6875
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16415 | 769.4531
66 | 16417 | 769.5469
66 | 16421 | 769.7344
66 | 16412 | 769.3125
66 | 16419 | 769.6406
66 | 16419 | 769.6406
66 | 16411 | 769.2656
66 | 16419 | 769.6406
66 | 16414 | 769.4062
66 | 16414 | 769.4062
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16420 | 769.6875
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16421 | 769.7344
66 | 16414 | 769.4062
66 | 16423 | 769.8281
66 | 16415 | 769.4531
66 | 16415 | 769.4531
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16419 | 769.6406
66 | 16411 | 769.2656
66 | 16419 | 769.6406
66 | 16418 | 769.5938
66 | 16414 | 769.4062
66 | 16421 | 769.7344
66 | 16415 | 769.4531
66 | 16418 | 769.5938
66 | 16423 | 769.8281
66 | 16412 | 769.3125
66 | 16421 | 769.7344
66 | 16421 | 769.7344
66 | 16410 | 769.2188
66 | 16420 | 769.6875
66 | 16417 | 769.5469
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16411 | 769.2656
66 | 16418 | 769.5938
66 | 16421 | 769.7344
66 | 16412 | 769.3125
66 | 16421 | 769.7344
66 | 16418 | 769.5938
66 | 16415 | 769.4531
66 | 16422 | 769.7812
66 | 16412 | 769.3125
66 | 16418 | 769.5938
66 | 16421 | 769.7344
66 | 16411 | 769.2656
66 | 16418 | 769.5938
66 | 16418 | 769.5938
66 | 16413 | 769.3594
66 | 16420 | 769.6875
66 | 16412 | 769.3125
66 | 16416 | 769.5
66 | 16420 | 769.6875
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16418 | 769.5938
66 | 16415 | 769.4531
66 | 16422 | 769.7812
66 | 16411 | 769.2656
66 | 16417 | 769.5469
66 | 16419 | 769.6406
66 | 16410 | 769.2188
66 | 16420 | 769.6875
66 | 16416 | 769.5
66 | 16411 | 769.2656
66 | 16420 | 769.6875
66 | 16412 | 769.3125
66 | 16414 | 769.4062
66 | 16421 | 769.7344
66 | 16412 | 769.3125
66 | 16418 | 769.5938
66 | 16421 | 769.7344
66 | 16411 | 769.2656
66 | 16420 | 769.6875
66 | 16416 | 769.5
66 | 16412 | 769.3125
66 | 16419 | 769.6406
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16417 | 769.5469
66 | 16410 | 769.2188
66 | 16419 | 769.6406
66 | 16411 | 769.2656
66 | 16415 | 769.4531
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16421 | 769.7344
66 | 16411 | 769.2656
66 | 16419 | 769.6406
66 | 16419 | 769.6406
66 | 16411 | 769.2656
66 | 16419 | 769.6406
66 | 16416 | 769.5
66 | 16410 | 769.2188
66 | 16419 | 769.6406
66 | 16412 | 769.3125
66 | 16411 | 769.2656
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16421 | 769.7344
66 | 16410 | 769.2188
66 | 16419 | 769.6406
66 | 16420 | 769.6875
66 | 16412 | 769.3125
66 | 16419 | 769.6406
66 | 16416 | 769.5
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16412 | 769.3125
66 | 16412 | 769.3125
66 | 16419 | 769.6406
66 | 16408 | 769.125
66 | 16417 | 769.5469
66 | 16419 | 769.6406
66 | 16410 | 769.2188
66 | 16419 | 769.6406
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16421 | 769.7344
66 | 16412 | 769.3125
66 | 16415 | 769.4531
66 | 16420 | 769.6875
66 | 16408 | 769.125
66 | 16416 | 769.5
66 | 16418 | 769.5938
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16416 | 769.5
66 | 16409 | 769.1719
66 | 16419 | 769.6406
66 | 16417 | 769.5469
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16418 | 769.5938
66 | 16413 | 769.3594
66 | 16419 | 769.6406
66 | 16415 | 769.4531
66 | 16414 | 769.4062
66 | 16419 | 769.6406
66 | 16408 | 769.125
66 | 16414 | 769.4062
66 | 16416 | 769.5
66 | 16408 | 769.125
66 | 16416 | 769.5
66 | 16417 | 769.5469
66 | 16410 | 769.2188
66 | 16419 | 769.6406
66 | 16417 | 769.5469
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16415 | 769.4531
66 | 16413 | 769.3594
66 | 16421 | 769.7344
66 | 16412 | 769.3125
66 | 16412 | 769.3125
66 | 16417 | 769.5469
66 | 16407 | 769.0781
66 | 16413 | 769.3594
66 | 16417 | 769.5469
66 | 16409 | 769.1719
66 | 16416 | 769.5
66 | 16417 | 769.5469
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16418 | 769.5938
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16417 | 769.5469
66 | 16409 | 769.1719
66 | 16415 | 769.4531
66 | 16417 | 769.5469
66 | 16408 | 769.125
66 | 16416 | 769.5
66 | 16416 | 769.5
66 | 16407 | 769.0781
66 | 16418 | 769.5938
66 | 16418 | 769.5938
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16418 | 769.5938
66 | 16411 | 769.2656
66 | 16417 | 769.5469
66 | 16421 | 769.7344
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16418 | 769.5938
66 | 16409 | 769.1719
66 | 16416 | 769.5
66 | 16418 | 769.5938
66 | 16407 | 769.0781
66 | 16416 | 769.5
66 | 16418 | 769.5938
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16418 | 769.5938
66 | 16409 | 769.1719
66 | 16418 | 769.5938
66 | 16418 | 769.5938
66 | 16411 | 769.2656
66 | 16418 | 769.5938
66 | 16417 | 769.5469
66 | 16409 | 769.1719
66 | 16416 | 769.5
66 | 16415 | 769.4531
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16411 | 769.2656
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16409 | 769.1719
66 | 16415 | 769.4531
66 | 16419 | 769.6406
66 | 16411 | 769.2656
66 | 16417 | 769.5469
66 | 16418 | 769.5938
66 | 16410 | 769.2188
66 | 16418 | 769.5938
66 | 16417 | 769.5469
66 | 16409 | 769.1719
66 | 16414 | 769.4062
66 | 16416 | 769.5
66 | 16407 | 769.0781
66 | 16415 | 769.4531
66 | 16413 | 769.3594
66 | 16408 | 769.125
66 | 16417 | 769.5469
66 | 16413 | 769.3594
66 | 16412 | 769.3125
66 | 16420 | 769.6875
66 | 16410 | 769.2188
66 | 16415 | 769.4531
66 | 16419 | 769.6406
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16416 | 769.5
66 | 16408 | 769.125
66 | 16416 | 769.5
66 | 16415 | 769.4531
66 | 16408 | 769.125
66 | 16417 | 769.5469
66 | 16413 | 769.3594
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16411 | 769.2656
66 | 16413 | 769.3594
66 | 16418 | 769.5938
66 | 16408 | 769.125
66 | 16418 | 769.5938
66 | 16418 | 769.5938
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16416 | 769.5
66 | 16405 | 768.9844
66 | 16412 | 769.3125
66 | 16416 | 769.5
66 | 16407 | 769.0781
66 | 16415 | 769.4531
66 | 16415 | 769.4531
66 | 16408 | 769.125
66 | 16417 | 769.5469
66 | 16415 | 769.4531
66 | 16410 | 769.2188
66 | 16416 | 769.5
66 | 16413 | 769.3594
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16408 | 769.125
66 | 16415 | 769.4531
66 | 16408 | 769.125
66 | 16410 | 769.2188
66 | 16415 | 769.4531
66 | 16406 | 769.0312
66 | 16415 | 769.4531
66 | 16416 | 769.5
66 | 16407 | 769.0781
66 | 16414 | 769.4062
66 | 16415 | 769.4531
66 | 16409 | 769.1719
66 | 16416 | 769.5
66 | 16417 | 769.5469
66 | 16407 | 769.0781
66 | 16415 | 769.4531
66 | 16413 | 769.3594
66 | 16408 | 769.125
66 | 16414 | 769.4062
66 | 16412 | 769.3125
66 | 16407 | 769.0781
66 | 16415 | 769.4531
66 | 16413 | 769.3594
66 | 16407 | 769.0781
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16411 | 769.2656
66 | 16409 | 769.1719
66 | 16416 | 769.5
66 | 16406 | 769.0312
66 | 16412 | 769.3125
66 | 16415 | 769.4531
66 | 16406 | 769.0312
66 | 16415 | 769.4531
66 | 16416 | 769.5
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16415 | 769.4531
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16413 | 769.3594
66 | 16408 | 769.125
66 | 16415 | 769.4531
66 | 16408 | 769.125
66 | 16410 | 769.2188
66 | 16416 | 769.5
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16407 | 769.0781
66 | 16416 | 769.5
66 | 16415 | 769.4531
66 | 16408 | 769.125
66 | 16417 | 769.5469
66 | 16414 | 769.4062
66 | 16410 | 769.2188
66 | 16417 | 769.5469
66 | 16411 | 769.2656
66 | 16409 | 769.1719
66 | 16416 | 769.5
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16404 | 768.9375
66 | 16414 | 769.4062
66 | 16415 | 769.4531
66 | 16407 | 769.0781
66 | 16416 | 769.5
66 | 16415 | 769.4531
66 | 16407 | 769.0781
66 | 16416 | 769.5
66 | 16415 | 769.4531
66 | 16408 | 769.125
66 | 16416 | 769.5
66 | 16413 | 769.3594
66 | 16407 | 769.0781
66 | 16414 | 769.4062
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16412 | 769.3125
66 | 16407 | 769.0781
66 | 16415 | 769.4531
66 | 16413 | 769.3594
66 | 16406 | 769.0312
66 | 16415 | 769.4531
66 | 16414 | 769.4062
66 | 16408 | 769.125
66 | 16417 | 769.5469
66 | 16413 | 769.3594
66 | 16409 | 769.1719
66 | 16414 | 769.4062
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16410 | 769.2188
66 | 16406 | 769.0312
66 | 16412 | 769.3125
66 | 16411 | 769.2656
66 | 16406 | 769.0312
66 | 16415 | 769.4531
66 | 16412 | 769.3125
66 | 16409 | 769.1719
66 | 16417 | 769.5469
66 | 16410 | 769.2188
66 | 16410 | 769.2188
66 | 16416 | 769.5
66 | 16405 | 768.9844
66 | 16410 | 769.2188
66 | 16414 | 769.4062
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16412 | 769.3125
66 | 16413 | 769.3594
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16413 | 769.3594
66 | 16406 | 769.0312
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16406 | 769.0312
66 | 16414 | 769.4062
66 | 16414 | 769.4062
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16413 | 769.3594
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16413 | 769.3594
66 | 16405 | 768.9844
66 | 16414 | 769.4062
66 | 16413 | 769.3594
66 | 16406 | 769.0312
66 | 16414 | 769.4062
66 | 16410 | 769.2188
66 | 16406 | 769.0312
66 | 16413 | 769.3594
66 | 16410 | 769.2188
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16409 | 769.1719
66 | 16405 | 768.9844
66 | 16412 | 769.3125
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16415 | 769.4531
66 | 16414 | 769.4062
66 | 16409 | 769.1719
66 | 16415 | 769.4531
66 | 16412 | 769.3125
66 | 16406 | 769.0312
66 | 16413 | 769.3594
66 | 16413 | 769.3594
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16412 | 769.3125
66 | 16415 | 769.4531
66 | 16404 | 768.9375
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16406 | 769.0312
66 | 16411 | 769.2656
66 | 16414 | 769.4062
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16413 | 769.3594
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16413 | 769.3594
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16414 | 769.4062
66 | 16406 | 769.0312
66 | 16413 | 769.3594
66 | 16413 | 769.3594
66 | 16406 | 769.0312
66 | 16412 | 769.3125
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16411 | 769.2656
66 | 16410 | 769.2188
66 | 16404 | 768.9375
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16404 | 768.9375
66 | 16413 | 769.3594
66 | 16413 | 769.3594
66 | 16407 | 769.0781
66 | 16415 | 769.4531
66 | 16414 | 769.4062
66 | 16406 | 769.0312
66 | 16413 | 769.3594
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16412 | 769.3125
66 | 16408 | 769.125
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16404 | 768.9375
66 | 16407 | 769.0781
66 | 16413 | 769.3594
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16414 | 769.4062
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16414 | 769.4062
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16410 | 769.2188
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16404 | 768.9375
66 | 16412 | 769.3125
66 | 16413 | 769.3594
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16414 | 769.4062
66 | 16404 | 768.9375
66 | 16409 | 769.1719
66 | 16413 | 769.3594
66 | 16400 | 768.75
66 | 16409 | 769.1719
66 | 16412 | 769.3125
66 | 16402 | 768.8438
66 | 16409 | 769.1719
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16413 | 769.3594
66 | 16412 | 769.3125
66 | 16404 | 768.9375
66 | 16413 | 769.3594
66 | 16411 | 769.2656
66 | 16404 | 768.9375
66 | 16411 | 769.2656
66 | 16408 | 769.125
66 | 16404 | 768.9375
66 | 16411 | 769.2656
66 | 16408 | 769.125
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16411 | 769.2656
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16414 | 769.4062
66 | 16404 | 768.9375
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16410 | 769.2188
66 | 16408 | 769.125
66 | 16402 | 768.8438
66 | 16409 | 769.1719
66 | 16411 | 769.2656
66 | 16402 | 768.8438
66 | 16410 | 769.2188
66 | 16412 | 769.3125
66 | 16402 | 768.8438
66 | 16412 | 769.3125
66 | 16413 | 769.3594
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16401 | 768.7969
66 | 16410 | 769.2188
66 | 16409 | 769.1719
66 | 16401 | 768.7969
66 | 16408 | 769.125
66 | 16408 | 769.125
66 | 16401 | 768.7969
66 | 16410 | 769.2188
66 | 16409 | 769.1719
66 | 16402 | 768.8438
66 | 16411 | 769.2656
66 | 16411 | 769.2656
66 | 16403 | 768.8906
66 | 16411 | 769.2656
66 | 16412 | 769.3125
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16410 | 769.2188
66 | 16410 | 769.2188
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16407 | 769.0781
66 | 16412 | 769.3125
66 | 16404 | 768.9375
66 | 16406 | 769.0312
66 | 16413 | 769.3594
66 | 16403 | 768.8906
66 | 16405 | 768.9844
66 | 16411 | 769.2656
66 | 16407 | 769.0781
66 | 16402 | 768.8438
66 | 16408 | 769.125
66 | 16409 | 769.1719
66 | 16399 | 768.7031
66 | 16408 | 769.125
66 | 16409 | 769.1719
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16406 | 769.0312
66 | 16412 | 769.3125
66 | 16405 | 768.9844
66 | 16406 | 769.0312
66 | 16411 | 769.2656
66 | 16405 | 768.9844
66 | 16402 | 768.8438
66 | 16409 | 769.1719
66 | 16407 | 769.0781
66 | 16399 | 768.7031
66 | 16406 | 769.0312
66 | 16409 | 769.1719
66 | 16399 | 768.7031
66 | 16408 | 769.125
66 | 16410 | 769.2188
66 | 16400 | 768.75
66 | 16409 | 769.1719
66 | 16411 | 769.2656
66 | 16401 | 768.7969
66 | 16409 | 769.1719
66 | 16411 | 769.2656
66 | 16401 | 768.7969
66 | 16409 | 769.1719
66 | 16409 | 769.1719
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16406 | 769.0312
66 | 16399 | 768.7031
66 | 16406 | 769.0312
66 | 16407 | 769.0781
66 | 16399 | 768.7031
66 | 16408 | 769.125
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16409 | 769.1719
66 | 16410 | 769.2188
66 | 16402 | 768.8438
66 | 16408 | 769.125
66 | 16409 | 769.1719
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16404 | 768.9375
66 | 16411 | 769.2656
66 | 16408 | 769.125
66 | 16403 | 768.8906
66 | 16408 | 769.125
66 | 16410 | 769.2188
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16406 | 769.0312
66 | 16399 | 768.7031
66 | 16407 | 769.0781
66 | 16407 | 769.0781
66 | 16399 | 768.7031
66 | 16406 | 769.0312
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16405 | 768.9844
66 | 16410 | 769.2188
66 | 16404 | 768.9375
66 | 16400 | 768.75
66 | 16408 | 769.125
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16408 | 769.125
66 | 16399 | 768.7031
66 | 16407 | 769.0781
66 | 16410 | 769.2188
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16409 | 769.1719
66 | 16399 | 768.7031
66 | 16405 | 768.9844
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16402 | 768.8438
66 | 16407 | 769.0781
66 | 16397 | 768.6094
66 | 16404 | 768.9375
66 | 16408 | 769.125
66 | 16400 | 768.75
66 | 16403 | 768.8906
66 | 16410 | 769.2188
66 | 16401 | 768.7969
66 | 16405 | 768.9844
66 | 16410 | 769.2188
66 | 16403 | 768.8906
66 | 16402 | 768.8438
66 | 16407 | 769.0781
66 | 16406 | 769.0312
66 | 16397 | 768.6094
66 | 16405 | 768.9844
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16404 | 768.9375
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16408 | 769.125
66 | 16399 | 768.7031
66 | 16405 | 768.9844
66 | 16410 | 769.2188
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16408 | 769.125
66 | 16400 | 768.75
66 | 16405 | 768.9844
66 | 16409 | 769.1719
66 | 16397 | 768.6094
66 | 16403 | 768.8906
66 | 16405 | 768.9844
66 | 16396 | 768.5625
66 | 16402 | 768.8438
66 | 16406 | 769.0312
66 | 16397 | 768.6094
66 | 16405 | 768.9844
66 | 16409 | 769.1719
66 | 16399 | 768.7031
66 | 16406 | 769.0312
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16406 | 769.0312
66 | 16406 | 769.0312
66 | 16397 | 768.6094
66 | 16406 | 769.0312
66 | 16406 | 769.0312
66 | 16396 | 768.5625
66 | 16404 | 768.9375
66 | 16404 | 768.9375
66 | 16398 | 768.6562
66 | 16404 | 768.9375
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16404 | 768.9375
66 | 16409 | 769.1719
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16410 | 769.2188
66 | 16400 | 768.75
66 | 16404 | 768.9375
66 | 16408 | 769.125
66 | 16396 | 768.5625
66 | 16402 | 768.8438
66 | 16406 | 769.0312
66 | 16397 | 768.6094
66 | 16403 | 768.8906
66 | 16406 | 769.0312
66 | 16396 | 768.5625
66 | 16404 | 768.9375
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16409 | 769.1719
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16407 | 769.0781
66 | 16397 | 768.6094
66 | 16403 | 768.8906
66 | 16407 | 769.0781
66 | 16395 | 768.5156
66 | 16402 | 768.8438
66 | 16406 | 769.0312
66 | 16398 | 768.6562
66 | 16400 | 768.75
66 | 16406 | 769.0312
66 | 16404 | 768.9375
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16407 | 769.0781
66 | 16405 | 768.9844
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16406 | 769.0312
66 | 16396 | 768.5625
66 | 16404 | 768.9375
66 | 16405 | 768.9844
66 | 16394 | 768.4688
66 | 16402 | 768.8438
66 | 16406 | 769.0312
66 | 16395 | 768.5156
66 | 16402 | 768.8438
66 | 16408 | 769.125
66 | 16398 | 768.6562
66 | 16402 | 768.8438
66 | 16407 | 769.0781
66 | 16403 | 768.8906
66 | 16399 | 768.7031
66 | 16406 | 769.0312
66 | 16405 | 768.9844
66 | 16397 | 768.6094
66 | 16402 | 768.8438
66 | 16404 | 768.9375
66 | 16393 | 768.4219
66 | 16401 | 768.7969
66 | 16406 | 769.0312
66 | 16398 | 768.6562
66 | 16401 | 768.7969
66 | 16407 | 769.0781
66 | 16404 | 768.9375
66 | 16398 | 768.6562
66 | 16407 | 769.0781
66 | 16407 | 769.0781
66 | 16396 | 768.5625
66 | 16403 | 768.8906
66 | 16407 | 769.0781
66 | 16401 | 768.7969
66 | 16397 | 768.6094
66 | 16402 | 768.8438
66 | 16404 | 768.9375
66 | 16394 | 768.4688
66 | 16402 | 768.8438
66 | 16405 | 768.9844
66 | 16397 | 768.6094
66 | 16400 | 768.75
66 | 16406 | 769.0312
66 | 16401 | 768.7969
66 | 16399 | 768.7031
66 | 16405 | 768.9844
66 | 16407 | 769.0781
66 | 16397 | 768.6094
66 | 16399 | 768.7031
66 | 16406 | 769.0312
66 | 16403 | 768.8906
66 | 16395 | 768.5156
66 | 16401 | 768.7969
66 | 16404 | 768.9375
66 | 16396 | 768.5625
66 | 16400 | 768.75
66 | 16405 | 768.9844
66 | 16402 | 768.8438
66 | 16397 | 768.6094
66 | 16403 | 768.8906
66 | 16407 | 769.0781
66 | 16398 | 768.6562
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16403 | 768.8906
66 | 16397 | 768.6094
66 | 16404 | 768.9375
66 | 16405 | 768.9844
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16405 | 768.9844
66 | 16401 | 768.7969
66 | 16395 | 768.5156
66 | 16403 | 768.8906
66 | 16407 | 769.0781
66 | 16396 | 768.5625
66 | 16402 | 768.8438
66 | 16406 | 769.0312
66 | 16401 | 768.7969
66 | 16399 | 768.7031
66 | 16404 | 768.9375
66 | 16406 | 769.0312
66 | 16396 | 768.5625
66 | 16400 | 768.75
66 | 16405 | 768.9844
66 | 16399 | 768.7031
66 | 16394 | 768.4688
66 | 16402 | 768.8438
66 | 16404 | 768.9375
66 | 16395 | 768.5156
66 | 16400 | 768.75
66 | 16406 | 769.0312
66 | 16402 | 768.8438
66 | 16397 | 768.6094
66 | 16403 | 768.8906
66 | 16406 | 769.0312
66 | 16399 | 768.7031
66 | 16398 | 768.6562
66 | 16404 | 768.9375
66 | 16402 | 768.8438
66 | 16395 | 768.5156
66 | 16400 | 768.75
66 | 16405 | 768.9844
66 | 16398 | 768.6562
66 | 16395 | 768.5156
66 | 16403 | 768.8906
66 | 16404 | 768.9375
66 | 16395 | 768.5156
66 | 16402 | 768.8438
66 | 16405 | 768.9844
66 | 16397 | 768.6094
66 | 16401 | 768.7969
66 | 16406 | 769.0312
66 | 16399 | 768.7031
66 | 16396 | 768.5625
66 | 16401 | 768.7969
66 | 16404 | 768.9375
66 | 16392 | 768.375
66 | 16399 | 768.7031
66 | 16402 | 768.8438
66 | 16397 | 768.6094
66 | 16393 | 768.4219
66 | 16402 | 768.8438
66 | 16403 | 768.8906
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16404 | 768.9375
66 | 16395 | 768.5156
66 | 16399 | 768.7031
66 | 16405 | 768.9844
66 | 16401 | 768.7969
66 | 16394 | 768.4688
66 | 16402 | 768.8438
66 | 16402 | 768.8438
66 | 16392 | 768.375
66 | 16398 | 768.6562
66 | 16403 | 768.8906
66 | 16399 | 768.7031
66 | 16394 | 768.4688
66 | 16402 | 768.8438
66 | 16404 | 768.9375
66 | 16395 | 768.5156
66 | 16397 | 768.6094
66 | 16404 | 768.9375
66 | 16404 | 768.9375
66 | 16395 | 768.5156
66 | 16398 | 768.6562
66 | 16405 | 768.9844
66 | 16401 | 768.7969
66 | 16395 | 768.5156
66 | 16401 | 768.7969
66 | 16402 | 768.8438
66 | 16394 | 768.4688
66 | 16394 | 768.4688
66 | 16401 | 768.7969
66 | 16404 | 768.9375
66 | 16396 | 768.5625
66 | 16398 | 768.6562
66 | 16403 | 768.8906
66 | 16404 | 768.9375
66 | 16397 | 768.6094
66 | 16400 | 768.75
66 | 16407 | 769.0781
66 | 16399 | 768.7031
66 | 16394 | 768.4688
66 | 16403 | 768.8906
66 | 16402 | 768.8438
66 | 16393 | 768.4219
66 | 16394 | 768.4688
66 | 16401 | 768.7969
66 | 16402 | 768.8438
66 | 16393 | 768.4219
66 | 16396 | 768.5625
66 | 16404 | 768.9375
66 | 16401 | 768.7969
66 | 16395 | 768.5156
66 | 16402 | 768.8438
66 | 16405 | 768.9844
66 | 16397 | 768.6094
66 | 16397 | 768.6094
66 | 16403 | 768.8906
66 | 16402 | 768.8438
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16401 | 768.7969
66 | 16395 | 768.5156
66 | 16393 | 768.4219
66 | 16400 | 768.75
66 | 16403 | 768.8906
66 | 16393 | 768.4219
66 | 16395 | 768.5156
66 | 16401 | 768.7969
66 | 16404 | 768.9375
66 | 16396 | 768.5625
66 | 16397 | 768.6094
66 | 16402 | 768.8438
66 | 16402 | 768.8438
66 | 16393 | 768.4219
66 | 16396 | 768.5625
66 | 16402 | 768.8438
66 | 16397 | 768.6094
66 | 16391 | 768.3281
66 | 16398 | 768.6562
66 | 16401 | 768.7969
66 | 16397 | 768.6094
66 | 16393 | 768.4219
66 | 16400 | 768.75
66 | 16404 | 768.9375
66 | 16398 | 768.6562
66 | 16394 | 768.4688
66 | 16401 | 768.7969
66 | 16404 | 768.9375
66 | 16395 | 768.5156
66 | 16395 | 768.5156
66 | 16399 | 768.7031
66 | 16402 | 768.8438
66 | 16396 | 768.5625
66 | 16392 | 768.375
66 | 16398 | 768.6562
66 | 16402 | 768.8438
66 | 16395 | 768.5156
66 | 16391 | 768.3281
66 | 16398 | 768.6562
66 | 16403 | 768.8906
66 | 16398 | 768.6562
66 | 16394 | 768.4688
66 | 16401 | 768.7969
66 | 16403 | 768.8906
66 | 16397 | 768.6094
66 | 16394 | 768.4688
66 | 16400 | 768.75
66 | 16402 | 768.8438
66 | 16395 | 768.5156
66 | 16393 | 768.4219
66 | 16398 | 768.6562
66 | 16403 | 768.8906
66 | 16397 | 768.6094
66 | 16393 | 768.4219
66 | 16399 | 768.7031
66 | 16403 | 768.8906
66 | 16399 | 768.7031
66 | 16392 | 768.375
66 | 16399 | 768.7031
66 | 16404 | 768.9375
66 | 16398 | 768.6562
66 | 16394 | 768.4688
66 | 16400 | 768.75
66 | 16402 | 768.8438
66 | 16392 | 768.375
66 | 16394 | 768.4688
66 | 16398 | 768.6562
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16400 | 768.75
66 | 16400 | 768.75
66 | 16391 | 768.3281
66 | 16396 | 768.5625
66 | 16404 | 768.9375
66 | 16400 | 768.75
66 | 16392 | 768.375
66 | 16396 | 768.5625
66 | 16402 | 768.8438
66 | 16398 | 768.6562
66 | 16392 | 768.375
66 | 16398 | 768.6562
66 | 16401 | 768.7969
66 | 16394 | 768.4688
66 | 16393 | 768.4219
66 | 16399 | 768.7031
66 | 16400 | 768.75
66 | 16393 | 768.4219
66 | 16394 | 768.4688
66 | 16400 | 768.75
66 | 16402 | 768.8438
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16402 | 768.8438
66 | 16400 | 768.75
66 | 16391 | 768.3281
66 | 16397 | 768.6094
66 | 16401 | 768.7969
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16398 | 768.6562
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16401 | 768.7969
66 | 16400 | 768.75
66 | 16392 | 768.375
66 | 16396 | 768.5625
66 | 16403 | 768.8906
66 | 16401 | 768.7969
66 | 16392 | 768.375
66 | 16394 | 768.4688
66 | 16400 | 768.75
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16400 | 768.75
66 | 16400 | 768.75
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16401 | 768.7969
66 | 16392 | 768.375
66 | 16397 | 768.6094
66 | 16402 | 768.8438
66 | 16399 | 768.7031
66 | 16391 | 768.3281
66 | 16396 | 768.5625
66 | 16400 | 768.75
66 | 16398 | 768.6562
66 | 16389 | 768.2344
66 | 16395 | 768.5156
66 | 16399 | 768.7031
66 | 16396 | 768.5625
66 | 16388 | 768.1875
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16399 | 768.7031
66 | 16392 | 768.375
66 | 16396 | 768.5625
66 | 16400 | 768.75
66 | 16400 | 768.75
66 | 16390 | 768.2812
66 | 16396 | 768.5625
66 | 16399 | 768.7031
66 | 16400 | 768.75
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16400 | 768.75
66 | 16397 | 768.6094
66 | 16388 | 768.1875
66 | 16392 | 768.375
66 | 16398 | 768.6562
66 | 16398 | 768.6562
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16398 | 768.6562
66 | 16400 | 768.75
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16401 | 768.7969
66 | 16400 | 768.75
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16399 | 768.7031
66 | 16397 | 768.6094
66 | 16389 | 768.2344
66 | 16393 | 768.4219
66 | 16399 | 768.7031
66 | 16396 | 768.5625
66 | 16387 | 768.1406
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16398 | 768.6562
66 | 16391 | 768.3281
66 | 16395 | 768.5156
66 | 16401 | 768.7969
66 | 16397 | 768.6094
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16398 | 768.6562
66 | 16395 | 768.5156
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16400 | 768.75
66 | 16401 | 768.7969
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16398 | 768.6562
66 | 16401 | 768.7969
66 | 16397 | 768.6094
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16398 | 768.6562
66 | 16399 | 768.7031
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16396 | 768.5625
66 | 16398 | 768.6562
66 | 16396 | 768.5625
66 | 16388 | 768.1875
66 | 16392 | 768.375
66 | 16399 | 768.7031
66 | 16401 | 768.7969
66 | 16393 | 768.4219
66 | 16391 | 768.3281
66 | 16398 | 768.6562
66 | 16401 | 768.7969
66 | 16401 | 768.7969
66 | 16391 | 768.3281
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16398 | 768.6562
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16395 | 768.5156
66 | 16399 | 768.7031
66 | 16397 | 768.6094
66 | 16389 | 768.2344
66 | 16392 | 768.375
66 | 16399 | 768.7031
66 | 16401 | 768.7969
66 | 16395 | 768.5156
66 | 16391 | 768.3281
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16400 | 768.75
66 | 16394 | 768.4688
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16396 | 768.5625
66 | 16398 | 768.6562
66 | 16393 | 768.4219
66 | 16389 | 768.2344
66 | 16395 | 768.5156
66 | 16398 | 768.6562
66 | 16398 | 768.6562
66 | 16393 | 768.4219
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16400 | 768.75
66 | 16396 | 768.5625
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16395 | 768.5156
66 | 16398 | 768.6562
66 | 16397 | 768.6094
66 | 16389 | 768.2344
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16399 | 768.7031
66 | 16398 | 768.6562
66 | 16392 | 768.375
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16399 | 768.7031
66 | 16401 | 768.7969
66 | 16394 | 768.4688
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16397 | 768.6094
66 | 16391 | 768.3281
66 | 16388 | 768.1875
66 | 16392 | 768.375
66 | 16398 | 768.6562
66 | 16396 | 768.5625
66 | 16388 | 768.1875
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16399 | 768.7031
66 | 16397 | 768.6094
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16398 | 768.6562
66 | 16399 | 768.7031
66 | 16392 | 768.375
66 | 16388 | 768.1875
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16394 | 768.4688
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16395 | 768.5156
66 | 16397 | 768.6094
66 | 16395 | 768.5156
66 | 16389 | 768.2344
66 | 16392 | 768.375
66 | 16397 | 768.6094
66 | 16398 | 768.6562
66 | 16393 | 768.4219
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16396 | 768.5625
66 | 16397 | 768.6094
66 | 16388 | 768.1875
66 | 16388 | 768.1875
66 | 16394 | 768.4688
66 | 16396 | 768.5625
66 | 16393 | 768.4219
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16397 | 768.6094
66 | 16396 | 768.5625
66 | 16391 | 768.3281
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16398 | 768.6562
66 | 16393 | 768.4219
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16395 | 768.5156
66 | 16388 | 768.1875
66 | 16387 | 768.1406
66 | 16393 | 768.4219
66 | 16396 | 768.5625
66 | 16395 | 768.5156
66 | 16386 | 768.0938
66 | 16391 | 768.3281
66 | 16396 | 768.5625
66 | 16398 | 768.6562
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16398 | 768.6562
66 | 16395 | 768.5156
66 | 16388 | 768.1875
66 | 16389 | 768.2344
66 | 16392 | 768.375
66 | 16396 | 768.5625
66 | 16395 | 768.5156
66 | 16386 | 768.0938
66 | 16387 | 768.1406
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16395 | 768.5156
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16395 | 768.5156
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16396 | 768.5625
66 | 16398 | 768.6562
66 | 16393 | 768.4219
66 | 16384 | 768
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16394 | 768.4688
66 | 16389 | 768.2344
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16397 | 768.6094
66 | 16391 | 768.3281
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16398 | 768.6562
66 | 16391 | 768.3281
66 | 16387 | 768.1406
66 | 16393 | 768.4219
66 | 16395 | 768.5156
66 | 16397 | 768.6094
66 | 16391 | 768.3281
66 | 16385 | 768.0469
66 | 16390 | 768.2812
66 | 16392 | 768.375
66 | 16396 | 768.5625
66 | 16396 | 768.5625
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16392 | 768.375
66 | 16395 | 768.5156
66 | 16398 | 768.6562
66 | 16396 | 768.5625
66 | 16388 | 768.1875
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16396 | 768.5625
66 | 16396 | 768.5625
66 | 16392 | 768.375
66 | 16385 | 768.0469
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16395 | 768.5156
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16396 | 768.5625
66 | 16397 | 768.6094
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16395 | 768.5156
66 | 16397 | 768.6094
66 | 16395 | 768.5156
66 | 16388 | 768.1875
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16396 | 768.5625
66 | 16395 | 768.5156
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16397 | 768.6094
66 | 16393 | 768.4219
66 | 16387 | 768.1406
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16396 | 768.5625
66 | 16391 | 768.3281
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16395 | 768.5156
66 | 16395 | 768.5156
66 | 16392 | 768.375
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16394 | 768.4688
66 | 16397 | 768.6094
66 | 16397 | 768.6094
66 | 16394 | 768.4688
66 | 16386 | 768.0938
66 | 16388 | 768.1875
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16394 | 768.4688
66 | 16395 | 768.5156
66 | 16392 | 768.375
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16392 | 768.375
66 | 16395 | 768.5156
66 | 16394 | 768.4688
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16396 | 768.5625
66 | 16396 | 768.5625
66 | 16392 | 768.375
66 | 16384 | 768
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16393 | 768.4219
66 | 16390 | 768.2812
66 | 16383 | 767.9531
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16395 | 768.5156
66 | 16394 | 768.4688
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16397 | 768.6094
66 | 16395 | 768.5156
66 | 16390 | 768.2812
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16389 | 768.2344
66 | 16392 | 768.375
66 | 16393 | 768.4219
66 | 16392 | 768.375
66 | 16386 | 768.0938
66 | 16383 | 767.9531
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16395 | 768.5156
66 | 16397 | 768.6094
66 | 16394 | 768.4688
66 | 16389 | 768.2344
66 | 16386 | 768.0938
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16392 | 768.375
66 | 16395 | 768.5156
66 | 16392 | 768.375
66 | 16387 | 768.1406
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16395 | 768.5156
66 | 16394 | 768.4688
66 | 16391 | 768.3281
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16391 | 768.3281
66 | 16394 | 768.4688
66 | 16396 | 768.5625
66 | 16396 | 768.5625
66 | 16389 | 768.2344
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16393 | 768.4219
66 | 16388 | 768.1875
66 | 16383 | 767.9531
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16393 | 768.4219
66 | 16395 | 768.5156
66 | 16391 | 768.3281
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16390 | 768.2812
66 | 16394 | 768.4688
66 | 16395 | 768.5156
66 | 16393 | 768.4219
66 | 16387 | 768.1406
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16392 | 768.375
66 | 16393 | 768.4219
66 | 16392 | 768.375
66 | 16386 | 768.0938
66 | 16382 | 767.9062
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16392 | 768.375
66 | 16392 | 768.375
66 | 16388 | 768.1875
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16391 | 768.3281
66 | 16395 | 768.5156
66 | 16392 | 768.375
66 | 16386 | 768.0938
66 | 16385 | 768.0469
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16382 | 767.9062
66 | 16386 | 768.0938
66 | 16392 | 768.375
66 | 16393 | 768.4219
66 | 16388 | 768.1875
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16391 | 768.3281
66 | 16395 | 768.5156
66 | 16394 | 768.4688
66 | 16388 | 768.1875
66 | 16382 | 767.9062
66 | 16386 | 768.0938
66 | 16392 | 768.375
66 | 16394 | 768.4688
66 | 16392 | 768.375
66 | 16387 | 768.1406
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16390 | 768.2812
66 | 16393 | 768.4219
66 | 16393 | 768.4219
66 | 16386 | 768.0938
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16394 | 768.4688
66 | 16391 | 768.3281
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16390 | 768.2812
66 | 16384 | 768
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16392 | 768.375
66 | 16388 | 768.1875
66 | 16382 | 767.9062
66 | 16386 | 768.0938
66 | 16388 | 768.1875
66 | 16392 | 768.375
66 | 16393 | 768.4219
66 | 16389 | 768.2344
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16393 | 768.4219
66 | 16394 | 768.4688
66 | 16389 | 768.2344
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16387 | 768.1406
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16391 | 768.3281
66 | 16386 | 768.0938
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16393 | 768.4219
66 | 16390 | 768.2812
66 | 16383 | 767.9531
66 | 16383 | 767.9531
66 | 16387 | 768.1406
66 | 16390 | 768.2812
66 | 16392 | 768.375
66 | 16391 | 768.3281
66 | 16386 | 768.0938
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16388 | 768.1875
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16390 | 768.2812
66 | 16391 | 768.3281
66 | 16393 | 768.4219
66 | 16391 | 768.3281
66 | 16385 | 768.0469
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16390 | 768.2812
66 | 16391 | 768.3281
66 | 16387 | 768.1406
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16392 | 768.375
66 | 16389 | 768.2344
66 | 16381 | 767.8594
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16391 | 768.3281
66 | 16391 | 768.3281
66 | 16390 | 768.2812
66 | 16382 | 767.9062
66 | 16380 | 767.8125
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16387 | 768.1406
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16391 | 768.3281
66 | 16392 | 768.375
66 | 16391 | 768.3281
66 | 16388 | 768.1875
66 | 16383 | 767.9531
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16390 | 768.2812
66 | 16387 | 768.1406
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16388 | 768.1875
66 | 16391 | 768.3281
66 | 16389 | 768.2344
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16391 | 768.3281
66 | 16391 | 768.3281
66 | 16391 | 768.3281
66 | 16390 | 768.2812
66 | 16384 | 768
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16387 | 768.1406
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16387 | 768.1406
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16385 | 768.0469
66 | 16388 | 768.1875
66 | 16390 | 768.2812
66 | 16392 | 768.375
66 | 16393 | 768.4219
66 | 16385 | 768.0469
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16390 | 768.2812
66 | 16388 | 768.1875
66 | 16384 | 768
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16388 | 768.1875
66 | 16390 | 768.2812
66 | 16392 | 768.375
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16390 | 768.2812
66 | 16387 | 768.1406
66 | 16382 | 767.9062
66 | 16378 | 767.7188
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16389 | 768.2344
66 | 16388 | 768.1875
66 | 16382 | 767.9062
66 | 16380 | 767.8125
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16390 | 768.2812
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16386 | 768.0938
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16387 | 768.1406
66 | 16388 | 768.1875
66 | 16387 | 768.1406
66 | 16386 | 768.0938
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16391 | 768.3281
66 | 16390 | 768.2812
66 | 16386 | 768.0938
66 | 16381 | 767.8594
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16386 | 768.0938
66 | 16389 | 768.2344
66 | 16389 | 768.2344
66 | 16385 | 768.0469
66 | 16382 | 767.9062
66 | 16379 | 767.7656
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16386 | 768.0938
66 | 16391 | 768.3281
66 | 16389 | 768.2344
66 | 16390 | 768.2812
66 | 16387 | 768.1406
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16385 | 768.0469
66 | 16388 | 768.1875
66 | 16388 | 768.1875
66 | 16387 | 768.1406
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16389 | 768.2344
66 | 16391 | 768.3281
66 | 16390 | 768.2812
66 | 16387 | 768.1406
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16383 | 767.9531
66 | 16381 | 767.8594
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16387 | 768.1406
66 | 16391 | 768.3281
66 | 16389 | 768.2344
66 | 16388 | 768.1875
66 | 16385 | 768.0469
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16388 | 768.1875
66 | 16387 | 768.1406
66 | 16385 | 768.0469
66 | 16381 | 767.8594
66 | 16378 | 767.7188
66 | 16378 | 767.7188
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16390 | 768.2812
66 | 16389 | 768.2344
66 | 16388 | 768.1875
66 | 16387 | 768.1406
66 | 16380 | 767.8125
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16386 | 768.0938
66 | 16387 | 768.1406
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16378 | 767.7188
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16388 | 768.1875
66 | 16389 | 768.2344
66 | 16385 | 768.0469
66 | 16383 | 767.9531
66 | 16378 | 767.7188
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16378 | 767.7188
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16380 | 767.8125
66 | 16378 | 767.7188
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16388 | 768.1875
66 | 16387 | 768.1406
66 | 16385 | 768.0469
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16387 | 768.1406
66 | 16386 | 768.0938
66 | 16386 | 768.0938
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16385 | 768.0469
66 | 16387 | 768.1406
66 | 16386 | 768.0938
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16381 | 767.8594
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16388 | 768.1875
66 | 16388 | 768.1875
66 | 16385 | 768.0469
66 | 16383 | 767.9531
66 | 16377 | 767.6719
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16386 | 768.0938
66 | 16384 | 768
66 | 16381 | 767.8594
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16383 | 767.9531
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16386 | 768.0938
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16381 | 767.8594
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16378 | 767.7188
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16384 | 768
66 | 16385 | 768.0469
66 | 16388 | 768.1875
66 | 16386 | 768.0938
66 | 16384 | 768
66 | 16379 | 767.7656
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16384 | 768
66 | 16383 | 767.9531
66 | 16379 | 767.7656
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16379 | 767.7656
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16386 | 768.0938
66 | 16387 | 768.1406
66 | 16387 | 768.1406
66 | 16384 | 768
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16386 | 768.0938
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16384 | 768
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16377 | 767.6719
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16384 | 768
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16384 | 768
66 | 16380 | 767.8125
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16387 | 768.1406
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16376 | 767.625
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16384 | 768
66 | 16383 | 767.9531
66 | 16380 | 767.8125
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16378 | 767.7188
66 | 16374 | 767.5312
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16377 | 767.6719
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16385 | 768.0469
66 | 16385 | 768.0469
66 | 16384 | 768
66 | 16384 | 768
66 | 16382 | 767.9062
66 | 16376 | 767.625
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16384 | 768
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16379 | 767.7656
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16378 | 767.7188
66 | 16376 | 767.625
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16377 | 767.6719
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16384 | 768
66 | 16384 | 768
66 | 16384 | 768
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16377 | 767.6719
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16376 | 767.625
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16376 | 767.625
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16378 | 767.7188
66 | 16377 | 767.6719
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16378 | 767.7188
66 | 16376 | 767.625
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16375 | 767.5781
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16378 | 767.7188
66 | 16377 | 767.6719
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16383 | 767.9531
66 | 16380 | 767.8125
66 | 16383 | 767.9531
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16382 | 767.9062
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16382 | 767.9062
66 | 16383 | 767.9531
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16379 | 767.7656
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16381 | 767.8594
66 | 16381 | 767.8594
66 | 16380 | 767.8125
66 | 16380 | 767.8125
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16378 | 767.7188
66 | 16377 | 767.6719
66 | 16377 | 767.6719
66 | 16377 | 767.6719
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16373 | 767.4844
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16378 | 767.7188
66 | 16378 | 767.7188
66 | 16378 | 767.7188
66 | 16379 | 767.7656
66 | 16379 | 767.7656
66 | 16377 | 767.6719
66 | 16377 | 767.6719
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16377 | 767.6719
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16372 | 767.4375
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16368 | 767.25
66 | 16371 | 767.3906
66 | 16373 | 767.4844
66 | 16376 | 767.625
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16371 | 767.3906
66 | 16370 | 767.3438
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16375 | 767.5781
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16369 | 767.2969
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16371 | 767.3906
66 | 16368 | 767.25
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16369 | 767.2969
66 | 16371 | 767.3906
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16368 | 767.25
66 | 16371 | 767.3906
66 | 16373 | 767.4844
66 | 16374 | 767.5312
66 | 16376 | 767.625
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16372 | 767.4375
66 | 16374 | 767.5312
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16366 | 767.1562
66 | 16362 | 766.9688
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16366 | 767.1562
66 | 16371 | 767.3906
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16375 | 767.5781
66 | 16374 | 767.5312
66 | 16374 | 767.5312
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16362 | 766.9688
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16367 | 767.2031
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16363 | 767.0156
66 | 16363 | 767.0156
66 | 16363 | 767.0156
66 | 16366 | 767.1562
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16367 | 767.2031
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16366 | 767.1562
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16373 | 767.4844
66 | 16372 | 767.4375
66 | 16372 | 767.4375
66 | 16371 | 767.3906
66 | 16370 | 767.3438
66 | 16367 | 767.2031
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16367 | 767.2031
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16372 | 767.4375
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16367 | 767.2031
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16365 | 767.1094
66 | 16368 | 767.25
66 | 16372 | 767.4375
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16367 | 767.2031
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16373 | 767.4844
66 | 16371 | 767.3906
66 | 16371 | 767.3906
66 | 16368 | 767.25
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16360 | 766.875
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16368 | 767.25
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16366 | 767.1562
66 | 16363 | 767.0156
66 | 16361 | 766.9219
66 | 16362 | 766.9688
66 | 16362 | 766.9688
66 | 16364 | 767.0625
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16363 | 767.0156
66 | 16365 | 767.1094
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16361 | 766.9219
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16366 | 767.1562
66 | 16361 | 766.9219
66 | 16360 | 766.875
66 | 16363 | 767.0156
66 | 16369 | 767.2969
66 | 16370 | 767.3438
66 | 16372 | 767.4375
66 | 16370 | 767.3438
66 | 16368 | 767.25
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16359 | 766.8281
66 | 16365 | 767.1094
66 | 16367 | 767.2031
66 | 16370 | 767.3438
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16360 | 766.875
66 | 16361 | 766.9219
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16368 | 767.25
66 | 16371 | 767.3906
66 | 16369 | 767.2969
66 | 16366 | 767.1562
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16364 | 767.0625
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16362 | 766.9688
66 | 16359 | 766.8281
66 | 16361 | 766.9219
66 | 16365 | 767.1094
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16361 | 766.9219
66 | 16358 | 766.7812
66 | 16359 | 766.8281
66 | 16359 | 766.8281
66 | 16365 | 767.1094
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16367 | 767.2031
66 | 16369 | 767.2969
66 | 16369 | 767.2969
66 | 16367 | 767.2031
66 | 16363 | 767.0156
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16363 | 767.0156
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16361 | 766.9219
66 | 16361 | 766.9219
66 | 16362 | 766.9688
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16364 | 767.0625
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16359 | 766.8281
66 | 16363 | 767.0156
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16361 | 766.9219
66 | 16360 | 766.875
66 | 16361 | 766.9219
66 | 16364 | 767.0625
66 | 16368 | 767.25
66 | 16369 | 767.2969
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16364 | 767.0625
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16359 | 766.8281
66 | 16359 | 766.8281
66 | 16364 | 767.0625
66 | 16366 | 767.1562
66 | 16370 | 767.3438
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16361 | 766.9219
66 | 16356 | 766.6875
66 | 16362 | 766.9688
66 | 16365 | 767.1094
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16358 | 766.7812
66 | 16358 | 766.7812
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16368 | 767.25
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16361 | 766.9219
66 | 16358 | 766.7812
66 | 16357 | 766.7344
66 | 16356 | 766.6875
66 | 16360 | 766.875
66 | 16365 | 767.1094
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16360 | 766.875
66 | 16365 | 767.1094
66 | 16368 | 767.25
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16363 | 767.0156
66 | 16359 | 766.8281
66 | 16356 | 766.6875
66 | 16359 | 766.8281
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16363 | 767.0156
66 | 16358 | 766.7812
66 | 16358 | 766.7812
66 | 16361 | 766.9219
66 | 16366 | 767.1562
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16357 | 766.7344
66 | 16360 | 766.875
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16359 | 766.8281
66 | 16356 | 766.6875
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16366 | 767.1562
66 | 16368 | 767.25
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16358 | 766.7812
66 | 16360 | 766.875
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16357 | 766.7344
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16361 | 766.9219
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16362 | 766.9688
66 | 16356 | 766.6875
66 | 16359 | 766.8281
66 | 16367 | 767.2031
66 | 16366 | 767.1562
66 | 16363 | 767.0156
66 | 16359 | 766.8281
66 | 16356 | 766.6875
66 | 16356 | 766.6875
66 | 16365 | 767.1094
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16362 | 766.9688
66 | 16359 | 766.8281
66 | 16358 | 766.7812
66 | 16363 | 767.0156
66 | 16369 | 767.2969
66 | 16368 | 767.25
66 | 16364 | 767.0625
66 | 16359 | 766.8281
66 | 16356 | 766.6875
66 | 16359 | 766.8281
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16357 | 766.7344
66 | 16366 | 767.1562
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16362 | 766.9688
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16362 | 766.9688
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16356 | 766.6875
66 | 16356 | 766.6875
66 | 16365 | 767.1094
66 | 16367 | 767.2031
66 | 16367 | 767.2031
66 | 16364 | 767.0625
66 | 16357 | 766.7344
66 | 16356 | 766.6875
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16358 | 766.7812
66 | 16354 | 766.5938
66 | 16358 | 766.7812
66 | 16364 | 767.0625
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16363 | 767.0156
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16355 | 766.6406
66 | 16358 | 766.7812
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16359 | 766.8281
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16355 | 766.6406
66 | 16355 | 766.6406
66 | 16362 | 766.9688
66 | 16366 | 767.1562
66 | 16366 | 767.1562
66 | 16362 | 766.9688
66 | 16356 | 766.6875
66 | 16360 | 766.875
66 | 16367 | 767.2031
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16356 | 766.6875
66 | 16362 | 766.9688
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16356 | 766.6875
66 | 16354 | 766.5938
66 | 16360 | 766.875
66 | 16365 | 767.1094
66 | 16365 | 767.1094
66 | 16362 | 766.9688
66 | 16355 | 766.6406
66 | 16359 | 766.8281
66 | 16368 | 767.25
66 | 16366 | 767.1562
66 | 16362 | 766.9688
66 | 16357 | 766.7344
66 | 16355 | 766.6406
66 | 16359 | 766.8281
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16356 | 766.6875
66 | 16353 | 766.5469
66 | 16360 | 766.875
66 | 16365 | 767.1094
66 | 16366 | 767.1562
66 | 16361 | 766.9219
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16363 | 767.0156
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16360 | 766.875
66 | 16354 | 766.5938
66 | 16359 | 766.8281
66 | 16362 | 766.9688
66 | 16362 | 766.9688
66 | 16360 | 766.875
66 | 16355 | 766.6406
66 | 16353 | 766.5469
66 | 16362 | 766.9688
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16362 | 766.9688
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16361 | 766.9219
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16364 | 767.0625
66 | 16363 | 767.0156
66 | 16361 | 766.9219
66 | 16356 | 766.6875
66 | 16359 | 766.8281
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16363 | 767.0156
66 | 16356 | 766.6875
66 | 16360 | 766.875
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16357 | 766.7344
66 | 16353 | 766.5469
66 | 16362 | 766.9688
66 | 16364 | 767.0625
66 | 16357 | 766.7344
66 | 16355 | 766.6406
66 | 16363 | 767.0156
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16355 | 766.6406
66 | 16359 | 766.8281
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16354 | 766.5938
66 | 16361 | 766.9219
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16355 | 766.6406
66 | 16354 | 766.5938
66 | 16361 | 766.9219
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16354 | 766.5938
66 | 16353 | 766.5469
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16356 | 766.6875
66 | 16355 | 766.6406
66 | 16366 | 767.1562
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16356 | 766.6875
66 | 16358 | 766.7812
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16352 | 766.5
66 | 16358 | 766.7812
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16355 | 766.6406
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16359 | 766.8281
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16366 | 767.1562
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16362 | 766.9688
66 | 16362 | 766.9688
66 | 16355 | 766.6406
66 | 16354 | 766.5938
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16356 | 766.6875
66 | 16356 | 766.6875
66 | 16366 | 767.1562
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16359 | 766.8281
66 | 16354 | 766.5938
66 | 16356 | 766.6875
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16353 | 766.5469
66 | 16355 | 766.6406
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16358 | 766.7812
66 | 16353 | 766.5469
66 | 16360 | 766.875
66 | 16363 | 767.0156
66 | 16359 | 766.8281
66 | 16354 | 766.5938
66 | 16355 | 766.6406
66 | 16362 | 766.9688
66 | 16362 | 766.9688
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16364 | 767.0625
66 | 16365 | 767.1094
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16360 | 766.875
66 | 16364 | 767.0625
66 | 16359 | 766.8281
66 | 16351 | 766.4531
66 | 16355 | 766.6406
66 | 16363 | 767.0156
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16354 | 766.5938
66 | 16364 | 767.0625
66 | 16364 | 767.0625
66 | 16358 | 766.7812
66 | 16356 | 766.6875
66 | 16362 | 766.9688
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16353 | 766.5469
66 | 16357 | 766.7344
66 | 16361 | 766.9219
66 | 16358 | 766.7812
66 | 16354 | 766.5938
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16364 | 767.0625
66 | 16362 | 766.9688
66 | 16353 | 766.5469
66 | 16358 | 766.7812
66 | 16364 | 767.0625
66 | 16360 | 766.875
66 | 16352 | 766.5
66 | 16357 | 766.7344
66 | 16362 | 766.9688
66 | 16360 | 766.875
66 | 16352 | 766.5
66 | 16353 | 766.5469
66 | 16365 | 767.1094
66 | 16363 | 767.0156
66 | 16357 | 766.7344
66 | 16354 | 766.5938
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16354 | 766.5938
66 | 16362 | 766.9688
66 | 16363 | 767.0156
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16362 | 766.9688
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16351 | 766.4531
66 | 16361 | 766.9219
66 | 16363 | 767.0156
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16365 | 767.1094
66 | 16364 | 767.0625
66 | 16356 | 766.6875
66 | 16353 | 766.5469
66 | 16362 | 766.9688
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16352 | 766.5
66 | 16361 | 766.9219
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16363 | 767.0156
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16354 | 766.5938
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16354 | 766.5938
66 | 16364 | 767.0625
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16362 | 766.9688
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16354 | 766.5938
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16360 | 766.875
66 | 16363 | 767.0156
66 | 16356 | 766.6875
66 | 16351 | 766.4531
66 | 16361 | 766.9219
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16360 | 766.875
66 | 16355 | 766.6406
66 | 16353 | 766.5469
66 | 16364 | 767.0625
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16354 | 766.5938
66 | 16363 | 767.0156
66 | 16362 | 766.9688
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16361 | 766.9219
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16352 | 766.5
66 | 16362 | 766.9688
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16359 | 766.8281
66 | 16361 | 766.9219
66 | 16355 | 766.6406
66 | 16356 | 766.6875
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16355 | 766.6406
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16359 | 766.8281
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16361 | 766.9219
66 | 16361 | 766.9219
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16362 | 766.9688
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16352 | 766.5
66 | 16354 | 766.5938
66 | 16361 | 766.9219
66 | 16357 | 766.7344
66 | 16351 | 766.4531
66 | 16362 | 766.9688
66 | 16361 | 766.9219
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16361 | 766.9219
66 | 16359 | 766.8281
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16357 | 766.7344
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16363 | 767.0156
66 | 16360 | 766.875
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16360 | 766.875
66 | 16359 | 766.8281
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16359 | 766.8281
66 | 16356 | 766.6875
66 | 16348 | 766.3125
66 | 16356 | 766.6875
66 | 16357 | 766.7344
66 | 16354 | 766.5938
66 | 16354 | 766.5938
66 | 16361 | 766.9219
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16358 | 766.7812
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16358 | 766.7812
66 | 16357 | 766.7344
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16361 | 766.9219
66 | 16353 | 766.5469
66 | 16350 | 766.4062
66 | 16360 | 766.875
66 | 16357 | 766.7344
66 | 16351 | 766.4531
66 | 16349 | 766.3594
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16351 | 766.4531
66 | 16352 | 766.5
66 | 16361 | 766.9219
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16359 | 766.8281
66 | 16360 | 766.875
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16359 | 766.8281
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16359 | 766.8281
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16358 | 766.7812
66 | 16360 | 766.875
66 | 16354 | 766.5938
66 | 16349 | 766.3594
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16348 | 766.3125
66 | 16356 | 766.6875
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16358 | 766.7812
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16350 | 766.4062
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16356 | 766.6875
66 | 16356 | 766.6875
66 | 16348 | 766.3125
66 | 16359 | 766.8281
66 | 16358 | 766.7812
66 | 16349 | 766.3594
66 | 16359 | 766.8281
66 | 16359 | 766.8281
66 | 16351 | 766.4531
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16348 | 766.3125
66 | 16358 | 766.7812
66 | 16356 | 766.6875
66 | 16347 | 766.2656
66 | 16358 | 766.7812
66 | 16357 | 766.7344
66 | 16347 | 766.2656
66 | 16358 | 766.7812
66 | 16357 | 766.7344
66 | 16349 | 766.3594
66 | 16362 | 766.9688
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16356 | 766.6875
66 | 16351 | 766.4531
66 | 16361 | 766.9219
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16356 | 766.6875
66 | 16352 | 766.5
66 | 16347 | 766.2656
66 | 16359 | 766.8281
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16360 | 766.875
66 | 16358 | 766.7812
66 | 16349 | 766.3594
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16349 | 766.3594
66 | 16360 | 766.875
66 | 16357 | 766.7344
66 | 16348 | 766.3125
66 | 16359 | 766.8281
66 | 16354 | 766.5938
66 | 16350 | 766.4062
66 | 16359 | 766.8281
66 | 16355 | 766.6406
66 | 16355 | 766.6406
66 | 16361 | 766.9219
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16356 | 766.6875
66 | 16357 | 766.7344
66 | 16348 | 766.3125
66 | 16352 | 766.5
66 | 16356 | 766.6875
66 | 16348 | 766.3125
66 | 16358 | 766.7812
66 | 16358 | 766.7812
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16360 | 766.875
66 | 16357 | 766.7344
66 | 16349 | 766.3594
66 | 16356 | 766.6875
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16348 | 766.3125
66 | 16356 | 766.6875
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16362 | 766.9688
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16360 | 766.875
66 | 16355 | 766.6406
66 | 16351 | 766.4531
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16359 | 766.8281
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16358 | 766.7812
66 | 16353 | 766.5469
66 | 16355 | 766.6406
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16360 | 766.875
66 | 16353 | 766.5469
66 | 16357 | 766.7344
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16349 | 766.3594
66 | 16357 | 766.7344
66 | 16351 | 766.4531
66 | 16355 | 766.6406
66 | 16358 | 766.7812
66 | 16350 | 766.4062
66 | 16355 | 766.6406
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16359 | 766.8281
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16359 | 766.8281
66 | 16353 | 766.5469
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16352 | 766.5
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16350 | 766.4062
66 | 16358 | 766.7812
66 | 16355 | 766.6406
66 | 16349 | 766.3594
66 | 16357 | 766.7344
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16356 | 766.6875
66 | 16352 | 766.5
66 | 16350 | 766.4062
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16357 | 766.7344
66 | 16355 | 766.6406
66 | 16353 | 766.5469
66 | 16359 | 766.8281
66 | 16352 | 766.5
66 | 16349 | 766.3594
66 | 16358 | 766.7812
66 | 16353 | 766.5469
66 | 16348 | 766.3125
66 | 16357 | 766.7344
66 | 16350 | 766.4062
66 | 16347 | 766.2656
66 | 16355 | 766.6406
66 | 16351 | 766.4531
66 | 16353 | 766.5469
66 | 16358 | 766.7812
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16357 | 766.7344
66 | 16358 | 766.7812
66 | 16349 | 766.3594
66 | 16353 | 766.5469
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16356 | 766.6875
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16352 | 766.5
66 | 16357 | 766.7344
66 | 16351 | 766.4531
66 | 16355 | 766.6406
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16358 | 766.7812
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16356 | 766.6875
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16357 | 766.7344
66 | 16354 | 766.5938
66 | 16348 | 766.3125
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16356 | 766.6875
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16347 | 766.2656
66 | 16356 | 766.6875
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16357 | 766.7344
66 | 16349 | 766.3594
66 | 16351 | 766.4531
66 | 16356 | 766.6875
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16355 | 766.6406
66 | 16346 | 766.2188
66 | 16355 | 766.6406
66 | 16352 | 766.5
66 | 16347 | 766.2656
66 | 16354 | 766.5938
66 | 16349 | 766.3594
66 | 16356 | 766.6875
66 | 16355 | 766.6406
66 | 16351 | 766.4531
66 | 16358 | 766.7812
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16355 | 766.6406
66 | 16355 | 766.6406
66 | 16346 | 766.2188
66 | 16350 | 766.4062
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16355 | 766.6406
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16357 | 766.7344
66 | 16349 | 766.3594
66 | 16354 | 766.5938
66 | 16355 | 766.6406
66 | 16348 | 766.3125
66 | 16357 | 766.7344
66 | 16352 | 766.5
66 | 16349 | 766.3594
66 | 16355 | 766.6406
66 | 16348 | 766.3125
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16355 | 766.6406
66 | 16352 | 766.5
66 | 16347 | 766.2656
66 | 16359 | 766.8281
66 | 16354 | 766.5938
66 | 16350 | 766.4062
66 | 16357 | 766.7344
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16354 | 766.5938
66 | 16345 | 766.1719
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16345 | 766.1719
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16350 | 766.4062
66 | 16356 | 766.6875
66 | 16347 | 766.2656
66 | 16358 | 766.7812
66 | 16353 | 766.5469
66 | 16348 | 766.3125
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16354 | 766.5938
66 | 16353 | 766.5469
66 | 16347 | 766.2656
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16354 | 766.5938
66 | 16343 | 766.0781
66 | 16353 | 766.5469
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16355 | 766.6406
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16356 | 766.6875
66 | 16347 | 766.2656
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16347 | 766.2656
66 | 16356 | 766.6875
66 | 16349 | 766.3594
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16345 | 766.1719
66 | 16354 | 766.5938
66 | 16347 | 766.2656
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16349 | 766.3594
66 | 16356 | 766.6875
66 | 16347 | 766.2656
66 | 16357 | 766.7344
66 | 16351 | 766.4531
66 | 16354 | 766.5938
66 | 16354 | 766.5938
66 | 16350 | 766.4062
66 | 16354 | 766.5938
66 | 16343 | 766.0781
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16355 | 766.6406
66 | 16354 | 766.5938
66 | 16348 | 766.3125
66 | 16357 | 766.7344
66 | 16348 | 766.3125
66 | 16355 | 766.6406
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16353 | 766.5469
66 | 16347 | 766.2656
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16354 | 766.5938
66 | 16349 | 766.3594
66 | 16354 | 766.5938
66 | 16353 | 766.5469
66 | 16347 | 766.2656
66 | 16356 | 766.6875
66 | 16346 | 766.2188
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16353 | 766.5469
66 | 16353 | 766.5469
66 | 16344 | 766.125
66 | 16356 | 766.6875
66 | 16351 | 766.4531
66 | 16349 | 766.3594
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16353 | 766.5469
66 | 16352 | 766.5
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16347 | 766.2656
66 | 16352 | 766.5
66 | 16344 | 766.125
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16354 | 766.5938
66 | 16346 | 766.2188
66 | 16356 | 766.6875
66 | 16348 | 766.3125
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16352 | 766.5
66 | 16350 | 766.4062
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16346 | 766.2188
66 | 16353 | 766.5469
66 | 16350 | 766.4062
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16345 | 766.1719
66 | 16355 | 766.6406
66 | 16349 | 766.3594
66 | 16351 | 766.4531
66 | 16352 | 766.5
66 | 16343 | 766.0781
66 | 16353 | 766.5469
66 | 16349 | 766.3594
66 | 16344 | 766.125
66 | 16354 | 766.5938
66 | 16347 | 766.2656
66 | 16355 | 766.6406
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16356 | 766.6875
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16352 | 766.5
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16344 | 766.125
66 | 16355 | 766.6406
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16354 | 766.5938
66 | 16344 | 766.125
66 | 16355 | 766.6406
66 | 16348 | 766.3125
66 | 16352 | 766.5
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16351 | 766.4531
66 | 16344 | 766.125
66 | 16351 | 766.4531
66 | 16347 | 766.2656
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16345 | 766.1719
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16348 | 766.3125
66 | 16353 | 766.5469
66 | 16350 | 766.4062
66 | 16347 | 766.2656
66 | 16352 | 766.5
66 | 16343 | 766.0781
66 | 16352 | 766.5
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16351 | 766.4531
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16347 | 766.2656
66 | 16356 | 766.6875
66 | 16352 | 766.5
66 | 16349 | 766.3594
66 | 16353 | 766.5469
66 | 16343 | 766.0781
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16343 | 766.0781
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16354 | 766.5938
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16346 | 766.2188
66 | 16354 | 766.5938
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16345 | 766.1719
66 | 16350 | 766.4062
66 | 16343 | 766.0781
66 | 16353 | 766.5469
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16353 | 766.5469
66 | 16350 | 766.4062
66 | 16353 | 766.5469
66 | 16343 | 766.0781
66 | 16352 | 766.5
66 | 16341 | 765.9844
66 | 16351 | 766.4531
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16352 | 766.5
66 | 16344 | 766.125
66 | 16355 | 766.6406
66 | 16347 | 766.2656
66 | 16353 | 766.5469
66 | 16347 | 766.2656
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16350 | 766.4062
66 | 16348 | 766.3125
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16351 | 766.4531
66 | 16349 | 766.3594
66 | 16351 | 766.4531
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16349 | 766.3594
66 | 16343 | 766.0781
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16354 | 766.5938
66 | 16347 | 766.2656
66 | 16352 | 766.5
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16349 | 766.3594
66 | 16347 | 766.2656
66 | 16350 | 766.4062
66 | 16343 | 766.0781
66 | 16349 | 766.3594
66 | 16341 | 765.9844
66 | 16352 | 766.5
66 | 16345 | 766.1719
66 | 16355 | 766.6406
66 | 16348 | 766.3125
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16353 | 766.5469
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16351 | 766.4531
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16348 | 766.3125
66 | 16352 | 766.5
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16343 | 766.0781
66 | 16351 | 766.4531
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16345 | 766.1719
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16346 | 766.2188
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16348 | 766.3125
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16347 | 766.2656
66 | 16346 | 766.2188
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16348 | 766.3125
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16348 | 766.3125
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16341 | 765.9844
66 | 16353 | 766.5469
66 | 16345 | 766.1719
66 | 16352 | 766.5
66 | 16349 | 766.3594
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16352 | 766.5
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16350 | 766.4062
66 | 16348 | 766.3125
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16346 | 766.2188
66 | 16343 | 766.0781
66 | 16349 | 766.3594
66 | 16342 | 766.0312
66 | 16351 | 766.4531
66 | 16343 | 766.0781
66 | 16353 | 766.5469
66 | 16345 | 766.1719
66 | 16353 | 766.5469
66 | 16345 | 766.1719
66 | 16351 | 766.4531
66 | 16348 | 766.3125
66 | 16349 | 766.3594
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16346 | 766.2188
66 | 16346 | 766.2188
66 | 16346 | 766.2188
66 | 16344 | 766.125
66 | 16349 | 766.3594
66 | 16342 | 766.0312
66 | 16350 | 766.4062
66 | 16343 | 766.0781
66 | 16353 | 766.5469
66 | 16346 | 766.2188
66 | 16351 | 766.4531
66 | 16350 | 766.4062
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16346 | 766.2188
66 | 16349 | 766.3594
66 | 16343 | 766.0781
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16352 | 766.5
66 | 16342 | 766.0312
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16348 | 766.3125
66 | 16343 | 766.0781
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16343 | 766.0781
66 | 16350 | 766.4062
66 | 16346 | 766.2188
66 | 16349 | 766.3594
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16349 | 766.3594
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16349 | 766.3594
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16351 | 766.4531
66 | 16349 | 766.3594
66 | 16350 | 766.4062
66 | 16347 | 766.2656
66 | 16350 | 766.4062
66 | 16344 | 766.125
66 | 16350 | 766.4062
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16345 | 766.1719
66 | 16349 | 766.3594
66 | 16344 | 766.125
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16349 | 766.3594
66 | 16348 | 766.3125
66 | 16352 | 766.5
66 | 16346 | 766.2188
66 | 16350 | 766.4062
66 | 16347 | 766.2656
66 | 16350 | 766.4062
66 | 16348 | 766.3125
66 | 16349 | 766.3594
66 | 16344 | 766.125
66 | 16349 | 766.3594
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16344 | 766.125
66 | 16349 | 766.3594
66 | 16343 | 766.0781
66 | 16350 | 766.4062
66 | 16346 | 766.2188
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16349 | 766.3594
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16347 | 766.2656
66 | 16348 | 766.3125
66 | 16346 | 766.2188
66 | 16347 | 766.2656
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16344 | 766.125
66 | 16346 | 766.2188
66 | 16344 | 766.125
66 | 16347 | 766.2656
66 | 16343 | 766.0781
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16349 | 766.3594
66 | 16346 | 766.2188
66 | 16348 | 766.3125
66 | 16345 | 766.1719
66 | 16348 | 766.3125
66 | 16344 | 766.125
66 | 16348 | 766.3125
66 | 16345 | 766.1719
66 | 16347 | 766.2656
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16343 | 766.0781
66 | 16347 | 766.2656
66 | 16344 | 766.125
66 | 16347 | 766.2656
66 | 16343 | 766.0781
66 | 16348 | 766.3125
66 | 16342 | 766.0312
66 | 16350 | 766.4062
66 | 16344 | 766.125
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16349 | 766.3594
66 | 16343 | 766.0781
66 | 16348 | 766.3125
66 | 16342 | 766.0312
66 | 16347 | 766.2656
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16342 | 766.0312
66 | 16348 | 766.3125
66 | 16341 | 765.9844
66 | 16349 | 766.3594
66 | 16341 | 765.9844
66 | 16349 | 766.3594
66 | 16342 | 766.0312
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16350 | 766.4062
66 | 16341 | 765.9844
66 | 16349 | 766.3594
66 | 16342 | 766.0312
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16341 | 765.9844
66 | 16348 | 766.3125
66 | 16341 | 765.9844
66 | 16350 | 766.4062
66 | 16343 | 766.0781
66 | 16351 | 766.4531
66 | 16344 | 766.125
66 | 16350 | 766.4062
66 | 16345 | 766.1719
66 | 16350 | 766.4062
66 | 16343 | 766.0781
66 | 16350 | 766.4062
66 | 16342 | 766.0312
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16349 | 766.3594
66 | 16341 | 765.9844
66 | 16350 | 766.4062
66 | 16341 | 765.9844
66 | 16348 | 766.3125
66 | 16342 | 766.0312
66 | 16350 | 766.4062
66 | 16344 | 766.125
66 | 16352 | 766.5
66 | 16343 | 766.0781
66 | 16350 | 766.4062
66 | 16344 | 766.125
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16339 | 765.8906
66 | 16348 | 766.3125
66 | 16339 | 765.8906
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16341 | 765.9844
66 | 16348 | 766.3125
66 | 16341 | 765.9844
66 | 16350 | 766.4062
66 | 16342 | 766.0312
66 | 16351 | 766.4531
66 | 16343 | 766.0781
66 | 16350 | 766.4062
66 | 16341 | 765.9844
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16351 | 766.4531
66 | 16341 | 765.9844
66 | 16351 | 766.4531
66 | 16341 | 765.9844
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16348 | 766.3125
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16338 | 765.8438
66 | 16347 | 766.2656
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16337 | 765.7969
66 | 16348 | 766.3125
66 | 16337 | 765.7969
66 | 16348 | 766.3125
66 | 16339 | 765.8906
66 | 16347 | 766.2656
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16351 | 766.4531
66 | 16342 | 766.0312
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16340 | 765.9375
66 | 16347 | 766.2656
66 | 16337 | 765.7969
66 | 16347 | 766.2656
66 | 16338 | 765.8438
66 | 16350 | 766.4062
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16350 | 766.4062
66 | 16340 | 765.9375
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16348 | 766.3125
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16337 | 765.7969
66 | 16348 | 766.3125
66 | 16338 | 765.8438
66 | 16349 | 766.3594
66 | 16337 | 765.7969
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16352 | 766.5
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16339 | 765.8906
66 | 16348 | 766.3125
66 | 16337 | 765.7969
66 | 16349 | 766.3594
66 | 16336 | 765.75
66 | 16349 | 766.3594
66 | 16338 | 765.8438
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16338 | 765.8438
66 | 16350 | 766.4062
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16341 | 765.9844
66 | 16350 | 766.4062
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16340 | 765.9375
66 | 16350 | 766.4062
66 | 16338 | 765.8438
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16338 | 765.8438
66 | 16347 | 766.2656
66 | 16335 | 765.7031
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16348 | 766.3125
66 | 16337 | 765.7969
66 | 16348 | 766.3125
66 | 16339 | 765.8906
66 | 16348 | 766.3125
66 | 16338 | 765.8438
66 | 16349 | 766.3594
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16337 | 765.7969
66 | 16349 | 766.3594
66 | 16337 | 765.7969
66 | 16348 | 766.3125
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16337 | 765.7969
66 | 16347 | 766.2656
66 | 16338 | 765.8438
66 | 16346 | 766.2188
66 | 16339 | 765.8906
66 | 16349 | 766.3594
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16344 | 766.125
66 | 16339 | 765.8906
66 | 16347 | 766.2656
66 | 16335 | 765.7031
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16338 | 765.8438
66 | 16347 | 766.2656
66 | 16344 | 766.125
66 | 16341 | 765.9844
66 | 16349 | 766.3594
66 | 16338 | 765.8438
66 | 16348 | 766.3125
66 | 16337 | 765.7969
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16342 | 766.0312
66 | 16337 | 765.7969
66 | 16347 | 766.2656
66 | 16335 | 765.7031
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16337 | 765.7969
66 | 16345 | 766.1719
66 | 16344 | 766.125
66 | 16339 | 765.8906
66 | 16347 | 766.2656
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16343 | 766.0781
66 | 16338 | 765.8438
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16337 | 765.7969
66 | 16346 | 766.2188
66 | 16345 | 766.1719
66 | 16337 | 765.7969
66 | 16348 | 766.3125
66 | 16335 | 765.7031
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16342 | 766.0312
66 | 16336 | 765.75
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16346 | 766.2188
66 | 16338 | 765.8438
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16345 | 766.1719
66 | 16343 | 766.0781
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16342 | 766.0312
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16345 | 766.1719
66 | 16346 | 766.2188
66 | 16338 | 765.8438
66 | 16347 | 766.2656
66 | 16336 | 765.75
66 | 16344 | 766.125
66 | 16345 | 766.1719
66 | 16337 | 765.7969
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16341 | 765.9844
66 | 16335 | 765.7031
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16340 | 765.9375
66 | 16339 | 765.8906
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16338 | 765.8438
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16342 | 766.0312
66 | 16336 | 765.75
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16345 | 766.1719
66 | 16337 | 765.7969
66 | 16346 | 766.2188
66 | 16337 | 765.7969
66 | 16345 | 766.1719
66 | 16342 | 766.0312
66 | 16338 | 765.8438
66 | 16347 | 766.2656
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16345 | 766.1719
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16345 | 766.1719
66 | 16337 | 765.7969
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16343 | 766.0781
66 | 16336 | 765.75
66 | 16347 | 766.2656
66 | 16335 | 765.7031
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16347 | 766.2656
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16344 | 766.125
66 | 16336 | 765.75
66 | 16346 | 766.2188
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16347 | 766.2656
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16338 | 765.8438
66 | 16337 | 765.7969
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16345 | 766.1719
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16341 | 765.9844
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16344 | 766.125
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16345 | 766.1719
66 | 16332 | 765.5625
66 | 16341 | 765.9844
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16341 | 765.9844
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16334 | 765.6562
66 | 16343 | 766.0781
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16337 | 765.7969
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16341 | 765.9844
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16344 | 766.125
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16340 | 765.9375
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16343 | 766.0781
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16344 | 766.125
66 | 16340 | 765.9375
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16341 | 765.9844
66 | 16335 | 765.7031
66 | 16343 | 766.0781
66 | 16341 | 765.9844
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16336 | 765.75
66 | 16336 | 765.75
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16345 | 766.1719
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16344 | 766.125
66 | 16331 | 765.5156
66 | 16338 | 765.8438
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16345 | 766.1719
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16345 | 766.1719
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16345 | 766.1719
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16346 | 766.2188
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16345 | 766.1719
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16344 | 766.125
66 | 16336 | 765.75
66 | 16337 | 765.7969
66 | 16344 | 766.125
66 | 16337 | 765.7969
66 | 16337 | 765.7969
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16344 | 766.125
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16345 | 766.1719
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16345 | 766.1719
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16342 | 766.0312
66 | 16346 | 766.2188
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16346 | 766.2188
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16344 | 766.125
66 | 16337 | 765.7969
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16337 | 765.7969
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16339 | 765.8906
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16344 | 766.125
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16334 | 765.6562
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16342 | 766.0312
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16335 | 765.7031
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16337 | 765.7969
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16342 | 766.0312
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16336 | 765.75
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16339 | 765.8906
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16341 | 765.9844
66 | 16329 | 765.4219
66 | 16339 | 765.8906
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16338 | 765.8438
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16338 | 765.8438
66 | 16343 | 766.0781
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16343 | 766.0781
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16334 | 765.6562
66 | 16332 | 765.5625
66 | 16341 | 765.9844
66 | 16344 | 766.125
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16332 | 765.5625
66 | 16331 | 765.5156
66 | 16338 | 765.8438
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16334 | 765.6562
66 | 16342 | 766.0312
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16333 | 765.6094
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16337 | 765.7969
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16335 | 765.7031
66 | 16332 | 765.5625
66 | 16341 | 765.9844
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16335 | 765.7031
66 | 16327 | 765.3281
66 | 16338 | 765.8438
66 | 16340 | 765.9375
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16341 | 765.9844
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16331 | 765.5156
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16333 | 765.6094
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16342 | 766.0312
66 | 16335 | 765.7031
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16341 | 765.9844
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16332 | 765.5625
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16337 | 765.7969
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16341 | 765.9844
66 | 16337 | 765.7969
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16343 | 766.0781
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16339 | 765.8906
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16341 | 765.9844
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16340 | 765.9375
66 | 16333 | 765.6094
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16336 | 765.75
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16333 | 765.6094
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16331 | 765.5156
66 | 16340 | 765.9375
66 | 16333 | 765.6094
66 | 16332 | 765.5625
66 | 16341 | 765.9844
66 | 16336 | 765.75
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16335 | 765.7031
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16334 | 765.6562
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16329 | 765.4219
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16334 | 765.6562
66 | 16328 | 765.375
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16338 | 765.8438
66 | 16330 | 765.4688
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16342 | 766.0312
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16335 | 765.7031
66 | 16328 | 765.375
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16336 | 765.75
66 | 16329 | 765.4219
66 | 16338 | 765.8438
66 | 16342 | 766.0312
66 | 16331 | 765.5156
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16335 | 765.7031
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16335 | 765.7031
66 | 16328 | 765.375
66 | 16337 | 765.7969
66 | 16339 | 765.8906
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16335 | 765.7031
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16332 | 765.5625
66 | 16329 | 765.4219
66 | 16338 | 765.8438
66 | 16340 | 765.9375
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16342 | 766.0312
66 | 16334 | 765.6562
66 | 16332 | 765.5625
66 | 16340 | 765.9375
66 | 16342 | 766.0312
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16337 | 765.7969
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16335 | 765.7031
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16334 | 765.6562
66 | 16329 | 765.4219
66 | 16338 | 765.8438
66 | 16340 | 765.9375
66 | 16328 | 765.375
66 | 16333 | 765.6094
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16338 | 765.8438
66 | 16334 | 765.6562
66 | 16329 | 765.4219
66 | 16337 | 765.7969
66 | 16340 | 765.9375
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16340 | 765.9375
66 | 16333 | 765.6094
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16335 | 765.7031
66 | 16339 | 765.8906
66 | 16330 | 765.4688
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16335 | 765.7031
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16331 | 765.5156
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16334 | 765.6562
66 | 16328 | 765.375
66 | 16338 | 765.8438
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16329 | 765.4219
66 | 16331 | 765.5156
66 | 16339 | 765.8906
66 | 16332 | 765.5625
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16330 | 765.4688
66 | 16330 | 765.4688
66 | 16338 | 765.8438
66 | 16337 | 765.7969
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16331 | 765.5156
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16334 | 765.6562
66 | 16329 | 765.4219
66 | 16338 | 765.8438
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16339 | 765.8906
66 | 16330 | 765.4688
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16331 | 765.5156
66 | 16327 | 765.3281
66 | 16335 | 765.7031
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16330 | 765.4688
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16338 | 765.8438
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16331 | 765.5156
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16336 | 765.75
66 | 16328 | 765.375
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16330 | 765.4688
66 | 16327 | 765.3281
66 | 16337 | 765.7969
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16333 | 765.6094
66 | 16330 | 765.4688
66 | 16338 | 765.8438
66 | 16339 | 765.8906
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16337 | 765.7969
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16332 | 765.5625
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16340 | 765.9375
66 | 16331 | 765.5156
66 | 16329 | 765.4219
66 | 16339 | 765.8906
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16340 | 765.9375
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16329 | 765.4219
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16341 | 765.9844
66 | 16329 | 765.4219
66 | 16331 | 765.5156
66 | 16337 | 765.7969
66 | 16338 | 765.8438
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16336 | 765.75
66 | 16337 | 765.7969
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16332 | 765.5625
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16337 | 765.7969
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16338 | 765.8438
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16332 | 765.5625
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16328 | 765.375
66 | 16327 | 765.3281
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16337 | 765.7969
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16339 | 765.8906
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16336 | 765.75
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16337 | 765.7969
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16339 | 765.8906
66 | 16329 | 765.4219
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16337 | 765.7969
66 | 16328 | 765.375
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16329 | 765.4219
66 | 16327 | 765.3281
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16337 | 765.7969
66 | 16327 | 765.3281
66 | 16326 | 765.2812
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16339 | 765.8906
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16334 | 765.6562
66 | 16338 | 765.8438
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16338 | 765.8438
66 | 16332 | 765.5625
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16329 | 765.4219
66 | 16330 | 765.4688
66 | 16337 | 765.7969
66 | 16338 | 765.8438
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16329 | 765.4219
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16339 | 765.8906
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16329 | 765.4219
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16337 | 765.7969
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16338 | 765.8438
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16330 | 765.4688
66 | 16336 | 765.75
66 | 16338 | 765.8438
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16335 | 765.7031
66 | 16329 | 765.4219
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16330 | 765.4688
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16328 | 765.375
66 | 16336 | 765.75
66 | 16337 | 765.7969
66 | 16327 | 765.3281
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16334 | 765.6562
66 | 16325 | 765.2344
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16330 | 765.4688
66 | 16327 | 765.3281
66 | 16335 | 765.7031
66 | 16338 | 765.8438
66 | 16329 | 765.4219
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16337 | 765.7969
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16335 | 765.7031
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16334 | 765.6562
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16328 | 765.375
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16336 | 765.75
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16329 | 765.4219
66 | 16325 | 765.2344
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16337 | 765.7969
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16327 | 765.3281
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16327 | 765.3281
66 | 16326 | 765.2812
66 | 16334 | 765.6562
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16327 | 765.3281
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16333 | 765.6094
66 | 16325 | 765.2344
66 | 16332 | 765.5625
66 | 16334 | 765.6562
66 | 16327 | 765.3281
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16333 | 765.6094
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16331 | 765.5156
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16337 | 765.7969
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16329 | 765.4219
66 | 16325 | 765.2344
66 | 16334 | 765.6562
66 | 16336 | 765.75
66 | 16329 | 765.4219
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16327 | 765.3281
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16335 | 765.7031
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16335 | 765.7031
66 | 16337 | 765.7969
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16323 | 765.1406
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16330 | 765.4688
66 | 16324 | 765.1875
66 | 16333 | 765.6094
66 | 16331 | 765.5156
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16331 | 765.5156
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16324 | 765.1875
66 | 16332 | 765.5625
66 | 16332 | 765.5625
66 | 16325 | 765.2344
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16334 | 765.6562
66 | 16334 | 765.6562
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16323 | 765.1406
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16335 | 765.7031
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16336 | 765.75
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16323 | 765.1406
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16330 | 765.4688
66 | 16323 | 765.1406
66 | 16332 | 765.5625
66 | 16334 | 765.6562
66 | 16328 | 765.375
66 | 16325 | 765.2344
66 | 16334 | 765.6562
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16329 | 765.4219
66 | 16323 | 765.1406
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16336 | 765.75
66 | 16325 | 765.2344
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16331 | 765.5156
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16333 | 765.6094
66 | 16327 | 765.3281
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16337 | 765.7969
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16330 | 765.4688
66 | 16321 | 765.0469
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16325 | 765.2344
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16329 | 765.4219
66 | 16335 | 765.7031
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16334 | 765.6562
66 | 16331 | 765.5156
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16324 | 765.1875
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16321 | 765.0469
66 | 16328 | 765.375
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16333 | 765.6094
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16335 | 765.7031
66 | 16326 | 765.2812
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16321 | 765.0469
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16328 | 765.375
66 | 16324 | 765.1875
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16323 | 765.1406
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16329 | 765.4219
66 | 16322 | 765.0938
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16330 | 765.4688
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16322 | 765.0938
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16328 | 765.375
66 | 16321 | 765.0469
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16332 | 765.5625
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16328 | 765.375
66 | 16320 | 765
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16334 | 765.6562
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16323 | 765.1406
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16332 | 765.5625
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16331 | 765.5156
66 | 16334 | 765.6562
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16321 | 765.0469
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16330 | 765.4688
66 | 16322 | 765.0938
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16327 | 765.3281
66 | 16323 | 765.1406
66 | 16332 | 765.5625
66 | 16334 | 765.6562
66 | 16325 | 765.2344
66 | 16326 | 765.2812
66 | 16333 | 765.6094
66 | 16328 | 765.375
66 | 16322 | 765.0938
66 | 16330 | 765.4688
66 | 16333 | 765.6094
66 | 16322 | 765.0938
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16329 | 765.4219
66 | 16324 | 765.1875
66 | 16332 | 765.5625
66 | 16335 | 765.7031
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16325 | 765.2344
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16327 | 765.3281
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16333 | 765.6094
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16330 | 765.4688
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16325 | 765.2344
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16333 | 765.6094
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16329 | 765.4219
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16327 | 765.3281
66 | 16322 | 765.0938
66 | 16331 | 765.5156
66 | 16332 | 765.5625
66 | 16323 | 765.1406
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16328 | 765.375
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16320 | 765
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16327 | 765.3281
66 | 16321 | 765.0469
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16322 | 765.0938
66 | 16322 | 765.0938
66 | 16330 | 765.4688
66 | 16329 | 765.4219
66 | 16320 | 765
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16324 | 765.1875
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16334 | 765.6562
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16333 | 765.6094
66 | 16323 | 765.1406
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16328 | 765.375
66 | 16319 | 764.9531
66 | 16328 | 765.375
66 | 16332 | 765.5625
66 | 16323 | 765.1406
66 | 16322 | 765.0938
66 | 16331 | 765.5156
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16334 | 765.6562
66 | 16323 | 765.1406
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16324 | 765.1875
66 | 16331 | 765.5156
66 | 16329 | 765.4219
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16326 | 765.2812
66 | 16331 | 765.5156
66 | 16322 | 765.0938
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16327 | 765.3281
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16325 | 765.2344
66 | 16331 | 765.5156
66 | 16331 | 765.5156
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16321 | 765.0469
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16325 | 765.2344
66 | 16332 | 765.5625
66 | 16330 | 765.4688
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16329 | 765.4219
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16333 | 765.6094
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16329 | 765.4219
66 | 16319 | 764.9531
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16327 | 765.3281
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16332 | 765.5625
66 | 16323 | 765.1406
66 | 16323 | 765.1406
66 | 16328 | 765.375
66 | 16332 | 765.5625
66 | 16322 | 765.0938
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16331 | 765.5156
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16328 | 765.375
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16328 | 765.375
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16320 | 765
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16330 | 765.4688
66 | 16319 | 764.9531
66 | 16323 | 765.1406
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16321 | 765.0469
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16332 | 765.5625
66 | 16320 | 765
66 | 16324 | 765.1875
66 | 16330 | 765.4688
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16329 | 765.4219
66 | 16320 | 765
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16325 | 765.2344
66 | 16320 | 765
66 | 16329 | 765.4219
66 | 16331 | 765.5156
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16330 | 765.4688
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16329 | 765.4219
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16331 | 765.5156
66 | 16321 | 765.0469
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16326 | 765.2812
66 | 16319 | 764.9531
66 | 16328 | 765.375
66 | 16331 | 765.5156
66 | 16319 | 764.9531
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16330 | 765.4688
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16328 | 765.375
66 | 16327 | 765.3281
66 | 16319 | 764.9531
66 | 16327 | 765.3281
66 | 16331 | 765.5156
66 | 16323 | 765.1406
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16329 | 765.4219
66 | 16320 | 765
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16324 | 765.1875
66 | 16319 | 764.9531
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16322 | 765.0938
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16330 | 765.4688
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16318 | 764.9062
66 | 16319 | 764.9531
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16325 | 765.2344
66 | 16320 | 765
66 | 16326 | 765.2812
66 | 16330 | 765.4688
66 | 16320 | 765
66 | 16321 | 765.0469
66 | 16329 | 765.4219
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16318 | 764.9062
66 | 16320 | 765
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16329 | 765.4219
66 | 16324 | 765.1875
66 | 16318 | 764.9062
66 | 16328 | 765.375
66 | 16330 | 765.4688
66 | 16321 | 765.0469
66 | 16322 | 765.0938
66 | 16330 | 765.4688
66 | 16324 | 765.1875
66 | 16319 | 764.9531
66 | 16326 | 765.2812
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16320 | 765
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16324 | 765.1875
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16327 | 765.3281
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16317 | 764.8594
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16322 | 765.0938
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16330 | 765.4688
66 | 16322 | 765.0938
66 | 16321 | 765.0469
66 | 16328 | 765.375
66 | 16326 | 765.2812
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16329 | 765.4219
66 | 16317 | 764.8594
66 | 16320 | 765
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16321 | 765.0469
66 | 16319 | 764.9531
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16319 | 764.9531
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16319 | 764.9531
66 | 16319 | 764.9531
66 | 16327 | 765.3281
66 | 16321 | 765.0469
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16330 | 765.4688
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16329 | 765.4219
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16328 | 765.375
66 | 16322 | 765.0938
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16317 | 764.8594
66 | 16320 | 765
66 | 16326 | 765.2812
66 | 16322 | 765.0938
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16328 | 765.375
66 | 16317 | 764.8594
66 | 16320 | 765
66 | 16329 | 765.4219
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16319 | 764.9531
66 | 16318 | 764.9062
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16329 | 765.4219
66 | 16321 | 765.0469
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16323 | 765.1406
66 | 16319 | 764.9531
66 | 16327 | 765.3281
66 | 16325 | 765.2344
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16328 | 765.375
66 | 16320 | 765
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16321 | 765.0469
66 | 16321 | 765.0469
66 | 16328 | 765.375
66 | 16326 | 765.2812
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16323 | 765.1406
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16323 | 765.1406
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16316 | 764.8125
66 | 16319 | 764.9531
66 | 16326 | 765.2812
66 | 16319 | 764.9531
66 | 16319 | 764.9531
66 | 16326 | 765.2812
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16326 | 765.2812
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16327 | 765.3281
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16321 | 765.0469
66 | 16318 | 764.9062
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16329 | 765.4219
66 | 16318 | 764.9062
66 | 16320 | 765
66 | 16326 | 765.2812
66 | 16322 | 765.0938
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16320 | 765
66 | 16318 | 764.9062
66 | 16327 | 765.3281
66 | 16323 | 765.1406
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16324 | 765.1875
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16315 | 764.7656
66 | 16319 | 764.9531
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16319 | 764.9531
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16323 | 765.1406
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16324 | 765.1875
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16326 | 765.2812
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16314 | 764.7188
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16317 | 764.8594
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16321 | 765.0469
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16324 | 765.1875
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16327 | 765.3281
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16314 | 764.7188
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16317 | 764.8594
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16321 | 765.0469
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16326 | 765.2812
66 | 16328 | 765.375
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16318 | 764.9062
66 | 16317 | 764.8594
66 | 16325 | 765.2344
66 | 16325 | 765.2344
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16324 | 765.1875
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16328 | 765.375
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16327 | 765.3281
66 | 16320 | 765
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16319 | 764.9531
66 | 16325 | 765.2344
66 | 16318 | 764.9062
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16323 | 765.1406
66 | 16315 | 764.7656
66 | 16324 | 765.1875
66 | 16326 | 765.2812
66 | 16315 | 764.7656
66 | 16322 | 765.0938
66 | 16326 | 765.2812
66 | 16314 | 764.7188
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16318 | 764.9062
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16324 | 765.1875
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16327 | 765.3281
66 | 16320 | 765
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16325 | 765.2344
66 | 16313 | 764.6719
66 | 16322 | 765.0938
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16327 | 765.3281
66 | 16317 | 764.8594
66 | 16323 | 765.1406
66 | 16325 | 765.2344
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16325 | 765.2344
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16325 | 765.2344
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16327 | 765.3281
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16325 | 765.2344
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16325 | 765.2344
66 | 16313 | 764.6719
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16314 | 764.7188
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16315 | 764.7656
66 | 16319 | 764.9531
66 | 16327 | 765.3281
66 | 16316 | 764.8125
66 | 16319 | 764.9531
66 | 16325 | 765.2344
66 | 16318 | 764.9062
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16317 | 764.8594
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16321 | 765.0469
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16319 | 764.9531
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16319 | 764.9531
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16317 | 764.8594
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16315 | 764.7656
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16316 | 764.8125
66 | 16315 | 764.7656
66 | 16324 | 765.1875
66 | 16317 | 764.8594
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16320 | 765
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16323 | 765.1406
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16313 | 764.6719
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16314 | 764.7188
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16317 | 764.8594
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16324 | 765.1875
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16313 | 764.6719
66 | 16317 | 764.8594
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16325 | 765.2344
66 | 16317 | 764.8594
66 | 16316 | 764.8125
66 | 16325 | 765.2344
66 | 16316 | 764.8125
66 | 16317 | 764.8594
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16321 | 765.0469
66 | 16318 | 764.9062
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16319 | 764.9531
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16323 | 765.1406
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16318 | 764.9062
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16322 | 765.0938
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16311 | 764.5781
66 | 16321 | 765.0469
66 | 16325 | 765.2344
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16326 | 765.2812
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16318 | 764.9062
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16315 | 764.7656
66 | 16314 | 764.7188
66 | 16321 | 765.0469
66 | 16316 | 764.8125
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16317 | 764.8594
66 | 16317 | 764.8594
66 | 16324 | 765.1875
66 | 16314 | 764.7188
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16316 | 764.8125
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16319 | 764.9531
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16316 | 764.8125
66 | 16312 | 764.625
66 | 16322 | 765.0938
66 | 16316 | 764.8125
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16316 | 764.8125
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16315 | 764.7656
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16316 | 764.8125
66 | 16315 | 764.7656
66 | 16324 | 765.1875
66 | 16315 | 764.7656
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16319 | 764.9531
66 | 16314 | 764.7188
66 | 16321 | 765.0469
66 | 16317 | 764.8594
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16317 | 764.8594
66 | 16313 | 764.6719
66 | 16320 | 765
66 | 16319 | 764.9531
66 | 16313 | 764.6719
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16324 | 765.1875
66 | 16314 | 764.7188
66 | 16321 | 765.0469
66 | 16324 | 765.1875
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16322 | 765.0938
66 | 16311 | 764.5781
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16311 | 764.5781
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16324 | 765.1875
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16325 | 765.2344
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16316 | 764.8125
66 | 16313 | 764.6719
66 | 16321 | 765.0469
66 | 16313 | 764.6719
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16311 | 764.5781
66 | 16320 | 765
66 | 16324 | 765.1875
66 | 16312 | 764.625
66 | 16322 | 765.0938
66 | 16322 | 765.0938
66 | 16313 | 764.6719
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16313 | 764.6719
66 | 16322 | 765.0938
66 | 16315 | 764.7656
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16315 | 764.7656
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16316 | 764.8125
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16318 | 764.9062
66 | 16315 | 764.7656
66 | 16322 | 765.0938
66 | 16314 | 764.7188
66 | 16315 | 764.7656
66 | 16322 | 765.0938
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16311 | 764.5781
66 | 16315 | 764.7656
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16319 | 764.9531
66 | 16324 | 765.1875
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16325 | 765.2344
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16324 | 765.1875
66 | 16313 | 764.6719
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16311 | 764.5781
66 | 16314 | 764.7188
66 | 16322 | 765.0938
66 | 16314 | 764.7188
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16316 | 764.8125
66 | 16324 | 765.1875
66 | 16312 | 764.625
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16324 | 765.1875
66 | 16313 | 764.6719
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16315 | 764.7656
66 | 16323 | 765.1406
66 | 16320 | 765
66 | 16316 | 764.8125
66 | 16321 | 765.0469
66 | 16314 | 764.7188
66 | 16313 | 764.6719
66 | 16321 | 765.0469
66 | 16316 | 764.8125
66 | 16312 | 764.625
66 | 16318 | 764.9062
66 | 16313 | 764.6719
66 | 16313 | 764.6719
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16323 | 765.1406
66 | 16311 | 764.5781
66 | 16320 | 765
66 | 16317 | 764.8594
66 | 16313 | 764.6719
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16322 | 765.0938
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16316 | 764.8125
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16321 | 765.0469
66 | 16313 | 764.6719
66 | 16323 | 765.1406
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16321 | 765.0469
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16317 | 764.8594
66 | 16311 | 764.5781
66 | 16319 | 764.9531
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16311 | 764.5781
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16319 | 764.9531
66 | 16313 | 764.6719
66 | 16320 | 765
66 | 16311 | 764.5781
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16309 | 764.4844
66 | 16318 | 764.9062
66 | 16320 | 765
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16313 | 764.6719
66 | 16313 | 764.6719
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16319 | 764.9531
66 | 16314 | 764.7188
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16311 | 764.5781
66 | 16316 | 764.8125
66 | 16319 | 764.9531
66 | 16311 | 764.5781
66 | 16322 | 765.0938
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16317 | 764.8594
66 | 16320 | 765
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16317 | 764.8594
66 | 16314 | 764.7188
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16317 | 764.8594
66 | 16317 | 764.8594
66 | 16313 | 764.6719
66 | 16323 | 765.1406
66 | 16313 | 764.6719
66 | 16321 | 765.0469
66 | 16320 | 765
66 | 16315 | 764.7656
66 | 16322 | 765.0938
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16317 | 764.8594
66 | 16319 | 764.9531
66 | 16311 | 764.5781
66 | 16320 | 765
66 | 16312 | 764.625
66 | 16314 | 764.7188
66 | 16323 | 765.1406
66 | 16312 | 764.625
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16316 | 764.8125
66 | 16322 | 765.0938
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16318 | 764.9062
66 | 16311 | 764.5781
66 | 16319 | 764.9531
66 | 16312 | 764.625
66 | 16312 | 764.625
66 | 16322 | 765.0938
66 | 16314 | 764.7188
66 | 16318 | 764.9062
66 | 16322 | 765.0938
66 | 16312 | 764.625
66 | 16317 | 764.8594
66 | 16322 | 765.0938
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16317 | 764.8594
66 | 16310 | 764.5312
66 | 16317 | 764.8594
66 | 16319 | 764.9531
66 | 16312 | 764.625
66 | 16318 | 764.9062
66 | 16314 | 764.7188
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16314 | 764.7188
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16311 | 764.5781
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16323 | 765.1406
66 | 16311 | 764.5781
66 | 16318 | 764.9062
66 | 16316 | 764.8125
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16311 | 764.5781
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16316 | 764.8125
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16314 | 764.7188
66 | 16311 | 764.5781
66 | 16321 | 765.0469
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16321 | 765.0469
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16321 | 765.0469
66 | 16312 | 764.625
66 | 16320 | 765
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16318 | 764.9062
66 | 16312 | 764.625
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16310 | 764.5312
66 | 16317 | 764.8594
66 | 16317 | 764.8594
66 | 16310 | 764.5312
66 | 16320 | 765
66 | 16315 | 764.7656
66 | 16314 | 764.7188
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16318 | 764.9062
66 | 16315 | 764.7656
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16315 | 764.7656
66 | 16311 | 764.5781
66 | 16318 | 764.9062
66 | 16316 | 764.8125
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16317 | 764.8594
66 | 16312 | 764.625
66 | 16311 | 764.5781
66 | 16320 | 765
66 | 16308 | 764.4375
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16318 | 764.9062
66 | 16319 | 764.9531
66 | 16313 | 764.6719
66 | 16320 | 765
66 | 16310 | 764.5312
66 | 16315 | 764.7656
66 | 16319 | 764.9531
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16310 | 764.5312
66 | 16309 | 764.4844
66 | 16318 | 764.9062
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16319 | 764.9531
66 | 16311 | 764.5781
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16313 | 764.6719
66 | 16321 | 765.0469
66 | 16309 | 764.4844
66 | 16318 | 764.9062
66 | 16319 | 764.9531
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16310 | 764.5312
66 | 16308 | 764.4375
66 | 16318 | 764.9062
66 | 16310 | 764.5312
66 | 16313 | 764.6719
66 | 16318 | 764.9062
66 | 16309 | 764.4844
66 | 16312 | 764.625
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16308 | 764.4375
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16318 | 764.9062
66 | 16316 | 764.8125
66 | 16310 | 764.5312
66 | 16317 | 764.8594
66 | 16310 | 764.5312
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16311 | 764.5781
66 | 16314 | 764.7188
66 | 16319 | 764.9531
66 | 16310 | 764.5312
66 | 16314 | 764.7188
66 | 16320 | 765
66 | 16311 | 764.5781
66 | 16315 | 764.7656
66 | 16320 | 765
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16314 | 764.7188
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16316 | 764.8125
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16309 | 764.4844
66 | 16314 | 764.7188
66 | 16321 | 765.0469
66 | 16311 | 764.5781
66 | 16318 | 764.9062
66 | 16319 | 764.9531
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16310 | 764.5312
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16309 | 764.4844
66 | 16313 | 764.6719
66 | 16319 | 764.9531
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16316 | 764.8125
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16312 | 764.625
66 | 16312 | 764.625
66 | 16318 | 764.9062
66 | 16310 | 764.5312
66 | 16310 | 764.5312
66 | 16319 | 764.9531
66 | 16308 | 764.4375
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16306 | 764.3438
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16314 | 764.7188
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16309 | 764.4844
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16312 | 764.625
66 | 16318 | 764.9062
66 | 16305 | 764.2969
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16310 | 764.5312
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16311 | 764.5781
66 | 16312 | 764.625
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16315 | 764.7656
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16317 | 764.8594
66 | 16307 | 764.3906
66 | 16311 | 764.5781
66 | 16317 | 764.8594
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16320 | 765
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16315 | 764.7656
66 | 16306 | 764.3438
66 | 16315 | 764.7656
66 | 16311 | 764.5781
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16309 | 764.4844
66 | 16317 | 764.8594
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16311 | 764.5781
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16311 | 764.5781
66 | 16309 | 764.4844
66 | 16317 | 764.8594
66 | 16308 | 764.4375
66 | 16312 | 764.625
66 | 16316 | 764.8125
66 | 16306 | 764.3438
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16306 | 764.3438
66 | 16317 | 764.8594
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16316 | 764.8125
66 | 16311 | 764.5781
66 | 16319 | 764.9531
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16312 | 764.625
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16312 | 764.625
66 | 16316 | 764.8125
66 | 16306 | 764.3438
66 | 16316 | 764.8125
66 | 16311 | 764.5781
66 | 16310 | 764.5312
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16312 | 764.625
66 | 16308 | 764.4375
66 | 16318 | 764.9062
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16314 | 764.7188
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16309 | 764.4844
66 | 16309 | 764.4844
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16311 | 764.5781
66 | 16312 | 764.625
66 | 16317 | 764.8594
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16308 | 764.4375
66 | 16311 | 764.5781
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16308 | 764.4375
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16309 | 764.4844
66 | 16309 | 764.4844
66 | 16318 | 764.9062
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16306 | 764.3438
66 | 16307 | 764.3906
66 | 16317 | 764.8594
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16308 | 764.4375
66 | 16308 | 764.4375
66 | 16319 | 764.9531
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16319 | 764.9531
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16311 | 764.5781
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16306 | 764.3438
66 | 16311 | 764.5781
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16318 | 764.9062
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16307 | 764.3906
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16311 | 764.5781
66 | 16308 | 764.4375
66 | 16318 | 764.9062
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16317 | 764.8594
66 | 16306 | 764.3438
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16315 | 764.7656
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16305 | 764.2969
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16311 | 764.5781
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16304 | 764.25
66 | 16312 | 764.625
66 | 16307 | 764.3906
66 | 16305 | 764.2969
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16306 | 764.3438
66 | 16313 | 764.6719
66 | 16310 | 764.5312
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16304 | 764.25
66 | 16314 | 764.7188
66 | 16312 | 764.625
66 | 16308 | 764.4375
66 | 16314 | 764.7188
66 | 16304 | 764.25
66 | 16312 | 764.625
66 | 16312 | 764.625
66 | 16305 | 764.2969
66 | 16316 | 764.8125
66 | 16309 | 764.4844
66 | 16313 | 764.6719
66 | 16318 | 764.9062
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16308 | 764.4375
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16308 | 764.4375
66 | 16309 | 764.4844
66 | 16317 | 764.8594
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16311 | 764.5781
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16308 | 764.4375
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16308 | 764.4375
66 | 16310 | 764.5312
66 | 16316 | 764.8125
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16312 | 764.625
66 | 16306 | 764.3438
66 | 16308 | 764.4375
66 | 16317 | 764.8594
66 | 16306 | 764.3438
66 | 16316 | 764.8125
66 | 16310 | 764.5312
66 | 16311 | 764.5781
66 | 16316 | 764.8125
66 | 16304 | 764.25
66 | 16311 | 764.5781
66 | 16311 | 764.5781
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16314 | 764.7188
66 | 16304 | 764.25
66 | 16312 | 764.625
66 | 16310 | 764.5312
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16317 | 764.8594
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16305 | 764.2969
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16314 | 764.7188
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16308 | 764.4375
66 | 16306 | 764.3438
66 | 16317 | 764.8594
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16314 | 764.7188
66 | 16308 | 764.4375
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16307 | 764.3906
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16308 | 764.4375
66 | 16309 | 764.4844
66 | 16316 | 764.8125
66 | 16304 | 764.25
66 | 16314 | 764.7188
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16303 | 764.2031
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16306 | 764.3438
66 | 16306 | 764.3438
66 | 16317 | 764.8594
66 | 16304 | 764.25
66 | 16315 | 764.7656
66 | 16311 | 764.5781
66 | 16311 | 764.5781
66 | 16315 | 764.7656
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16304 | 764.25
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16313 | 764.6719
66 | 16310 | 764.5312
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16317 | 764.8594
66 | 16305 | 764.2969
66 | 16314 | 764.7188
66 | 16311 | 764.5781
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16305 | 764.2969
66 | 16309 | 764.4844
66 | 16314 | 764.7188
66 | 16302 | 764.1562
66 | 16313 | 764.6719
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16308 | 764.4375
66 | 16307 | 764.3906
66 | 16316 | 764.8125
66 | 16305 | 764.2969
66 | 16311 | 764.5781
66 | 16313 | 764.6719
66 | 16305 | 764.2969
66 | 16312 | 764.625
66 | 16304 | 764.25
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16302 | 764.1562
66 | 16313 | 764.6719
66 | 16311 | 764.5781
66 | 16308 | 764.4375
66 | 16316 | 764.8125
66 | 16305 | 764.2969
66 | 16312 | 764.625
66 | 16306 | 764.3438
66 | 16306 | 764.3438
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16308 | 764.4375
66 | 16306 | 764.3438
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16304 | 764.25
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16314 | 764.7188
66 | 16308 | 764.4375
66 | 16311 | 764.5781
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16304 | 764.25
66 | 16307 | 764.3906
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16312 | 764.625
66 | 16304 | 764.25
66 | 16308 | 764.4375
66 | 16315 | 764.7656
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16306 | 764.3438
66 | 16308 | 764.4375
66 | 16313 | 764.6719
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16303 | 764.2031
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16303 | 764.2031
66 | 16313 | 764.6719
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16314 | 764.7188
66 | 16304 | 764.25
66 | 16312 | 764.625
66 | 16307 | 764.3906
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16308 | 764.4375
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16305 | 764.2969
66 | 16307 | 764.3906
66 | 16314 | 764.7188
66 | 16304 | 764.25
66 | 16313 | 764.6719
66 | 16309 | 764.4844
66 | 16308 | 764.4375
66 | 16313 | 764.6719
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16303 | 764.2031
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16305 | 764.2969
66 | 16309 | 764.4844
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16305 | 764.2969
66 | 16308 | 764.4375
66 | 16314 | 764.7188
66 | 16302 | 764.1562
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16308 | 764.4375
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16304 | 764.25
66 | 16306 | 764.3438
66 | 16315 | 764.7656
66 | 16305 | 764.2969
66 | 16313 | 764.6719
66 | 16308 | 764.4375
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16306 | 764.3438
66 | 16309 | 764.4844
66 | 16302 | 764.1562
66 | 16313 | 764.6719
66 | 16302 | 764.1562
66 | 16311 | 764.5781
66 | 16304 | 764.25
66 | 16311 | 764.5781
66 | 16311 | 764.5781
66 | 16308 | 764.4375
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16308 | 764.4375
66 | 16303 | 764.2031
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16311 | 764.5781
66 | 16302 | 764.1562
66 | 16312 | 764.625
66 | 16305 | 764.2969
66 | 16312 | 764.625
66 | 16305 | 764.2969
66 | 16310 | 764.5312
66 | 16313 | 764.6719
66 | 16305 | 764.2969
66 | 16312 | 764.625
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16307 | 764.3906
66 | 16311 | 764.5781
66 | 16303 | 764.2031
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16313 | 764.6719
66 | 16304 | 764.25
66 | 16309 | 764.4844
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16308 | 764.4375
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16313 | 764.6719
66 | 16301 | 764.1094
66 | 16312 | 764.625
66 | 16306 | 764.3438
66 | 16309 | 764.4844
66 | 16313 | 764.6719
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16302 | 764.1562
66 | 16312 | 764.625
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16304 | 764.25
66 | 16310 | 764.5312
66 | 16314 | 764.7188
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16301 | 764.1094
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16307 | 764.3906
66 | 16309 | 764.4844
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16313 | 764.6719
66 | 16302 | 764.1562
66 | 16312 | 764.625
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16309 | 764.4844
66 | 16306 | 764.3438
66 | 16312 | 764.625
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16302 | 764.1562
66 | 16303 | 764.2031
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16311 | 764.5781
66 | 16305 | 764.2969
66 | 16307 | 764.3906
66 | 16313 | 764.6719
66 | 16301 | 764.1094
66 | 16309 | 764.4844
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16305 | 764.2969
66 | 16305 | 764.2969
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16310 | 764.5312
66 | 16305 | 764.2969
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16304 | 764.25
66 | 16301 | 764.1094
66 | 16303 | 764.2031
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16302 | 764.1562
66 | 16311 | 764.5781
66 | 16305 | 764.2969
66 | 16309 | 764.4844
66 | 16310 | 764.5312
66 | 16306 | 764.3438
66 | 16310 | 764.5312
66 | 16302 | 764.1562
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16301 | 764.1094
66 | 16305 | 764.2969
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16312 | 764.625
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16304 | 764.25
66 | 16309 | 764.4844
66 | 16309 | 764.4844
66 | 16304 | 764.25
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16300 | 764.0625
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16311 | 764.5781
66 | 16301 | 764.1094
66 | 16311 | 764.5781
66 | 16304 | 764.25
66 | 16308 | 764.4375
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16309 | 764.4844
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16309 | 764.4844
66 | 16309 | 764.4844
66 | 16298 | 763.9688
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16311 | 764.5781
66 | 16302 | 764.1562
66 | 16311 | 764.5781
66 | 16303 | 764.2031
66 | 16310 | 764.5312
66 | 16304 | 764.25
66 | 16307 | 764.3906
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16302 | 764.1562
66 | 16303 | 764.2031
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16302 | 764.1562
66 | 16308 | 764.4375
66 | 16310 | 764.5312
66 | 16305 | 764.2969
66 | 16312 | 764.625
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16306 | 764.3438
66 | 16304 | 764.25
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16309 | 764.4844
66 | 16303 | 764.2031
66 | 16309 | 764.4844
66 | 16309 | 764.4844
66 | 16308 | 764.4375
66 | 16311 | 764.5781
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16306 | 764.3438
66 | 16299 | 764.0156
66 | 16306 | 764.3438
66 | 16310 | 764.5312
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16301 | 764.1094
66 | 16311 | 764.5781
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16305 | 764.2969
66 | 16306 | 764.3438
66 | 16302 | 764.1562
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16311 | 764.5781
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16300 | 764.0625
66 | 16304 | 764.25
66 | 16303 | 764.2031
66 | 16301 | 764.1094
66 | 16309 | 764.4844
66 | 16298 | 763.9688
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16306 | 764.3438
66 | 16303 | 764.2031
66 | 16303 | 764.2031
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16309 | 764.4844
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16298 | 763.9688
66 | 16303 | 764.2031
66 | 16303 | 764.2031
66 | 16301 | 764.1094
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16296 | 763.875
66 | 16309 | 764.4844
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16301 | 764.1094
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16295 | 763.8281
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16295 | 763.8281
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16309 | 764.4844
66 | 16296 | 763.875
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16309 | 764.4844
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16310 | 764.5312
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16309 | 764.4844
66 | 16298 | 763.9688
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16305 | 764.2969
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16299 | 764.0156
66 | 16305 | 764.2969
66 | 16300 | 764.0625
66 | 16304 | 764.25
66 | 16306 | 764.3438
66 | 16304 | 764.25
66 | 16308 | 764.4375
66 | 16302 | 764.1562
66 | 16310 | 764.5312
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16298 | 763.9688
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16298 | 763.9688
66 | 16306 | 764.3438
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16295 | 763.8281
66 | 16306 | 764.3438
66 | 16294 | 763.7812
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16309 | 764.4844
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16310 | 764.5312
66 | 16300 | 764.0625
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16308 | 764.4375
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16295 | 763.8281
66 | 16306 | 764.3438
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16294 | 763.7812
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16298 | 763.9688
66 | 16306 | 764.3438
66 | 16299 | 764.0156
66 | 16307 | 764.3906
66 | 16301 | 764.1094
66 | 16307 | 764.3906
66 | 16301 | 764.1094
66 | 16308 | 764.4375
66 | 16300 | 764.0625
66 | 16306 | 764.3438
66 | 16300 | 764.0625
66 | 16305 | 764.2969
66 | 16301 | 764.1094
66 | 16305 | 764.2969
66 | 16299 | 764.0156
66 | 16304 | 764.25
66 | 16301 | 764.1094
66 | 16304 | 764.25
66 | 16302 | 764.1562
66 | 16304 | 764.25
66 | 16303 | 764.2031
66 | 16305 | 764.2969
66 | 16304 | 764.25
66 | 16307 | 764.3906
66 | 16304 | 764.25
66 | 16307 | 764.3906
66 | 16301 | 764.1094
66 | 16305 | 764.2969
66 | 16298 | 763.9688
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16298 | 763.9688
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16303 | 764.2031
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16296 | 763.875
66 | 16307 | 764.3906
66 | 16297 | 763.9219
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16306 | 764.3438
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16299 | 764.0156
66 | 16306 | 764.3438
66 | 16301 | 764.1094
66 | 16302 | 764.1562
66 | 16301 | 764.1094
66 | 16295 | 763.8281
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16306 | 764.3438
66 | 16297 | 763.9219
66 | 16304 | 764.25
66 | 16300 | 764.0625
66 | 16302 | 764.1562
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16300 | 764.0625
66 | 16305 | 764.2969
66 | 16300 | 764.0625
66 | 16304 | 764.25
66 | 16300 | 764.0625
66 | 16302 | 764.1562
66 | 16300 | 764.0625
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16296 | 763.875
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16294 | 763.7812
66 | 16307 | 764.3906
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16298 | 763.9688
66 | 16303 | 764.2031
66 | 16300 | 764.0625
66 | 16296 | 763.875
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16306 | 764.3438
66 | 16298 | 763.9688
66 | 16304 | 764.25
66 | 16298 | 763.9688
66 | 16307 | 764.3906
66 | 16300 | 764.0625
66 | 16301 | 764.1094
66 | 16301 | 764.1094
66 | 16294 | 763.7812
66 | 16301 | 764.1094
66 | 16291 | 763.6406
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16299 | 764.0156
66 | 16301 | 764.1094
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16298 | 763.9688
66 | 16304 | 764.25
66 | 16300 | 764.0625
66 | 16301 | 764.1094
66 | 16302 | 764.1562
66 | 16298 | 763.9688
66 | 16305 | 764.2969
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16292 | 763.6875
66 | 16304 | 764.25
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16303 | 764.2031
66 | 16292 | 763.6875
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16298 | 763.9688
66 | 16297 | 763.9219
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16301 | 764.1094
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16296 | 763.875
66 | 16305 | 764.2969
66 | 16299 | 764.0156
66 | 16298 | 763.9688
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16305 | 764.2969
66 | 16293 | 763.7344
66 | 16305 | 764.2969
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16293 | 763.7344
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16301 | 764.1094
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16297 | 763.9219
66 | 16299 | 764.0156
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16297 | 763.9219
66 | 16300 | 764.0625
66 | 16292 | 763.6875
66 | 16305 | 764.2969
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16300 | 764.0625
66 | 16299 | 764.0156
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16292 | 763.6875
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16304 | 764.25
66 | 16298 | 763.9688
66 | 16296 | 763.875
66 | 16303 | 764.2031
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16296 | 763.875
66 | 16302 | 764.1562
66 | 16298 | 763.9688
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16292 | 763.6875
66 | 16304 | 764.25
66 | 16295 | 763.8281
66 | 16304 | 764.25
66 | 16296 | 763.875
66 | 16294 | 763.7812
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16299 | 764.0156
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16300 | 764.0625
66 | 16296 | 763.875
66 | 16304 | 764.25
66 | 16294 | 763.7812
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16294 | 763.7812
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16293 | 763.7344
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16305 | 764.2969
66 | 16299 | 764.0156
66 | 16304 | 764.25
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16303 | 764.2031
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16295 | 763.8281
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16301 | 764.1094
66 | 16295 | 763.8281
66 | 16302 | 764.1562
66 | 16296 | 763.875
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16300 | 764.0625
66 | 16296 | 763.875
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16300 | 764.0625
66 | 16297 | 763.9219
66 | 16296 | 763.875
66 | 16302 | 764.1562
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16303 | 764.2031
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16299 | 764.0156
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16294 | 763.7812
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16302 | 764.1562
66 | 16295 | 763.8281
66 | 16295 | 763.8281
66 | 16299 | 764.0156
66 | 16289 | 763.5469
66 | 16300 | 764.0625
66 | 16290 | 763.5938
66 | 16301 | 764.1094
66 | 16295 | 763.8281
66 | 16301 | 764.1094
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16303 | 764.2031
66 | 16296 | 763.875
66 | 16298 | 763.9688
66 | 16299 | 764.0156
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16290 | 763.5938
66 | 16301 | 764.1094
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16296 | 763.875
66 | 16300 | 764.0625
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16301 | 764.1094
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16302 | 764.1562
66 | 16292 | 763.6875
66 | 16304 | 764.25
66 | 16293 | 763.7344
66 | 16299 | 764.0156
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16296 | 763.875
66 | 16295 | 763.8281
66 | 16300 | 764.0625
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16299 | 764.0156
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16300 | 764.0625
66 | 16289 | 763.5469
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16290 | 763.5938
66 | 16298 | 763.9688
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16292 | 763.6875
66 | 16302 | 764.1562
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16300 | 764.0625
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16290 | 763.5938
66 | 16299 | 764.0156
66 | 16294 | 763.7812
66 | 16289 | 763.5469
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16293 | 763.7344
66 | 16301 | 764.1094
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16292 | 763.6875
66 | 16292 | 763.6875
66 | 16297 | 763.9219
66 | 16289 | 763.5469
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16300 | 764.0625
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16300 | 764.0625
66 | 16288 | 763.5
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16291 | 763.6406
66 | 16302 | 764.1562
66 | 16297 | 763.9219
66 | 16295 | 763.8281
66 | 16300 | 764.0625
66 | 16289 | 763.5469
66 | 16301 | 764.1094
66 | 16289 | 763.5469
66 | 16298 | 763.9688
66 | 16293 | 763.7344
66 | 16289 | 763.5469
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16297 | 763.9219
66 | 16298 | 763.9688
66 | 16301 | 764.1094
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16297 | 763.9219
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16298 | 763.9688
66 | 16290 | 763.5938
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16289 | 763.5469
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16292 | 763.6875
66 | 16301 | 764.1094
66 | 16298 | 763.9688
66 | 16289 | 763.5469
66 | 16300 | 764.0625
66 | 16289 | 763.5469
66 | 16296 | 763.875
66 | 16294 | 763.7812
66 | 16289 | 763.5469
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16290 | 763.5938
66 | 16301 | 764.1094
66 | 16295 | 763.8281
66 | 16291 | 763.6406
66 | 16301 | 764.1094
66 | 16293 | 763.7344
66 | 16298 | 763.9688
66 | 16298 | 763.9688
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16289 | 763.5469
66 | 16289 | 763.5469
66 | 16297 | 763.9219
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16300 | 764.0625
66 | 16290 | 763.5938
66 | 16299 | 764.0156
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16302 | 764.1562
66 | 16295 | 763.8281
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16290 | 763.5938
66 | 16292 | 763.6875
66 | 16297 | 763.9219
66 | 16288 | 763.5
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16299 | 764.0156
66 | 16289 | 763.5469
66 | 16302 | 764.1562
66 | 16294 | 763.7812
66 | 16292 | 763.6875
66 | 16299 | 764.0156
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16298 | 763.9688
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16294 | 763.7812
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16290 | 763.5938
66 | 16299 | 764.0156
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16302 | 764.1562
66 | 16293 | 763.7344
66 | 16290 | 763.5938
66 | 16298 | 763.9688
66 | 16289 | 763.5469
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16298 | 763.9688
66 | 16288 | 763.5
66 | 16299 | 764.0156
66 | 16293 | 763.7344
66 | 16289 | 763.5469
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16297 | 763.9219
66 | 16289 | 763.5469
66 | 16300 | 764.0625
66 | 16290 | 763.5938
66 | 16292 | 763.6875
66 | 16297 | 763.9219
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16293 | 763.7344
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16289 | 763.5469
66 | 16290 | 763.5938
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16291 | 763.6406
66 | 16296 | 763.875
66 | 16298 | 763.9688
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16288 | 763.5
66 | 16299 | 764.0156
66 | 16292 | 763.6875
66 | 16292 | 763.6875
66 | 16298 | 763.9688
66 | 16288 | 763.5
66 | 16291 | 763.6406
66 | 16296 | 763.875
66 | 16286 | 763.4062
66 | 16298 | 763.9688
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16301 | 764.1094
66 | 16294 | 763.7812
66 | 16289 | 763.5469
66 | 16299 | 764.0156
66 | 16295 | 763.8281
66 | 16289 | 763.5469
66 | 16298 | 763.9688
66 | 16292 | 763.6875
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16289 | 763.5469
66 | 16295 | 763.8281
66 | 16297 | 763.9219
66 | 16288 | 763.5
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16295 | 763.8281
66 | 16288 | 763.5
66 | 16300 | 764.0625
66 | 16291 | 763.6406
66 | 16289 | 763.5469
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16295 | 763.8281
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16292 | 763.6875
66 | 16299 | 764.0156
66 | 16288 | 763.5
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16287 | 763.4531
66 | 16296 | 763.875
66 | 16290 | 763.5938
66 | 16296 | 763.875
66 | 16295 | 763.8281
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16288 | 763.5
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16287 | 763.4531
66 | 16299 | 764.0156
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16298 | 763.9688
66 | 16288 | 763.5
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16287 | 763.4531
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16295 | 763.8281
66 | 16287 | 763.4531
66 | 16300 | 764.0625
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16300 | 764.0625
66 | 16293 | 763.7344
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16290 | 763.5938
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16300 | 764.0625
66 | 16294 | 763.7812
66 | 16288 | 763.5
66 | 16299 | 764.0156
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16299 | 764.0156
66 | 16289 | 763.5469
66 | 16290 | 763.5938
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16298 | 763.9688
66 | 16296 | 763.875
66 | 16286 | 763.4062
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16295 | 763.8281
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16294 | 763.7812
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16288 | 763.5
66 | 16298 | 763.9688
66 | 16296 | 763.875
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16295 | 763.8281
66 | 16289 | 763.5469
66 | 16284 | 763.3125
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16298 | 763.9688
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16293 | 763.7344
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16291 | 763.6406
66 | 16297 | 763.9219
66 | 16297 | 763.9219
66 | 16287 | 763.4531
66 | 16292 | 763.6875
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16295 | 763.8281
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16299 | 764.0156
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16294 | 763.7812
66 | 16293 | 763.7344
66 | 16282 | 763.2188
66 | 16296 | 763.875
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16294 | 763.7812
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16294 | 763.7812
66 | 16285 | 763.3594
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16289 | 763.5469
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16285 | 763.3594
66 | 16290 | 763.5938
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16285 | 763.3594
66 | 16297 | 763.9219
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16290 | 763.5938
66 | 16286 | 763.4062
66 | 16295 | 763.8281
66 | 16289 | 763.5469
66 | 16287 | 763.4531
66 | 16295 | 763.8281
66 | 16285 | 763.3594
66 | 16286 | 763.4062
66 | 16294 | 763.7812
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16295 | 763.8281
66 | 16284 | 763.3125
66 | 16288 | 763.5
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16292 | 763.6875
66 | 16289 | 763.5469
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16289 | 763.5469
66 | 16296 | 763.875
66 | 16288 | 763.5
66 | 16296 | 763.875
66 | 16295 | 763.8281
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16289 | 763.5469
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16290 | 763.5938
66 | 16287 | 763.4531
66 | 16296 | 763.875
66 | 16287 | 763.4531
66 | 16297 | 763.9219
66 | 16295 | 763.8281
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16288 | 763.5
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16297 | 763.9219
66 | 16288 | 763.5
66 | 16286 | 763.4062
66 | 16295 | 763.8281
66 | 16288 | 763.5
66 | 16296 | 763.875
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16297 | 763.9219
66 | 16292 | 763.6875
66 | 16287 | 763.4531
66 | 16298 | 763.9688
66 | 16292 | 763.6875
66 | 16288 | 763.5
66 | 16295 | 763.8281
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16297 | 763.9219
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16292 | 763.6875
66 | 16284 | 763.3125
66 | 16296 | 763.875
66 | 16294 | 763.7812
66 | 16284 | 763.3125
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16294 | 763.7812
66 | 16285 | 763.3594
66 | 16296 | 763.875
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16291 | 763.6406
66 | 16283 | 763.2656
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16283 | 763.2656
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16293 | 763.7344
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16286 | 763.4062
66 | 16296 | 763.875
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16293 | 763.7344
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16290 | 763.5938
66 | 16281 | 763.1719
66 | 16293 | 763.7344
66 | 16288 | 763.5
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16284 | 763.3125
66 | 16297 | 763.9219
66 | 16293 | 763.7344
66 | 16283 | 763.2656
66 | 16292 | 763.6875
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16293 | 763.7344
66 | 16292 | 763.6875
66 | 16283 | 763.2656
66 | 16295 | 763.8281
66 | 16290 | 763.5938
66 | 16283 | 763.2656
66 | 16295 | 763.8281
66 | 16292 | 763.6875
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16290 | 763.5938
66 | 16283 | 763.2656
66 | 16293 | 763.7344
66 | 16290 | 763.5938
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16294 | 763.7812
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16290 | 763.5938
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16291 | 763.6406
66 | 16293 | 763.7344
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16292 | 763.6875
66 | 16288 | 763.5
66 | 16287 | 763.4531
66 | 16295 | 763.8281
66 | 16288 | 763.5
66 | 16285 | 763.3594
66 | 16295 | 763.8281
66 | 16295 | 763.8281
66 | 16284 | 763.3125
66 | 16290 | 763.5938
66 | 16290 | 763.5938
66 | 16283 | 763.2656
66 | 16292 | 763.6875
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16286 | 763.4062
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16288 | 763.5
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16294 | 763.7812
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16293 | 763.7344
66 | 16293 | 763.7344
66 | 16284 | 763.3125
66 | 16295 | 763.8281
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16293 | 763.7344
66 | 16292 | 763.6875
66 | 16283 | 763.2656
66 | 16291 | 763.6406
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16294 | 763.7812
66 | 16294 | 763.7812
66 | 16287 | 763.4531
66 | 16283 | 763.2656
66 | 16293 | 763.7344
66 | 16292 | 763.6875
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16291 | 763.6406
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16293 | 763.7344
66 | 16292 | 763.6875
66 | 16284 | 763.3125
66 | 16292 | 763.6875
66 | 16290 | 763.5938
66 | 16282 | 763.2188
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16281 | 763.1719
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16293 | 763.7344
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16292 | 763.6875
66 | 16292 | 763.6875
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16285 | 763.3594
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16288 | 763.5
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16294 | 763.7812
66 | 16292 | 763.6875
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16290 | 763.5938
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16293 | 763.7344
66 | 16293 | 763.7344
66 | 16282 | 763.2188
66 | 16284 | 763.3125
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16280 | 763.125
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16289 | 763.5469
66 | 16280 | 763.125
66 | 16291 | 763.6406
66 | 16289 | 763.5469
66 | 16280 | 763.125
66 | 16286 | 763.4062
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16292 | 763.6875
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16294 | 763.7812
66 | 16291 | 763.6406
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16289 | 763.5469
66 | 16289 | 763.5469
66 | 16281 | 763.1719
66 | 16283 | 763.2656
66 | 16289 | 763.5469
66 | 16286 | 763.4062
66 | 16281 | 763.1719
66 | 16291 | 763.6406
66 | 16290 | 763.5938
66 | 16283 | 763.2656
66 | 16291 | 763.6406
66 | 16293 | 763.7344
66 | 16287 | 763.4531
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16289 | 763.5469
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16280 | 763.125
66 | 16291 | 763.6406
66 | 16290 | 763.5938
66 | 16283 | 763.2656
66 | 16286 | 763.4062
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16289 | 763.5469
66 | 16281 | 763.1719
66 | 16284 | 763.3125
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16280 | 763.125
66 | 16290 | 763.5938
66 | 16290 | 763.5938
66 | 16282 | 763.2188
66 | 16288 | 763.5
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16283 | 763.2656
66 | 16292 | 763.6875
66 | 16292 | 763.6875
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16289 | 763.5469
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16292 | 763.6875
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16293 | 763.7344
66 | 16293 | 763.7344
66 | 16286 | 763.4062
66 | 16281 | 763.1719
66 | 16292 | 763.6875
66 | 16290 | 763.5938
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16286 | 763.4062
66 | 16280 | 763.125
66 | 16292 | 763.6875
66 | 16288 | 763.5
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16292 | 763.6875
66 | 16282 | 763.2188
66 | 16284 | 763.3125
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16281 | 763.1719
66 | 16291 | 763.6406
66 | 16289 | 763.5469
66 | 16281 | 763.1719
66 | 16290 | 763.5938
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16288 | 763.5
66 | 16294 | 763.7812
66 | 16286 | 763.4062
66 | 16281 | 763.1719
66 | 16290 | 763.5938
66 | 16289 | 763.5469
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16290 | 763.5938
66 | 16289 | 763.5469
66 | 16280 | 763.125
66 | 16286 | 763.4062
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16291 | 763.6406
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16279 | 763.0781
66 | 16283 | 763.2656
66 | 16290 | 763.5938
66 | 16287 | 763.4531
66 | 16280 | 763.125
66 | 16287 | 763.4531
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16279 | 763.0781
66 | 16285 | 763.3594
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16278 | 763.0312
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16293 | 763.7344
66 | 16291 | 763.6406
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16292 | 763.6875
66 | 16291 | 763.6406
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16289 | 763.5469
66 | 16286 | 763.4062
66 | 16277 | 762.9844
66 | 16288 | 763.5
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16290 | 763.5938
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16278 | 763.0312
66 | 16281 | 763.1719
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16277 | 762.9844
66 | 16290 | 763.5938
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16291 | 763.6406
66 | 16292 | 763.6875
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16290 | 763.5938
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16291 | 763.6406
66 | 16282 | 763.2188
66 | 16278 | 763.0312
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16292 | 763.6875
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16289 | 763.5469
66 | 16286 | 763.4062
66 | 16278 | 763.0312
66 | 16288 | 763.5
66 | 16288 | 763.5
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16293 | 763.7344
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16291 | 763.6406
66 | 16290 | 763.5938
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16289 | 763.5469
66 | 16287 | 763.4531
66 | 16278 | 763.0312
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16288 | 763.5
66 | 16289 | 763.5469
66 | 16284 | 763.3125
66 | 16280 | 763.125
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16280 | 763.125
66 | 16286 | 763.4062
66 | 16288 | 763.5
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16288 | 763.5
66 | 16287 | 763.4531
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16284 | 763.3125
66 | 16289 | 763.5469
66 | 16287 | 763.4531
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16289 | 763.5469
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16290 | 763.5938
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16291 | 763.6406
66 | 16290 | 763.5938
66 | 16281 | 763.1719
66 | 16287 | 763.4531
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16277 | 762.9844
66 | 16285 | 763.3594
66 | 16287 | 763.4531
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16288 | 763.5
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16282 | 763.2188
66 | 16283 | 763.2656
66 | 16290 | 763.5938
66 | 16287 | 763.4531
66 | 16280 | 763.125
66 | 16285 | 763.3594
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16289 | 763.5469
66 | 16284 | 763.3125
66 | 16280 | 763.125
66 | 16288 | 763.5
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16281 | 763.1719
66 | 16289 | 763.5469
66 | 16290 | 763.5938
66 | 16286 | 763.4062
66 | 16279 | 763.0781
66 | 16283 | 763.2656
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16281 | 763.1719
66 | 16291 | 763.6406
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16289 | 763.5469
66 | 16291 | 763.6406
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16290 | 763.5938
66 | 16286 | 763.4062
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16286 | 763.4062
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16283 | 763.2656
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16290 | 763.5938
66 | 16290 | 763.5938
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16288 | 763.5
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16284 | 763.3125
66 | 16288 | 763.5
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16282 | 763.2188
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16279 | 763.0781
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16287 | 763.4531
66 | 16278 | 763.0312
66 | 16289 | 763.5469
66 | 16287 | 763.4531
66 | 16283 | 763.2656
66 | 16276 | 762.9375
66 | 16285 | 763.3594
66 | 16287 | 763.4531
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16290 | 763.5938
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16283 | 763.2656
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16285 | 763.3594
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16289 | 763.5469
66 | 16288 | 763.5
66 | 16279 | 763.0781
66 | 16281 | 763.1719
66 | 16288 | 763.5
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16276 | 762.9375
66 | 16280 | 763.125
66 | 16286 | 763.4062
66 | 16286 | 763.4062
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16288 | 763.5
66 | 16288 | 763.5
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16291 | 763.6406
66 | 16288 | 763.5
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16285 | 763.3594
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16278 | 763.0312
66 | 16282 | 763.2188
66 | 16288 | 763.5
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16285 | 763.3594
66 | 16291 | 763.6406
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16289 | 763.5469
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16284 | 763.3125
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16284 | 763.3125
66 | 16288 | 763.5
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16285 | 763.3594
66 | 16290 | 763.5938
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16283 | 763.2656
66 | 16289 | 763.5469
66 | 16286 | 763.4062
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16282 | 763.2188
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16280 | 763.125
66 | 16284 | 763.3125
66 | 16290 | 763.5938
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16275 | 762.8906
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16284 | 763.3125
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16287 | 763.4531
66 | 16278 | 763.0312
66 | 16281 | 763.1719
66 | 16288 | 763.5
66 | 16286 | 763.4062
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16287 | 763.4531
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16288 | 763.5
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16281 | 763.1719
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16284 | 763.3125
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16280 | 763.125
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16287 | 763.4531
66 | 16287 | 763.4531
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16284 | 763.3125
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16286 | 763.4062
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16284 | 763.3125
66 | 16286 | 763.4062
66 | 16283 | 763.2656
66 | 16277 | 762.9844
66 | 16282 | 763.2188
66 | 16288 | 763.5
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16280 | 763.125
66 | 16288 | 763.5
66 | 16284 | 763.3125
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16282 | 763.2188
66 | 16288 | 763.5
66 | 16285 | 763.3594
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16279 | 763.0781
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16282 | 763.2188
66 | 16288 | 763.5
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16286 | 763.4062
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16286 | 763.4062
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16286 | 763.4062
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16277 | 762.9844
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16286 | 763.4062
66 | 16286 | 763.4062
66 | 16283 | 763.2656
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16287 | 763.4531
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16281 | 763.1719
66 | 16289 | 763.5469
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16288 | 763.5
66 | 16286 | 763.4062
66 | 16283 | 763.2656
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16287 | 763.4531
66 | 16284 | 763.3125
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16285 | 763.3594
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16277 | 762.9844
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16277 | 762.9844
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16286 | 763.4062
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16285 | 763.3594
66 | 16285 | 763.3594
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16275 | 762.8906
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16282 | 763.2188
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16286 | 763.4062
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16286 | 763.4062
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16282 | 763.2188
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16285 | 763.3594
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16283 | 763.2656
66 | 16287 | 763.4531
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16279 | 763.0781
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16286 | 763.4062
66 | 16283 | 763.2656
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16282 | 763.2188
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16272 | 762.75
66 | 16280 | 763.125
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16279 | 763.0781
66 | 16287 | 763.4531
66 | 16285 | 763.3594
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16282 | 763.2188
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16281 | 763.1719
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16278 | 763.0312
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16280 | 763.125
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16284 | 763.3125
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16279 | 763.0781
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16282 | 763.2188
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16282 | 763.2188
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16284 | 763.3125
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16279 | 763.0781
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16282 | 763.2188
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16282 | 763.2188
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16283 | 763.2656
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16278 | 763.0312
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16281 | 763.1719
66 | 16285 | 763.3594
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16275 | 762.8906
66 | 16283 | 763.2656
66 | 16284 | 763.3125
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16281 | 763.1719
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16275 | 762.8906
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16279 | 763.0781
66 | 16274 | 762.8438
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16275 | 762.8906
66 | 16272 | 762.75
66 | 16277 | 762.9844
66 | 16286 | 763.4062
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16277 | 762.9844
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16286 | 763.4062
66 | 16285 | 763.3594
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16280 | 763.125
66 | 16285 | 763.3594
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16283 | 763.2656
66 | 16285 | 763.3594
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16275 | 762.8906
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16280 | 763.125
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16284 | 763.3125
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16278 | 763.0312
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16281 | 763.1719
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16281 | 763.1719
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16272 | 762.75
66 | 16281 | 763.1719
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16276 | 762.9375
66 | 16283 | 763.2656
66 | 16284 | 763.3125
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16274 | 762.8438
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16279 | 763.0781
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16279 | 763.0781
66 | 16284 | 763.3125
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16278 | 763.0312
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16276 | 762.9375
66 | 16283 | 763.2656
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16279 | 763.0781
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16278 | 763.0312
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16277 | 762.9844
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16277 | 762.9844
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16274 | 762.8438
66 | 16280 | 763.125
66 | 16283 | 763.2656
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16274 | 762.8438
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16275 | 762.8906
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16274 | 762.8438
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16272 | 762.75
66 | 16280 | 763.125
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16271 | 762.7031
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16273 | 762.7969
66 | 16279 | 763.0781
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16282 | 763.2188
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16278 | 763.0312
66 | 16283 | 763.2656
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16273 | 762.7969
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16277 | 762.9844
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16275 | 762.8906
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16274 | 762.8438
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16271 | 762.7031
66 | 16275 | 762.8906
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16282 | 763.2188
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16273 | 762.7969
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16266 | 762.4688
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16272 | 762.75
66 | 16277 | 762.9844
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16274 | 762.8438
66 | 16279 | 763.0781
66 | 16281 | 763.1719
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16281 | 763.1719
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16272 | 762.75
66 | 16278 | 763.0312
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16268 | 762.5625
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16273 | 762.7969
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16279 | 763.0781
66 | 16280 | 763.125
66 | 16279 | 763.0781
66 | 16279 | 763.0781
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16279 | 763.0781
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16272 | 762.75
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16278 | 763.0312
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16278 | 763.0312
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16278 | 763.0312
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16277 | 762.9844
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16277 | 762.9844
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16276 | 762.9375
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16276 | 762.9375
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16275 | 762.8906
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16275 | 762.8906
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16264 | 762.375
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16266 | 762.4688
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16267 | 762.5156
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16272 | 762.75
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16274 | 762.8438
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16268 | 762.5625
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16267 | 762.5156
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16264 | 762.375
66 | 16262 | 762.2812
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16273 | 762.7969
66 | 16268 | 762.5625
66 | 16264 | 762.375
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16266 | 762.4688
66 | 16263 | 762.3281
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16273 | 762.7969
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16265 | 762.4219
66 | 16263 | 762.3281
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16270 | 762.6562
66 | 16263 | 762.3281
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16271 | 762.7031
66 | 16268 | 762.5625
66 | 16262 | 762.2812
66 | 16260 | 762.1875
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16272 | 762.75
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16260 | 762.1875
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16266 | 762.4688
66 | 16263 | 762.3281
66 | 16260 | 762.1875
66 | 16260 | 762.1875
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16264 | 762.375
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16257 | 762.0469
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16272 | 762.75
66 | 16270 | 762.6562
66 | 16266 | 762.4688
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16269 | 762.6094
66 | 16266 | 762.4688
66 | 16263 | 762.3281
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16261 | 762.2344
66 | 16258 | 762.0938
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16271 | 762.7031
66 | 16268 | 762.5625
66 | 16263 | 762.3281
66 | 16260 | 762.1875
66 | 16258 | 762.0938
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16268 | 762.5625
66 | 16262 | 762.2812
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16265 | 762.4219
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16264 | 762.375
66 | 16261 | 762.2344
66 | 16258 | 762.0938
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16262 | 762.2812
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16267 | 762.5156
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16270 | 762.6562
66 | 16266 | 762.4688
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16270 | 762.6562
66 | 16265 | 762.4219
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16263 | 762.3281
66 | 16259 | 762.1406
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16259 | 762.1406
66 | 16264 | 762.375
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16271 | 762.7031
66 | 16269 | 762.6094
66 | 16263 | 762.3281
66 | 16260 | 762.1875
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16263 | 762.3281
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16266 | 762.4688
66 | 16261 | 762.2344
66 | 16258 | 762.0938
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16265 | 762.4219
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16263 | 762.3281
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16260 | 762.1875
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16269 | 762.6094
66 | 16264 | 762.375
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16268 | 762.5625
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16265 | 762.4219
66 | 16258 | 762.0938
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16269 | 762.6094
66 | 16269 | 762.6094
66 | 16271 | 762.7031
66 | 16265 | 762.4219
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16259 | 762.1406
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16266 | 762.4688
66 | 16261 | 762.2344
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16260 | 762.1875
66 | 16258 | 762.0938
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16261 | 762.2344
66 | 16258 | 762.0938
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16269 | 762.6094
66 | 16267 | 762.5156
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16265 | 762.4219
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16270 | 762.6562
66 | 16263 | 762.3281
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16265 | 762.4219
66 | 16259 | 762.1406
66 | 16258 | 762.0938
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16256 | 762
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16259 | 762.1406
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16269 | 762.6094
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16260 | 762.1875
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16260 | 762.1875
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16260 | 762.1875
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16258 | 762.0938
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16257 | 762.0469
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16262 | 762.2812
66 | 16256 | 762
66 | 16256 | 762
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16256 | 762
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16268 | 762.5625
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16259 | 762.1406
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16257 | 762.0469
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16259 | 762.1406
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16257 | 762.0469
66 | 16253 | 761.8594
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16267 | 762.5156
66 | 16264 | 762.375
66 | 16259 | 762.1406
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16265 | 762.4219
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16262 | 762.2812
66 | 16258 | 762.0938
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16264 | 762.375
66 | 16264 | 762.375
66 | 16258 | 762.0938
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16259 | 762.1406
66 | 16256 | 762
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16266 | 762.4688
66 | 16262 | 762.2812
66 | 16255 | 761.9531
66 | 16254 | 761.9062
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16256 | 762
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16264 | 762.375
66 | 16257 | 762.0469
66 | 16256 | 762
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16262 | 762.2812
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16267 | 762.5156
66 | 16263 | 762.3281
66 | 16258 | 762.0938
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16263 | 762.3281
66 | 16256 | 762
66 | 16255 | 761.9531
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16266 | 762.4688
66 | 16259 | 762.1406
66 | 16255 | 761.9531
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16268 | 762.5625
66 | 16258 | 762.0938
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16259 | 762.1406
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16263 | 762.3281
66 | 16267 | 762.5156
66 | 16268 | 762.5625
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16268 | 762.5625
66 | 16267 | 762.5156
66 | 16259 | 762.1406
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16255 | 761.9531
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16269 | 762.6094
66 | 16257 | 762.0469
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16258 | 762.0938
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16256 | 762
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16259 | 762.1406
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16259 | 762.1406
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16267 | 762.5156
66 | 16257 | 762.0469
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16261 | 762.2344
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16267 | 762.5156
66 | 16259 | 762.1406
66 | 16256 | 762
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16256 | 762
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16268 | 762.5625
66 | 16258 | 762.0938
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16260 | 762.1875
66 | 16255 | 761.9531
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16268 | 762.5625
66 | 16260 | 762.1875
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16263 | 762.3281
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16257 | 762.0469
66 | 16253 | 761.8594
66 | 16254 | 761.9062
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16264 | 762.375
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16261 | 762.2344
66 | 16255 | 761.9531
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16261 | 762.2344
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16268 | 762.5625
66 | 16258 | 762.0938
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16264 | 762.375
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16263 | 762.3281
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16255 | 761.9531
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16258 | 762.0938
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16265 | 762.4219
66 | 16266 | 762.4688
66 | 16258 | 762.0938
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16257 | 762.0469
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16254 | 761.9062
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16258 | 762.0938
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16256 | 762
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16259 | 762.1406
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16256 | 762
66 | 16254 | 761.9062
66 | 16259 | 762.1406
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16258 | 762.0938
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16265 | 762.4219
66 | 16264 | 762.375
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16254 | 761.9062
66 | 16252 | 761.8125
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16255 | 761.9531
66 | 16251 | 761.7656
66 | 16259 | 762.1406
66 | 16265 | 762.4219
66 | 16263 | 762.3281
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16264 | 762.375
66 | 16264 | 762.375
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16253 | 761.8594
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16257 | 762.0469
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16258 | 762.0938
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16255 | 761.9531
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16266 | 762.4688
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16264 | 762.375
66 | 16265 | 762.4219
66 | 16257 | 762.0469
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16251 | 761.7656
66 | 16254 | 761.9062
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16255 | 761.9531
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16265 | 762.4219
66 | 16265 | 762.4219
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16251 | 761.7656
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16254 | 761.9062
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16266 | 762.4688
66 | 16256 | 762
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16261 | 762.2344
66 | 16259 | 762.1406
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16256 | 762
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16256 | 762
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16263 | 762.3281
66 | 16266 | 762.4688
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16262 | 762.2812
66 | 16265 | 762.4219
66 | 16252 | 761.8125
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16253 | 761.8594
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16263 | 762.3281
66 | 16259 | 762.1406
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16260 | 762.1875
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16256 | 762
66 | 16251 | 761.7656
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16255 | 761.9531
66 | 16252 | 761.8125
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16255 | 761.9531
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16265 | 762.4219
66 | 16255 | 761.9531
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16264 | 762.375
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16262 | 762.2812
66 | 16263 | 762.3281
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16256 | 762
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16264 | 762.375
66 | 16253 | 761.8594
66 | 16253 | 761.8594
66 | 16262 | 762.2812
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16257 | 762.0469
66 | 16249 | 761.6719
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16254 | 761.9062
66 | 16251 | 761.7656
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16253 | 761.8594
66 | 16250 | 761.7188
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16262 | 762.2812
66 | 16264 | 762.375
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16263 | 762.3281
66 | 16263 | 762.3281
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16255 | 761.9531
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16263 | 762.3281
66 | 16253 | 761.8594
66 | 16254 | 761.9062
66 | 16263 | 762.3281
66 | 16265 | 762.4219
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16251 | 761.7656
66 | 16253 | 761.8594
66 | 16261 | 762.2344
66 | 16257 | 762.0469
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16261 | 762.2344
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16260 | 762.1875
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16255 | 761.9531
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16262 | 762.2812
66 | 16259 | 762.1406
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16258 | 762.0938
66 | 16248 | 761.625
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16250 | 761.7188
66 | 16257 | 762.0469
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16262 | 762.2812
66 | 16259 | 762.1406
66 | 16253 | 761.8594
66 | 16261 | 762.2344
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16250 | 761.7188
66 | 16251 | 761.7656
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16263 | 762.3281
66 | 16255 | 761.9531
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16250 | 761.7188
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16255 | 761.9531
66 | 16250 | 761.7188
66 | 16259 | 762.1406
66 | 16263 | 762.3281
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16263 | 762.3281
66 | 16254 | 761.9062
66 | 16252 | 761.8125
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16261 | 762.2344
66 | 16249 | 761.6719
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16263 | 762.3281
66 | 16251 | 761.7656
66 | 16254 | 761.9062
66 | 16262 | 762.2812
66 | 16254 | 761.9062
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16250 | 761.7188
66 | 16251 | 761.7656
66 | 16260 | 762.1875
66 | 16258 | 762.0938
66 | 16250 | 761.7188
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16253 | 761.8594
66 | 16252 | 761.8125
66 | 16260 | 762.1875
66 | 16263 | 762.3281
66 | 16251 | 761.7656
66 | 16254 | 761.9062
66 | 16260 | 762.1875
66 | 16256 | 762
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16260 | 762.1875
66 | 16249 | 761.6719
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16251 | 761.7656
66 | 16249 | 761.6719
66 | 16260 | 762.1875
66 | 16262 | 762.2812
66 | 16251 | 761.7656
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16253 | 761.8594
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16247 | 761.5781
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16259 | 762.1406
66 | 16262 | 762.2812
66 | 16249 | 761.6719
66 | 16256 | 762
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16261 | 762.2344
66 | 16260 | 762.1875
66 | 16250 | 761.7188
66 | 16254 | 761.9062
66 | 16260 | 762.1875
66 | 16248 | 761.625
66 | 16249 | 761.6719
66 | 16258 | 762.0938
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16261 | 762.2344
66 | 16250 | 761.7188
66 | 16257 | 762.0469
66 | 16261 | 762.2344
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16254 | 761.9062
66 | 16248 | 761.625
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16251 | 761.7656
66 | 16250 | 761.7188
66 | 16260 | 762.1875
66 | 16261 | 762.2344
66 | 16251 | 761.7656
66 | 16255 | 761.9531
66 | 16261 | 762.2344
66 | 16253 | 761.8594
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16254 | 761.9062
66 | 16248 | 761.625
66 | 16256 | 762
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16256 | 762
66 | 16250 | 761.7188
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16253 | 761.8594
66 | 16247 | 761.5781
66 | 16256 | 762
66 | 16260 | 762.1875
66 | 16248 | 761.625
66 | 16252 | 761.8125
66 | 16260 | 762.1875
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16255 | 761.9531
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16259 | 762.1406
66 | 16248 | 761.625
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16254 | 761.9062
66 | 16249 | 761.6719
66 | 16259 | 762.1406
66 | 16261 | 762.2344
66 | 16249 | 761.6719
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16252 | 761.8125
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16254 | 761.9062
66 | 16246 | 761.5312
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16259 | 762.1406
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16259 | 762.1406
66 | 16249 | 761.6719
66 | 16251 | 761.7656
66 | 16259 | 762.1406
66 | 16249 | 761.6719
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16252 | 761.8125
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16250 | 761.7188
66 | 16252 | 761.8125
66 | 16261 | 762.2344
66 | 16253 | 761.8594
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16259 | 762.1406
66 | 16249 | 761.6719
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16250 | 761.7188
66 | 16248 | 761.625
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16249 | 761.6719
66 | 16251 | 761.7656
66 | 16261 | 762.2344
66 | 16258 | 762.0938
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16262 | 762.2812
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16258 | 762.0938
66 | 16249 | 761.6719
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16250 | 761.7188
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16249 | 761.6719
66 | 16254 | 761.9062
66 | 16261 | 762.2344
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16252 | 761.8125
66 | 16250 | 761.7188
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16247 | 761.5781
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16248 | 761.625
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16251 | 761.7656
66 | 16252 | 761.8125
66 | 16260 | 762.1875
66 | 16257 | 762.0469
66 | 16248 | 761.625
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16249 | 761.6719
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16257 | 762.0469
66 | 16249 | 761.6719
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16253 | 761.8594
66 | 16249 | 761.6719
66 | 16259 | 762.1406
66 | 16260 | 762.1875
66 | 16248 | 761.625
66 | 16252 | 761.8125
66 | 16258 | 762.0938
66 | 16251 | 761.7656
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16248 | 761.625
66 | 16247 | 761.5781
66 | 16258 | 762.0938
66 | 16260 | 762.1875
66 | 16247 | 761.5781
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16254 | 761.9062
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16247 | 761.5781
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16251 | 761.7656
66 | 16248 | 761.625
66 | 16257 | 762.0469
66 | 16260 | 762.1875
66 | 16248 | 761.625
66 | 16253 | 761.8594
66 | 16260 | 762.1875
66 | 16249 | 761.6719
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16249 | 761.6719
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16249 | 761.6719
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16247 | 761.5781
66 | 16255 | 761.9531
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16259 | 762.1406
66 | 16248 | 761.625
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16248 | 761.625
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16248 | 761.625
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16246 | 761.5312
66 | 16251 | 761.7656
66 | 16258 | 762.0938
66 | 16249 | 761.6719
66 | 16255 | 761.9531
66 | 16259 | 762.1406
66 | 16248 | 761.625
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16247 | 761.5781
66 | 16255 | 761.9531
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16249 | 761.6719
66 | 16255 | 761.9531
66 | 16260 | 762.1875
66 | 16247 | 761.5781
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16248 | 761.625
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16246 | 761.5312
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16259 | 762.1406
66 | 16248 | 761.625
66 | 16256 | 762
66 | 16258 | 762.0938
66 | 16248 | 761.625
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16247 | 761.5781
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16246 | 761.5312
66 | 16256 | 762
66 | 16258 | 762.0938
66 | 16247 | 761.5781
66 | 16253 | 761.8594
66 | 16258 | 762.0938
66 | 16247 | 761.5781
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16244 | 761.4375
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16248 | 761.625
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16249 | 761.6719
66 | 16249 | 761.6719
66 | 16257 | 762.0469
66 | 16248 | 761.625
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16245 | 761.4844
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16243 | 761.3906
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16248 | 761.625
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16246 | 761.5312
66 | 16251 | 761.7656
66 | 16257 | 762.0469
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16247 | 761.5781
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16243 | 761.3906
66 | 16249 | 761.6719
66 | 16256 | 762
66 | 16244 | 761.4375
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16252 | 761.8125
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16250 | 761.7188
66 | 16254 | 761.9062
66 | 16243 | 761.3906
66 | 16249 | 761.6719
66 | 16255 | 761.9531
66 | 16245 | 761.4844
66 | 16251 | 761.7656
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16243 | 761.3906
66 | 16250 | 761.7188
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16258 | 762.0938
66 | 16247 | 761.5781
66 | 16253 | 761.8594
66 | 16255 | 761.9531
66 | 16246 | 761.5312
66 | 16255 | 761.9531
66 | 16257 | 762.0469
66 | 16244 | 761.4375
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16244 | 761.4375
66 | 16251 | 761.7656
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16255 | 761.9531
66 | 16256 | 762
66 | 16244 | 761.4375
66 | 16255 | 761.9531
66 | 16254 | 761.9062
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16254 | 761.9062
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16251 | 761.7656
66 | 16244 | 761.4375
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16244 | 761.4375
66 | 16254 | 761.9062
66 | 16252 | 761.8125
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16254 | 761.9062
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16253 | 761.8594
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16250 | 761.7188
66 | 16244 | 761.4375
66 | 16254 | 761.9062
66 | 16253 | 761.8594
66 | 16244 | 761.4375
66 | 16255 | 761.9531
66 | 16253 | 761.8594
66 | 16246 | 761.5312
66 | 16255 | 761.9531
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16253 | 761.8594
66 | 16254 | 761.9062
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16257 | 762.0469
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16257 | 762.0469
66 | 16246 | 761.5312
66 | 16253 | 761.8594
66 | 16254 | 761.9062
66 | 16244 | 761.4375
66 | 16254 | 761.9062
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16251 | 761.7656
66 | 16254 | 761.9062
66 | 16244 | 761.4375
66 | 16254 | 761.9062
66 | 16256 | 762
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16247 | 761.5781
66 | 16255 | 761.9531
66 | 16258 | 762.0938
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16243 | 761.3906
66 | 16253 | 761.8594
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16253 | 761.8594
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16243 | 761.3906
66 | 16250 | 761.7188
66 | 16252 | 761.8125
66 | 16244 | 761.4375
66 | 16251 | 761.7656
66 | 16254 | 761.9062
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16246 | 761.5312
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16242 | 761.3438
66 | 16253 | 761.8594
66 | 16254 | 761.9062
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16256 | 762
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16254 | 761.9062
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16253 | 761.8594
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16253 | 761.8594
66 | 16241 | 761.2969
66 | 16251 | 761.7656
66 | 16255 | 761.9531
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16255 | 761.9531
66 | 16243 | 761.3906
66 | 16253 | 761.8594
66 | 16255 | 761.9531
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16253 | 761.8594
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16249 | 761.6719
66 | 16242 | 761.3438
66 | 16253 | 761.8594
66 | 16253 | 761.8594
66 | 16242 | 761.3438
66 | 16253 | 761.8594
66 | 16249 | 761.6719
66 | 16244 | 761.4375
66 | 16254 | 761.9062
66 | 16251 | 761.7656
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16245 | 761.4844
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16244 | 761.4375
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16254 | 761.9062
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16249 | 761.6719
66 | 16243 | 761.3906
66 | 16253 | 761.8594
66 | 16247 | 761.5781
66 | 16245 | 761.4844
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16247 | 761.5781
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16251 | 761.7656
66 | 16252 | 761.8125
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16251 | 761.7656
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16252 | 761.8125
66 | 16243 | 761.3906
66 | 16247 | 761.5781
66 | 16252 | 761.8125
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16251 | 761.7656
66 | 16249 | 761.6719
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16246 | 761.5312
66 | 16254 | 761.9062
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16248 | 761.625
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16243 | 761.3906
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16249 | 761.6719
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16247 | 761.5781
66 | 16242 | 761.3438
66 | 16252 | 761.8125
66 | 16250 | 761.7188
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16250 | 761.7188
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16251 | 761.7656
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16246 | 761.5312
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16245 | 761.4844
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16249 | 761.6719
66 | 16241 | 761.2969
66 | 16252 | 761.8125
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16248 | 761.625
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16249 | 761.6719
66 | 16241 | 761.2969
66 | 16251 | 761.7656
66 | 16249 | 761.6719
66 | 16244 | 761.4375
66 | 16251 | 761.7656
66 | 16244 | 761.4375
66 | 16245 | 761.4844
66 | 16254 | 761.9062
66 | 16242 | 761.3438
66 | 16251 | 761.7656
66 | 16253 | 761.8594
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16250 | 761.7188
66 | 16241 | 761.2969
66 | 16248 | 761.625
66 | 16250 | 761.7188
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16252 | 761.8125
66 | 16245 | 761.4844
66 | 16247 | 761.5781
66 | 16254 | 761.9062
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16248 | 761.625
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16250 | 761.7188
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16241 | 761.2969
66 | 16251 | 761.7656
66 | 16249 | 761.6719
66 | 16244 | 761.4375
66 | 16252 | 761.8125
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16250 | 761.7188
66 | 16244 | 761.4375
66 | 16246 | 761.5312
66 | 16251 | 761.7656
66 | 16241 | 761.2969
66 | 16248 | 761.625
66 | 16241 | 761.2969
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16248 | 761.625
66 | 16244 | 761.4375
66 | 16251 | 761.7656
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16253 | 761.8594
66 | 16243 | 761.3906
66 | 16250 | 761.7188
66 | 16251 | 761.7656
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16251 | 761.7656
66 | 16239 | 761.2031
66 | 16249 | 761.6719
66 | 16248 | 761.625
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16251 | 761.7656
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16253 | 761.8594
66 | 16242 | 761.3438
66 | 16250 | 761.7188
66 | 16244 | 761.4375
66 | 16243 | 761.3906
66 | 16252 | 761.8125
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16243 | 761.3906
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16241 | 761.2969
66 | 16248 | 761.625
66 | 16251 | 761.7656
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16251 | 761.7656
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16247 | 761.5781
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16243 | 761.3906
66 | 16250 | 761.7188
66 | 16239 | 761.2031
66 | 16249 | 761.6719
66 | 16247 | 761.5781
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16241 | 761.2969
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16251 | 761.7656
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16244 | 761.4375
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16247 | 761.5781
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16248 | 761.625
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16247 | 761.5781
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16242 | 761.3438
66 | 16240 | 761.25
66 | 16251 | 761.7656
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16249 | 761.6719
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16251 | 761.7656
66 | 16245 | 761.4844
66 | 16244 | 761.4375
66 | 16250 | 761.7188
66 | 16241 | 761.2969
66 | 16247 | 761.5781
66 | 16251 | 761.7656
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16250 | 761.7188
66 | 16238 | 761.1562
66 | 16247 | 761.5781
66 | 16245 | 761.4844
66 | 16239 | 761.2031
66 | 16248 | 761.625
66 | 16244 | 761.4375
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16252 | 761.8125
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16250 | 761.7188
66 | 16241 | 761.2969
66 | 16248 | 761.625
66 | 16245 | 761.4844
66 | 16239 | 761.2031
66 | 16249 | 761.6719
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16247 | 761.5781
66 | 16248 | 761.625
66 | 16239 | 761.2031
66 | 16247 | 761.5781
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16247 | 761.5781
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16241 | 761.2969
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16246 | 761.5312
66 | 16239 | 761.2031
66 | 16248 | 761.625
66 | 16239 | 761.2031
66 | 16247 | 761.5781
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16250 | 761.7188
66 | 16248 | 761.625
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16249 | 761.6719
66 | 16238 | 761.1562
66 | 16247 | 761.5781
66 | 16247 | 761.5781
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16243 | 761.3906
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16251 | 761.7656
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16242 | 761.3438
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16239 | 761.2031
66 | 16247 | 761.5781
66 | 16249 | 761.6719
66 | 16239 | 761.2031
66 | 16246 | 761.5312
66 | 16241 | 761.2969
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16239 | 761.2031
66 | 16248 | 761.625
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16245 | 761.4844
66 | 16248 | 761.625
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16242 | 761.3438
66 | 16249 | 761.6719
66 | 16238 | 761.1562
66 | 16249 | 761.6719
66 | 16243 | 761.3906
66 | 16245 | 761.4844
66 | 16252 | 761.8125
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16245 | 761.4844
66 | 16241 | 761.2969
66 | 16250 | 761.7188
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16244 | 761.4375
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16245 | 761.4844
66 | 16242 | 761.3438
66 | 16251 | 761.7656
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16246 | 761.5312
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16244 | 761.4375
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16246 | 761.5312
66 | 16238 | 761.1562
66 | 16249 | 761.6719
66 | 16239 | 761.2031
66 | 16245 | 761.4844
66 | 16250 | 761.7188
66 | 16239 | 761.2031
66 | 16246 | 761.5312
66 | 16245 | 761.4844
66 | 16239 | 761.2031
66 | 16249 | 761.6719
66 | 16240 | 761.25
66 | 16246 | 761.5312
66 | 16250 | 761.7188
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16244 | 761.4375
66 | 16239 | 761.2031
66 | 16248 | 761.625
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16248 | 761.625
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16238 | 761.1562
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16238 | 761.1562
66 | 16243 | 761.3906
66 | 16247 | 761.5781
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16247 | 761.5781
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16247 | 761.5781
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16242 | 761.3438
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16240 | 761.25
66 | 16243 | 761.3906
66 | 16248 | 761.625
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16240 | 761.25
66 | 16250 | 761.7188
66 | 16239 | 761.2031
66 | 16246 | 761.5312
66 | 16248 | 761.625
66 | 16239 | 761.2031
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16242 | 761.3438
66 | 16248 | 761.625
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16243 | 761.3906
66 | 16238 | 761.1562
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16246 | 761.5312
66 | 16248 | 761.625
66 | 16249 | 761.6719
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16239 | 761.2031
66 | 16242 | 761.3438
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16244 | 761.4375
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16244 | 761.4375
66 | 16237 | 761.1094
66 | 16249 | 761.6719
66 | 16241 | 761.2969
66 | 16245 | 761.4844
66 | 16250 | 761.7188
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16239 | 761.2031
66 | 16242 | 761.3438
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16238 | 761.1562
66 | 16244 | 761.4375
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16237 | 761.1094
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16243 | 761.3906
66 | 16239 | 761.2031
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16238 | 761.1562
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16239 | 761.2031
66 | 16243 | 761.3906
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16244 | 761.4375
66 | 16237 | 761.1094
66 | 16248 | 761.625
66 | 16239 | 761.2031
66 | 16244 | 761.4375
66 | 16248 | 761.625
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16238 | 761.1562
66 | 16241 | 761.2969
66 | 16249 | 761.6719
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16246 | 761.5312
66 | 16240 | 761.25
66 | 16246 | 761.5312
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16239 | 761.2031
66 | 16240 | 761.25
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16246 | 761.5312
66 | 16240 | 761.25
66 | 16240 | 761.25
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16244 | 761.4375
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16240 | 761.25
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16239 | 761.2031
66 | 16246 | 761.5312
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16241 | 761.2969
66 | 16239 | 761.2031
66 | 16245 | 761.4844
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16238 | 761.1562
66 | 16247 | 761.5781
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16240 | 761.25
66 | 16242 | 761.3438
66 | 16248 | 761.625
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16238 | 761.1562
66 | 16248 | 761.625
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16239 | 761.2031
66 | 16240 | 761.25
66 | 16249 | 761.6719
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16245 | 761.4844
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16240 | 761.25
66 | 16242 | 761.3438
66 | 16247 | 761.5781
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16246 | 761.5312
66 | 16246 | 761.5312
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16245 | 761.4844
66 | 16237 | 761.1094
66 | 16247 | 761.5781
66 | 16241 | 761.2969
66 | 16242 | 761.3438
66 | 16248 | 761.625
66 | 16237 | 761.1094
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16239 | 761.2031
66 | 16248 | 761.625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16247 | 761.5781
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16240 | 761.25
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16244 | 761.4375
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16237 | 761.1094
66 | 16242 | 761.3438
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16242 | 761.3438
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16238 | 761.1562
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16246 | 761.5312
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16236 | 761.0625
66 | 16241 | 761.2969
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16244 | 761.4375
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16246 | 761.5312
66 | 16240 | 761.25
66 | 16244 | 761.4375
66 | 16247 | 761.5781
66 | 16237 | 761.1094
66 | 16243 | 761.3906
66 | 16246 | 761.5312
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16240 | 761.25
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16237 | 761.1094
66 | 16239 | 761.2031
66 | 16247 | 761.5781
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16247 | 761.5781
66 | 16236 | 761.0625
66 | 16247 | 761.5781
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16248 | 761.625
66 | 16237 | 761.1094
66 | 16243 | 761.3906
66 | 16247 | 761.5781
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16238 | 761.1562
66 | 16240 | 761.25
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16238 | 761.1562
66 | 16241 | 761.2969
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16237 | 761.1094
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16247 | 761.5781
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16244 | 761.4375
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16244 | 761.4375
66 | 16241 | 761.2969
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16238 | 761.1562
66 | 16234 | 760.9688
66 | 16244 | 761.4375
66 | 16236 | 761.0625
66 | 16239 | 761.2031
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16245 | 761.4844
66 | 16245 | 761.4844
66 | 16239 | 761.2031
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16236 | 761.0625
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16238 | 761.1562
66 | 16237 | 761.1094
66 | 16245 | 761.4844
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16240 | 761.25
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16241 | 761.2969
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16244 | 761.4375
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16238 | 761.1562
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16244 | 761.4375
66 | 16238 | 761.1562
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16242 | 761.3438
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16236 | 761.0625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16245 | 761.4844
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16242 | 761.3438
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16239 | 761.2031
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16236 | 761.0625
66 | 16238 | 761.1562
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16241 | 761.2969
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16234 | 760.9688
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16241 | 761.2969
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16237 | 761.1094
66 | 16240 | 761.25
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16246 | 761.5312
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16237 | 761.1094
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16238 | 761.1562
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16241 | 761.2969
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16238 | 761.1562
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16239 | 761.2031
66 | 16246 | 761.5312
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16245 | 761.4844
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16246 | 761.5312
66 | 16235 | 761.0156
66 | 16241 | 761.2969
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16240 | 761.25
66 | 16241 | 761.2969
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16234 | 760.9688
66 | 16244 | 761.4375
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16244 | 761.4375
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16241 | 761.2969
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16242 | 761.3438
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16244 | 761.4375
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16244 | 761.4375
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16237 | 761.1094
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16241 | 761.2969
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16240 | 761.25
66 | 16247 | 761.5781
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16234 | 760.9688
66 | 16238 | 761.1562
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16234 | 760.9688
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16238 | 761.1562
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16237 | 761.1094
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16239 | 761.2031
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16234 | 760.9688
66 | 16239 | 761.2031
66 | 16243 | 761.3906
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16234 | 760.9688
66 | 16238 | 761.1562
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16236 | 761.0625
66 | 16239 | 761.2031
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16235 | 761.0156
66 | 16239 | 761.2031
66 | 16241 | 761.2969
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16236 | 761.0625
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16241 | 761.2969
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16237 | 761.1094
66 | 16233 | 760.9219
66 | 16244 | 761.4375
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16241 | 761.2969
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16232 | 760.875
66 | 16237 | 761.1094
66 | 16238 | 761.1562
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16234 | 760.9688
66 | 16240 | 761.25
66 | 16242 | 761.3438
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16241 | 761.2969
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16241 | 761.2969
66 | 16234 | 760.9688
66 | 16243 | 761.3906
66 | 16235 | 761.0156
66 | 16238 | 761.1562
66 | 16244 | 761.4375
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16238 | 761.1562
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16238 | 761.1562
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16235 | 761.0156
66 | 16244 | 761.4375
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16237 | 761.1094
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16239 | 761.2031
66 | 16235 | 761.0156
66 | 16245 | 761.4844
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16242 | 761.3438
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16235 | 761.0156
66 | 16237 | 761.1094
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16233 | 760.9219
66 | 16238 | 761.1562
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16235 | 761.0156
66 | 16238 | 761.1562
66 | 16245 | 761.4844
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16234 | 760.9688
66 | 16237 | 761.1094
66 | 16241 | 761.2969
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16237 | 761.1094
66 | 16240 | 761.25
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16237 | 761.1094
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16237 | 761.1094
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16234 | 760.9688
66 | 16236 | 761.0625
66 | 16243 | 761.3906
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16242 | 761.3438
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16238 | 761.1562
66 | 16233 | 760.9219
66 | 16243 | 761.3906
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16234 | 760.9688
66 | 16237 | 761.1094
66 | 16243 | 761.3906
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16238 | 761.1562
66 | 16236 | 761.0625
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16236 | 761.0625
66 | 16238 | 761.1562
66 | 16244 | 761.4375
66 | 16233 | 760.9219
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16233 | 760.9219
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16235 | 761.0156
66 | 16235 | 761.0156
66 | 16243 | 761.3906
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16236 | 761.0625
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16233 | 760.9219
66 | 16235 | 761.0156
66 | 16241 | 761.2969
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16238 | 761.1562
66 | 16241 | 761.2969
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16233 | 760.9219
66 | 16237 | 761.1094
66 | 16242 | 761.3438
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16230 | 760.7812
66 | 16235 | 761.0156
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16237 | 761.1094
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16237 | 761.1094
66 | 16234 | 760.9688
66 | 16242 | 761.3438
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16235 | 761.0156
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16230 | 760.7812
66 | 16237 | 761.1094
66 | 16237 | 761.1094
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16236 | 761.0625
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16233 | 760.9219
66 | 16236 | 761.0625
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16235 | 761.0156
66 | 16237 | 761.1094
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16235 | 761.0156
66 | 16240 | 761.25
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16232 | 760.875
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16237 | 761.1094
66 | 16238 | 761.1562
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16236 | 761.0625
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16232 | 760.875
66 | 16235 | 761.0156
66 | 16240 | 761.25
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16237 | 761.1094
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16232 | 760.875
66 | 16237 | 761.1094
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16242 | 761.3438
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16232 | 760.875
66 | 16236 | 761.0625
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16237 | 761.1094
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16233 | 760.9219
66 | 16236 | 761.0625
66 | 16241 | 761.2969
66 | 16228 | 760.6875
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16230 | 760.7812
66 | 16236 | 761.0625
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16232 | 760.875
66 | 16236 | 761.0625
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16237 | 761.1094
66 | 16232 | 760.875
66 | 16241 | 761.2969
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16226 | 760.5938
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16233 | 760.9219
66 | 16233 | 760.9219
66 | 16239 | 761.2031
66 | 16232 | 760.875
66 | 16240 | 761.25
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16234 | 760.9688
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16235 | 761.0156
66 | 16233 | 760.9219
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16236 | 761.0625
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16229 | 760.7344
66 | 16241 | 761.2969
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16232 | 760.875
66 | 16234 | 760.9688
66 | 16234 | 760.9688
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16230 | 760.7812
66 | 16236 | 761.0625
66 | 16239 | 761.2031
66 | 16231 | 760.8281
66 | 16241 | 761.2969
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16234 | 760.9688
66 | 16232 | 760.875
66 | 16233 | 760.9219
66 | 16237 | 761.1094
66 | 16228 | 760.6875
66 | 16239 | 761.2031
66 | 16228 | 760.6875
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16236 | 761.0625
66 | 16233 | 760.9219
66 | 16235 | 761.0156
66 | 16238 | 761.1562
66 | 16231 | 760.8281
66 | 16239 | 761.2031
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16230 | 760.7812
66 | 16233 | 760.9219
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16239 | 761.2031
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16236 | 761.0625
66 | 16232 | 760.875
66 | 16239 | 761.2031
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16230 | 760.7812
66 | 16232 | 760.875
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16231 | 760.8281
66 | 16235 | 761.0156
66 | 16238 | 761.1562
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16230 | 760.7812
66 | 16235 | 761.0156
66 | 16236 | 761.0625
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16236 | 761.0625
66 | 16231 | 760.8281
66 | 16242 | 761.3438
66 | 16229 | 760.7344
66 | 16240 | 761.25
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16228 | 760.6875
66 | 16234 | 760.9688
66 | 16232 | 760.875
66 | 16233 | 760.9219
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16240 | 761.25
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16229 | 760.7344
66 | 16233 | 760.9219
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16235 | 761.0156
66 | 16233 | 760.9219
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16229 | 760.7344
66 | 16234 | 760.9688
66 | 16233 | 760.9219
66 | 16231 | 760.8281
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16235 | 761.0156
66 | 16234 | 760.9688
66 | 16241 | 761.2969
66 | 16231 | 760.8281
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16226 | 760.5938
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16235 | 761.0156
66 | 16235 | 761.0156
66 | 16233 | 760.9219
66 | 16240 | 761.25
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16233 | 760.9219
66 | 16232 | 760.875
66 | 16238 | 761.1562
66 | 16230 | 760.7812
66 | 16241 | 761.2969
66 | 16230 | 760.7812
66 | 16240 | 761.25
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16231 | 760.8281
66 | 16232 | 760.875
66 | 16235 | 761.0156
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16229 | 760.7344
66 | 16239 | 761.2031
66 | 16230 | 760.7812
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16235 | 761.0156
66 | 16233 | 760.9219
66 | 16230 | 760.7812
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16235 | 761.0156
66 | 16231 | 760.8281
66 | 16235 | 761.0156
66 | 16235 | 761.0156
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16233 | 760.9219
66 | 16233 | 760.9219
66 | 16229 | 760.7344
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16239 | 761.2031
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16232 | 760.875
66 | 16229 | 760.7344
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16229 | 760.7344
66 | 16237 | 761.1094
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16231 | 760.8281
66 | 16235 | 761.0156
66 | 16234 | 760.9688
66 | 16230 | 760.7812
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16225 | 760.5469
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16225 | 760.5469
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16234 | 760.9688
66 | 16230 | 760.7812
66 | 16233 | 760.9219
66 | 16234 | 760.9688
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16238 | 761.1562
66 | 16226 | 760.5938
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16238 | 761.1562
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16229 | 760.7344
66 | 16232 | 760.875
66 | 16233 | 760.9219
66 | 16233 | 760.9219
66 | 16237 | 761.1094
66 | 16230 | 760.7812
66 | 16239 | 761.2031
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16228 | 760.6875
66 | 16236 | 761.0625
66 | 16230 | 760.7812
66 | 16233 | 760.9219
66 | 16235 | 761.0156
66 | 16227 | 760.6406
66 | 16238 | 761.1562
66 | 16226 | 760.5938
66 | 16238 | 761.1562
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16233 | 760.9219
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16230 | 760.7812
66 | 16235 | 761.0156
66 | 16233 | 760.9219
66 | 16231 | 760.8281
66 | 16237 | 761.1094
66 | 16228 | 760.6875
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16224 | 760.5
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16236 | 761.0625
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16228 | 760.6875
66 | 16235 | 761.0156
66 | 16228 | 760.6875
66 | 16235 | 761.0156
66 | 16228 | 760.6875
66 | 16233 | 760.9219
66 | 16229 | 760.7344
66 | 16232 | 760.875
66 | 16232 | 760.875
66 | 16228 | 760.6875
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16239 | 761.2031
66 | 16226 | 760.5938
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16231 | 760.8281
66 | 16231 | 760.8281
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16226 | 760.5938
66 | 16232 | 760.875
66 | 16226 | 760.5938
66 | 16232 | 760.875
66 | 16229 | 760.7344
66 | 16229 | 760.7344
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16236 | 761.0625
66 | 16225 | 760.5469
66 | 16237 | 761.1094
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16225 | 760.5469
66 | 16231 | 760.8281
66 | 16231 | 760.8281
66 | 16231 | 760.8281
66 | 16234 | 760.9688
66 | 16228 | 760.6875
66 | 16238 | 761.1562
66 | 16227 | 760.6406
66 | 16237 | 761.1094
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16234 | 760.9688
66 | 16222 | 760.4062
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16235 | 761.0156
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16226 | 760.5938
66 | 16233 | 760.9219
66 | 16227 | 760.6406
66 | 16233 | 760.9219
66 | 16228 | 760.6875
66 | 16229 | 760.7344
66 | 16230 | 760.7812
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16226 | 760.5938
66 | 16233 | 760.9219
66 | 16227 | 760.6406
66 | 16233 | 760.9219
66 | 16230 | 760.7812
66 | 16229 | 760.7344
66 | 16228 | 760.6875
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16225 | 760.5469
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16233 | 760.9219
66 | 16228 | 760.6875
66 | 16229 | 760.7344
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16235 | 761.0156
66 | 16227 | 760.6406
66 | 16234 | 760.9688
66 | 16230 | 760.7812
66 | 16229 | 760.7344
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16231 | 760.8281
66 | 16225 | 760.5469
66 | 16230 | 760.7812
66 | 16233 | 760.9219
66 | 16226 | 760.5938
66 | 16238 | 761.1562
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16230 | 760.7812
66 | 16224 | 760.5
66 | 16229 | 760.7344
66 | 16229 | 760.7344
66 | 16229 | 760.7344
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16226 | 760.5938
66 | 16233 | 760.9219
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16227 | 760.6406
66 | 16232 | 760.875
66 | 16226 | 760.5938
66 | 16230 | 760.7812
66 | 16228 | 760.6875
66 | 16226 | 760.5938
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16235 | 761.0156
66 | 16224 | 760.5
66 | 16235 | 761.0156
66 | 16224 | 760.5
66 | 16235 | 761.0156
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16233 | 760.9219
66 | 16225 | 760.5469
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16235 | 761.0156
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16235 | 761.0156
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16234 | 760.9688
66 | 16225 | 760.5469
66 | 16234 | 760.9688
66 | 16231 | 760.8281
66 | 16226 | 760.5938
66 | 16227 | 760.6406
66 | 16230 | 760.7812
66 | 16225 | 760.5469
66 | 16229 | 760.7344
66 | 16223 | 760.4531
66 | 16229 | 760.7344
66 | 16224 | 760.5
66 | 16228 | 760.6875
66 | 16230 | 760.7812
66 | 16225 | 760.5469
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16236 | 761.0625
66 | 16226 | 760.5938
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16235 | 761.0156
66 | 16223 | 760.4531
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16223 | 760.4531
66 | 16228 | 760.6875
66 | 16223 | 760.4531
66 | 16227 | 760.6406
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16235 | 761.0156
66 | 16225 | 760.5469
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16225 | 760.5469
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16228 | 760.6875
66 | 16226 | 760.5938
66 | 16227 | 760.6406
66 | 16230 | 760.7812
66 | 16222 | 760.4062
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16234 | 760.9688
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16223 | 760.4531
66 | 16229 | 760.7344
66 | 16223 | 760.4531
66 | 16227 | 760.6406
66 | 16229 | 760.7344
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16222 | 760.4062
66 | 16234 | 760.9688
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16231 | 760.8281
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16221 | 760.3594
66 | 16231 | 760.8281
66 | 16221 | 760.3594
66 | 16231 | 760.8281
66 | 16223 | 760.4531
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16226 | 760.5938
66 | 16230 | 760.7812
66 | 16229 | 760.7344
66 | 16225 | 760.5469
66 | 16231 | 760.8281
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16220 | 760.3125
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16224 | 760.5
66 | 16233 | 760.9219
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16234 | 760.9688
66 | 16222 | 760.4062
66 | 16233 | 760.9219
66 | 16222 | 760.4062
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16233 | 760.9219
66 | 16220 | 760.3125
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16231 | 760.8281
66 | 16221 | 760.3594
66 | 16231 | 760.8281
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16230 | 760.7812
66 | 16222 | 760.4062
66 | 16227 | 760.6406
66 | 16221 | 760.3594
66 | 16226 | 760.5938
66 | 16225 | 760.5469
66 | 16225 | 760.5469
66 | 16229 | 760.7344
66 | 16223 | 760.4531
66 | 16233 | 760.9219
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16223 | 760.4531
66 | 16231 | 760.8281
66 | 16223 | 760.4531
66 | 16227 | 760.6406
66 | 16227 | 760.6406
66 | 16224 | 760.5
66 | 16232 | 760.875
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16219 | 760.2656
66 | 16231 | 760.8281
66 | 16221 | 760.3594
66 | 16230 | 760.7812
66 | 16222 | 760.4062
66 | 16230 | 760.7812
66 | 16224 | 760.5
66 | 16229 | 760.7344
66 | 16229 | 760.7344
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16219 | 760.2656
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16230 | 760.7812
66 | 16222 | 760.4062
66 | 16230 | 760.7812
66 | 16228 | 760.6875
66 | 16226 | 760.5938
66 | 16233 | 760.9219
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16221 | 760.3594
66 | 16229 | 760.7344
66 | 16223 | 760.4531
66 | 16225 | 760.5469
66 | 16228 | 760.6875
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16221 | 760.3594
66 | 16231 | 760.8281
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16225 | 760.5469
66 | 16225 | 760.5469
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16222 | 760.4062
66 | 16228 | 760.6875
66 | 16225 | 760.5469
66 | 16224 | 760.5
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16217 | 760.1719
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16221 | 760.3594
66 | 16229 | 760.7344
66 | 16224 | 760.5
66 | 16228 | 760.6875
66 | 16227 | 760.6406
66 | 16222 | 760.4062
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16230 | 760.7812
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16221 | 760.3594
66 | 16231 | 760.8281
66 | 16220 | 760.3125
66 | 16231 | 760.8281
66 | 16221 | 760.3594
66 | 16230 | 760.7812
66 | 16222 | 760.4062
66 | 16226 | 760.5938
66 | 16222 | 760.4062
66 | 16225 | 760.5469
66 | 16226 | 760.5938
66 | 16221 | 760.3594
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16232 | 760.875
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16220 | 760.3125
66 | 16231 | 760.8281
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16217 | 760.1719
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16223 | 760.4531
66 | 16227 | 760.6406
66 | 16230 | 760.7812
66 | 16223 | 760.4531
66 | 16232 | 760.875
66 | 16222 | 760.4062
66 | 16231 | 760.8281
66 | 16219 | 760.2656
66 | 16230 | 760.7812
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16216 | 760.125
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16221 | 760.3594
66 | 16228 | 760.6875
66 | 16226 | 760.5938
66 | 16225 | 760.5469
66 | 16227 | 760.6406
66 | 16221 | 760.3594
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16217 | 760.1719
66 | 16227 | 760.6406
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16230 | 760.7812
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16221 | 760.3594
66 | 16232 | 760.875
66 | 16220 | 760.3125
66 | 16231 | 760.8281
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16230 | 760.7812
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16217 | 760.1719
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16230 | 760.7812
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16220 | 760.3125
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16220 | 760.3125
66 | 16225 | 760.5469
66 | 16225 | 760.5469
66 | 16221 | 760.3594
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16220 | 760.3125
66 | 16227 | 760.6406
66 | 16222 | 760.4062
66 | 16224 | 760.5
66 | 16223 | 760.4531
66 | 16221 | 760.3594
66 | 16225 | 760.5469
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16231 | 760.8281
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16227 | 760.6406
66 | 16219 | 760.2656
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16229 | 760.7344
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16224 | 760.5
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16217 | 760.1719
66 | 16229 | 760.7344
66 | 16218 | 760.2188
66 | 16230 | 760.7812
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16219 | 760.2656
66 | 16226 | 760.5938
66 | 16225 | 760.5469
66 | 16222 | 760.4062
66 | 16230 | 760.7812
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16227 | 760.6406
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16217 | 760.1719
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16217 | 760.1719
66 | 16224 | 760.5
66 | 16221 | 760.3594
66 | 16223 | 760.4531
66 | 16223 | 760.4531
66 | 16222 | 760.4062
66 | 16226 | 760.5938
66 | 16220 | 760.3125
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16228 | 760.6875
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16228 | 760.6875
66 | 16217 | 760.1719
66 | 16228 | 760.6875
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16221 | 760.3594
66 | 16223 | 760.4531
66 | 16226 | 760.5938
66 | 16221 | 760.3594
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16227 | 760.6406
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16214 | 760.0312
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16219 | 760.2656
66 | 16225 | 760.5469
66 | 16223 | 760.4531
66 | 16222 | 760.4062
66 | 16227 | 760.6406
66 | 16218 | 760.2188
66 | 16225 | 760.5469
66 | 16216 | 760.125
66 | 16227 | 760.6406
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16218 | 760.2188
66 | 16225 | 760.5469
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16225 | 760.5469
66 | 16219 | 760.2656
66 | 16228 | 760.6875
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16220 | 760.3125
66 | 16223 | 760.4531
66 | 16224 | 760.5
66 | 16221 | 760.3594
66 | 16225 | 760.5469
66 | 16216 | 760.125
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16214 | 760.0312
66 | 16225 | 760.5469
66 | 16213 | 759.9844
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16214 | 760.0312
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16223 | 760.4531
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16224 | 760.5
66 | 16220 | 760.3125
66 | 16228 | 760.6875
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16224 | 760.5
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16220 | 760.3125
66 | 16220 | 760.3125
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16224 | 760.5
66 | 16215 | 760.0781
66 | 16223 | 760.4531
66 | 16221 | 760.3594
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16224 | 760.5
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16227 | 760.6406
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16223 | 760.4531
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16214 | 760.0312
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16217 | 760.1719
66 | 16223 | 760.4531
66 | 16222 | 760.4062
66 | 16217 | 760.1719
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16222 | 760.4062
66 | 16220 | 760.3125
66 | 16226 | 760.5938
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16214 | 760.0312
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16223 | 760.4531
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16221 | 760.3594
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16215 | 760.0781
66 | 16222 | 760.4062
66 | 16220 | 760.3125
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16219 | 760.2656
66 | 16217 | 760.1719
66 | 16226 | 760.5938
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16222 | 760.4062
66 | 16221 | 760.3594
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16223 | 760.4531
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16215 | 760.0781
66 | 16223 | 760.4531
66 | 16218 | 760.2188
66 | 16219 | 760.2656
66 | 16227 | 760.6406
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16226 | 760.5938
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16226 | 760.5938
66 | 16213 | 759.9844
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16218 | 760.2188
66 | 16214 | 760.0312
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16226 | 760.5938
66 | 16213 | 759.9844
66 | 16223 | 760.4531
66 | 16218 | 760.2188
66 | 16218 | 760.2188
66 | 16225 | 760.5469
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16216 | 760.125
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16215 | 760.0781
66 | 16226 | 760.5938
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16224 | 760.5
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16217 | 760.1719
66 | 16217 | 760.1719
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16223 | 760.4531
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16224 | 760.5
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16225 | 760.5469
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16215 | 760.0781
66 | 16215 | 760.0781
66 | 16224 | 760.5
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16219 | 760.2656
66 | 16215 | 760.0781
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16223 | 760.4531
66 | 16213 | 759.9844
66 | 16223 | 760.4531
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16226 | 760.5938
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16219 | 760.2656
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16223 | 760.4531
66 | 16215 | 760.0781
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16221 | 760.3594
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16225 | 760.5469
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16215 | 760.0781
66 | 16225 | 760.5469
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16218 | 760.2188
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16221 | 760.3594
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16222 | 760.4062
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16218 | 760.2188
66 | 16213 | 759.9844
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16223 | 760.4531
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16216 | 760.125
66 | 16212 | 759.9375
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16219 | 760.2656
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16217 | 760.1719
66 | 16215 | 760.0781
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16211 | 759.8906
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16221 | 760.3594
66 | 16219 | 760.2656
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16220 | 760.3125
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16222 | 760.4062
66 | 16214 | 760.0312
66 | 16222 | 760.4062
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16219 | 760.2656
66 | 16213 | 759.9844
66 | 16223 | 760.4531
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16225 | 760.5469
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16217 | 760.1719
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16220 | 760.3125
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16220 | 760.3125
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16218 | 760.2188
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16217 | 760.1719
66 | 16212 | 759.9375
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16216 | 760.125
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16223 | 760.4531
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16217 | 760.1719
66 | 16209 | 759.7969
66 | 16219 | 760.2656
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16222 | 760.4062
66 | 16216 | 760.125
66 | 16211 | 759.8906
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16219 | 760.2656
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16223 | 760.4531
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16225 | 760.5469
66 | 16211 | 759.8906
66 | 16220 | 760.3125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16224 | 760.5
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16223 | 760.4531
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16219 | 760.2656
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16207 | 759.7031
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16224 | 760.5
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16215 | 760.0781
66 | 16210 | 759.8438
66 | 16221 | 760.3594
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16212 | 759.9375
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16223 | 760.4531
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16216 | 760.125
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16208 | 759.75
66 | 16217 | 760.1719
66 | 16222 | 760.4062
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16215 | 760.0781
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16223 | 760.4531
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16209 | 759.7969
66 | 16218 | 760.2188
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16219 | 760.2656
66 | 16222 | 760.4062
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16214 | 760.0312
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16214 | 760.0312
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16221 | 760.3594
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16222 | 760.4062
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16212 | 759.9375
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16221 | 760.3594
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16217 | 760.1719
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16213 | 759.9844
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16220 | 760.3125
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16221 | 760.3594
66 | 16213 | 759.9844
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16215 | 760.0781
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16209 | 759.7969
66 | 16218 | 760.2188
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16208 | 759.75
66 | 16217 | 760.1719
66 | 16218 | 760.2188
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16212 | 759.9375
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16218 | 760.2188
66 | 16206 | 759.6562
66 | 16208 | 759.75
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16221 | 760.3594
66 | 16212 | 759.9375
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16213 | 759.9844
66 | 16207 | 759.7031
66 | 16216 | 760.125
66 | 16219 | 760.2656
66 | 16212 | 759.9375
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16220 | 760.3125
66 | 16209 | 759.7969
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16211 | 759.8906
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16218 | 760.2188
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16217 | 760.1719
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16214 | 760.0312
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16217 | 760.1719
66 | 16218 | 760.2188
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16209 | 759.7969
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16212 | 759.9375
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16215 | 760.0781
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16214 | 760.0312
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16211 | 759.8906
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16214 | 760.0312
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16209 | 759.7969
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16211 | 759.8906
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16220 | 760.3125
66 | 16212 | 759.9375
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16212 | 759.9375
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16213 | 759.9844
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16221 | 760.3594
66 | 16210 | 759.8438
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16211 | 759.8906
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16209 | 759.7969
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16218 | 760.2188
66 | 16212 | 759.9375
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16219 | 760.2656
66 | 16210 | 759.8438
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16214 | 760.0312
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16214 | 760.0312
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16212 | 759.9375
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16220 | 760.3125
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16216 | 760.125
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16219 | 760.2656
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16220 | 760.3125
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16212 | 759.9375
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16216 | 760.125
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16219 | 760.2656
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16211 | 759.8906
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16218 | 760.2188
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16214 | 760.0312
66 | 16218 | 760.2188
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16216 | 760.125
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16216 | 760.125
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16218 | 760.2188
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16217 | 760.1719
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16218 | 760.2188
66 | 16205 | 759.6094
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16203 | 759.5156
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16202 | 759.4688
66 | 16209 | 759.7969
66 | 16216 | 760.125
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16216 | 760.125
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16217 | 760.1719
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16214 | 760.0312
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16215 | 760.0781
66 | 16211 | 759.8906
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16214 | 760.0312
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16214 | 760.0312
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16215 | 760.0781
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16211 | 759.8906
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16215 | 760.0781
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16215 | 760.0781
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16214 | 760.0312
66 | 16209 | 759.7969
66 | 16204 | 759.5625
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16205 | 759.6094
66 | 16214 | 760.0312
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16216 | 760.125
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16209 | 759.7969
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16209 | 759.7969
66 | 16215 | 760.0781
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16217 | 760.1719
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16214 | 760.0312
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16214 | 760.0312
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16212 | 759.9375
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16207 | 759.7031
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16212 | 759.9375
66 | 16204 | 759.5625
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16209 | 759.7969
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16208 | 759.75
66 | 16216 | 760.125
66 | 16205 | 759.6094
66 | 16214 | 760.0312
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16216 | 760.125
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16211 | 759.8906
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16206 | 759.6562
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16215 | 760.0781
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16211 | 759.8906
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16204 | 759.5625
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16214 | 760.0312
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16205 | 759.6094
66 | 16207 | 759.7031
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16211 | 759.8906
66 | 16204 | 759.5625
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16204 | 759.5625
66 | 16213 | 759.9844
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16202 | 759.4688
66 | 16213 | 759.9844
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16209 | 759.7969
66 | 16208 | 759.75
66 | 16214 | 760.0312
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16205 | 759.6094
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16214 | 760.0312
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16213 | 759.9844
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16201 | 759.4219
66 | 16212 | 759.9375
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16213 | 759.9844
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16205 | 759.6094
66 | 16207 | 759.7031
66 | 16215 | 760.0781
66 | 16202 | 759.4688
66 | 16210 | 759.8438
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16208 | 759.75
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16212 | 759.9375
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16211 | 759.8906
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16201 | 759.4219
66 | 16212 | 759.9375
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16202 | 759.4688
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16208 | 759.75
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16213 | 759.9844
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16209 | 759.7969
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16209 | 759.7969
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16207 | 759.7031
66 | 16209 | 759.7969
66 | 16204 | 759.5625
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16212 | 759.9375
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16204 | 759.5625
66 | 16208 | 759.75
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16211 | 759.8906
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16209 | 759.7969
66 | 16203 | 759.5156
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16202 | 759.4688
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16200 | 759.375
66 | 16211 | 759.8906
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16213 | 759.9844
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16203 | 759.5156
66 | 16209 | 759.7969
66 | 16209 | 759.7969
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16205 | 759.6094
66 | 16213 | 759.9844
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16208 | 759.75
66 | 16202 | 759.4688
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16205 | 759.6094
66 | 16209 | 759.7969
66 | 16207 | 759.7031
66 | 16207 | 759.7031
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16211 | 759.8906
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16211 | 759.8906
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16212 | 759.9375
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16204 | 759.5625
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16207 | 759.7031
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16203 | 759.5156
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16211 | 759.8906
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16211 | 759.8906
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16206 | 759.6562
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16198 | 759.2812
66 | 16209 | 759.7969
66 | 16198 | 759.2812
66 | 16209 | 759.7969
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16202 | 759.4688
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16201 | 759.4219
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16209 | 759.7969
66 | 16203 | 759.5156
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16209 | 759.7969
66 | 16199 | 759.3281
66 | 16210 | 759.8438
66 | 16201 | 759.4219
66 | 16208 | 759.75
66 | 16202 | 759.4688
66 | 16201 | 759.4219
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16206 | 759.6562
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16210 | 759.8438
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16202 | 759.4688
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16204 | 759.5625
66 | 16202 | 759.4688
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16209 | 759.7969
66 | 16201 | 759.4219
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16202 | 759.4688
66 | 16197 | 759.2344
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16200 | 759.375
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16205 | 759.6094
66 | 16194 | 759.0938
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16208 | 759.75
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16202 | 759.4688
66 | 16203 | 759.5156
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16198 | 759.2812
66 | 16203 | 759.5156
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16203 | 759.5156
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16205 | 759.6094
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16194 | 759.0938
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16204 | 759.5625
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16202 | 759.4688
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16201 | 759.4219
66 | 16204 | 759.5625
66 | 16194 | 759.0938
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16200 | 759.375
66 | 16204 | 759.5625
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16207 | 759.7031
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16206 | 759.6562
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16199 | 759.3281
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16209 | 759.7969
66 | 16200 | 759.375
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16208 | 759.75
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16207 | 759.7031
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16204 | 759.5625
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16199 | 759.3281
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16203 | 759.5156
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16202 | 759.4688
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16200 | 759.375
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16201 | 759.4219
66 | 16202 | 759.4688
66 | 16195 | 759.1406
66 | 16201 | 759.4219
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16201 | 759.4219
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16206 | 759.6562
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16208 | 759.75
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16202 | 759.4688
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16207 | 759.7031
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16203 | 759.5156
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16201 | 759.4219
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16206 | 759.6562
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16199 | 759.3281
66 | 16203 | 759.5156
66 | 16193 | 759.0469
66 | 16200 | 759.375
66 | 16203 | 759.5156
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16206 | 759.6562
66 | 16202 | 759.4688
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16199 | 759.3281
66 | 16202 | 759.4688
66 | 16195 | 759.1406
66 | 16198 | 759.2812
66 | 16203 | 759.5156
66 | 16194 | 759.0938
66 | 16201 | 759.4219
66 | 16202 | 759.4688
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16205 | 759.6094
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16206 | 759.6562
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16194 | 759.0938
66 | 16198 | 759.2812
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16205 | 759.6094
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16198 | 759.2812
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16207 | 759.7031
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16193 | 759.0469
66 | 16205 | 759.6094
66 | 16203 | 759.5156
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16196 | 759.1875
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16195 | 759.1406
66 | 16204 | 759.5625
66 | 16202 | 759.4688
66 | 16193 | 759.0469
66 | 16197 | 759.2344
66 | 16202 | 759.4688
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16204 | 759.5625
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16197 | 759.2344
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16207 | 759.7031
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16203 | 759.5156
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16193 | 759.0469
66 | 16200 | 759.375
66 | 16204 | 759.5625
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16204 | 759.5625
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16197 | 759.2344
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16199 | 759.3281
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16205 | 759.6094
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16191 | 758.9531
66 | 16201 | 759.4219
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16204 | 759.5625
66 | 16203 | 759.5156
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16206 | 759.6562
66 | 16201 | 759.4219
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16199 | 759.3281
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16197 | 759.2344
66 | 16205 | 759.6094
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16196 | 759.1875
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16196 | 759.1875
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16204 | 759.5625
66 | 16203 | 759.5156
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16198 | 759.2812
66 | 16203 | 759.5156
66 | 16197 | 759.2344
66 | 16190 | 758.9062
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16205 | 759.6094
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16196 | 759.1875
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16204 | 759.5625
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16204 | 759.5625
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16189 | 758.8594
66 | 16193 | 759.0469
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16195 | 759.1406
66 | 16189 | 758.8594
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16199 | 759.3281
66 | 16204 | 759.5625
66 | 16202 | 759.4688
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16190 | 758.9062
66 | 16193 | 759.0469
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16203 | 759.5156
66 | 16202 | 759.4688
66 | 16196 | 759.1875
66 | 16192 | 759
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16202 | 759.4688
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16203 | 759.5156
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16190 | 758.9062
66 | 16197 | 759.2344
66 | 16201 | 759.4219
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16201 | 759.4219
66 | 16192 | 759
66 | 16192 | 759
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16195 | 759.1406
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16195 | 759.1406
66 | 16201 | 759.4219
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16196 | 759.1875
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16202 | 759.4688
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16203 | 759.5156
66 | 16200 | 759.375
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16201 | 759.4219
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16195 | 759.1406
66 | 16203 | 759.5156
66 | 16201 | 759.4219
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16201 | 759.4219
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16194 | 759.0938
66 | 16187 | 758.7656
66 | 16193 | 759.0469
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16196 | 759.1875
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16201 | 759.4219
66 | 16201 | 759.4219
66 | 16198 | 759.2812
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16197 | 759.2344
66 | 16200 | 759.375
66 | 16196 | 759.1875
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16192 | 759
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16195 | 759.1406
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16201 | 759.4219
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16187 | 758.7656
66 | 16193 | 759.0469
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16198 | 759.2812
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16194 | 759.0938
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16197 | 759.2344
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16193 | 759.0469
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16197 | 759.2344
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16197 | 759.2344
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16195 | 759.1406
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16200 | 759.375
66 | 16196 | 759.1875
66 | 16192 | 759
66 | 16189 | 758.8594
66 | 16195 | 759.1406
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16200 | 759.375
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16203 | 759.5156
66 | 16199 | 759.3281
66 | 16196 | 759.1875
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16197 | 759.2344
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16197 | 759.2344
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16198 | 759.2812
66 | 16200 | 759.375
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16192 | 759
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16197 | 759.2344
66 | 16199 | 759.3281
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16200 | 759.375
66 | 16202 | 759.4688
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16191 | 758.9531
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16196 | 759.1875
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16196 | 759.1875
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16194 | 759.0938
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16195 | 759.1406
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16194 | 759.0938
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16195 | 759.1406
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16194 | 759.0938
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16195 | 759.1406
66 | 16199 | 759.3281
66 | 16201 | 759.4219
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16194 | 759.0938
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16195 | 759.1406
66 | 16190 | 758.9062
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16195 | 759.1406
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16196 | 759.1875
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16195 | 759.1406
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16200 | 759.375
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16199 | 759.3281
66 | 16199 | 759.3281
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16194 | 759.0938
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16196 | 759.1875
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16197 | 759.2344
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16198 | 759.2812
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16192 | 759
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16196 | 759.1875
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16195 | 759.1406
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16192 | 759
66 | 16192 | 759
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16192 | 759
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16192 | 759
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16192 | 759
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16183 | 758.5781
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16183 | 758.5781
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16195 | 759.1406
66 | 16196 | 759.1875
66 | 16195 | 759.1406
66 | 16195 | 759.1406
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16188 | 758.8125
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16195 | 759.1406
66 | 16195 | 759.1406
66 | 16194 | 759.0938
66 | 16189 | 758.8594
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16195 | 759.1406
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16193 | 759.0469
66 | 16194 | 759.0938
66 | 16194 | 759.0938
66 | 16188 | 758.8125
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16187 | 758.7656
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16185 | 758.6719
66 | 16181 | 758.4844
66 | 16182 | 758.5312
66 | 16186 | 758.7188
66 | 16192 | 759
66 | 16194 | 759.0938
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16182 | 758.5312
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16192 | 759
66 | 16189 | 758.8594
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16193 | 759.0469
66 | 16195 | 759.1406
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16181 | 758.4844
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16189 | 758.8594
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16194 | 759.0938
66 | 16193 | 759.0469
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16185 | 758.6719
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16183 | 758.5781
66 | 16183 | 758.5781
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16192 | 759
66 | 16192 | 759
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16184 | 758.625
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16185 | 758.6719
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16193 | 759.0469
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16184 | 758.625
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16190 | 758.9062
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16183 | 758.5781
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16187 | 758.7656
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16189 | 758.8594
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16186 | 758.7188
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16189 | 758.8594
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16192 | 759
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16183 | 758.5781
66 | 16181 | 758.4844
66 | 16187 | 758.7656
66 | 16191 | 758.9531
66 | 16193 | 759.0469
66 | 16188 | 758.8125
66 | 16181 | 758.4844
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16184 | 758.625
66 | 16189 | 758.8594
66 | 16193 | 759.0469
66 | 16190 | 758.9062
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16181 | 758.4844
66 | 16180 | 758.4375
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16189 | 758.8594
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16181 | 758.4844
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16182 | 758.5312
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16183 | 758.5781
66 | 16182 | 758.5312
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16193 | 759.0469
66 | 16183 | 758.5781
66 | 16179 | 758.3906
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16187 | 758.7656
66 | 16192 | 759
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16183 | 758.5781
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16194 | 759.0938
66 | 16185 | 758.6719
66 | 16181 | 758.4844
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16186 | 758.7188
66 | 16181 | 758.4844
66 | 16183 | 758.5781
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16184 | 758.625
66 | 16189 | 758.8594
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16187 | 758.7656
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16192 | 759
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16181 | 758.4844
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16182 | 758.5312
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16187 | 758.7656
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16192 | 759
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16188 | 758.8125
66 | 16190 | 758.9062
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16184 | 758.625
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16192 | 759
66 | 16182 | 758.5312
66 | 16182 | 758.5312
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16181 | 758.4844
66 | 16182 | 758.5312
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16186 | 758.7188
66 | 16192 | 759
66 | 16191 | 758.9531
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16190 | 758.9062
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16191 | 758.9531
66 | 16181 | 758.4844
66 | 16185 | 758.6719
66 | 16191 | 758.9531
66 | 16191 | 758.9531
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16189 | 758.8594
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16181 | 758.4844
66 | 16181 | 758.4844
66 | 16187 | 758.7656
66 | 16192 | 759
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16182 | 758.5312
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16191 | 758.9531
66 | 16182 | 758.5312
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16191 | 758.9531
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16192 | 759
66 | 16186 | 758.7188
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16183 | 758.5781
66 | 16189 | 758.8594
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16181 | 758.4844
66 | 16187 | 758.7656
66 | 16191 | 758.9531
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16182 | 758.5312
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16179 | 758.3906
66 | 16182 | 758.5312
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16190 | 758.9062
66 | 16190 | 758.9062
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16182 | 758.5312
66 | 16190 | 758.9062
66 | 16177 | 758.2969
66 | 16182 | 758.5312
66 | 16189 | 758.8594
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16183 | 758.5781
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16190 | 758.9062
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16180 | 758.4375
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16190 | 758.9062
66 | 16180 | 758.4375
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16191 | 758.9531
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16190 | 758.9062
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16189 | 758.8594
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16186 | 758.7188
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16189 | 758.8594
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16189 | 758.8594
66 | 16188 | 758.8125
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16186 | 758.7188
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16187 | 758.7656
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16188 | 758.8125
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16188 | 758.8125
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16188 | 758.8125
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16187 | 758.7656
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16187 | 758.7656
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16182 | 758.5312
66 | 16180 | 758.4375
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16172 | 758.0625
66 | 16184 | 758.625
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16187 | 758.7656
66 | 16174 | 758.1562
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16187 | 758.7656
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16178 | 758.3438
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16185 | 758.6719
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16185 | 758.6719
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16185 | 758.6719
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16178 | 758.3438
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16181 | 758.4844
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16179 | 758.3906
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16173 | 758.1094
66 | 16186 | 758.7188
66 | 16174 | 758.1562
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16177 | 758.2969
66 | 16186 | 758.7188
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16177 | 758.2969
66 | 16179 | 758.3906
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16183 | 758.5781
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16183 | 758.5781
66 | 16178 | 758.3438
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16185 | 758.6719
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16185 | 758.6719
66 | 16176 | 758.25
66 | 16185 | 758.6719
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16186 | 758.7188
66 | 16176 | 758.25
66 | 16184 | 758.625
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16182 | 758.5312
66 | 16178 | 758.3438
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16181 | 758.4844
66 | 16180 | 758.4375
66 | 16182 | 758.5312
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16179 | 758.3906
66 | 16180 | 758.4375
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16179 | 758.3906
66 | 16174 | 758.1562
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16175 | 758.2031
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16184 | 758.625
66 | 16175 | 758.2031
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16183 | 758.5781
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16179 | 758.3906
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16168 | 757.875
66 | 16180 | 758.4375
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16182 | 758.5312
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16183 | 758.5781
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16168 | 757.875
66 | 16181 | 758.4844
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16183 | 758.5781
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16176 | 758.25
66 | 16173 | 758.1094
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16172 | 758.0625
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16174 | 758.1562
66 | 16177 | 758.2969
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16180 | 758.4375
66 | 16179 | 758.3906
66 | 16174 | 758.1562
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16172 | 758.0625
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16171 | 758.0156
66 | 16183 | 758.5781
66 | 16173 | 758.1094
66 | 16177 | 758.2969
66 | 16180 | 758.4375
66 | 16173 | 758.1094
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16173 | 758.1094
66 | 16179 | 758.3906
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16182 | 758.5312
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16176 | 758.25
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16183 | 758.5781
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16183 | 758.5781
66 | 16170 | 757.9688
66 | 16182 | 758.5312
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16168 | 757.875
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16181 | 758.4844
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16171 | 758.0156
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16172 | 758.0625
66 | 16183 | 758.5781
66 | 16170 | 757.9688
66 | 16182 | 758.5312
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16171 | 758.0156
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16172 | 758.0625
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16179 | 758.3906
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16178 | 758.3438
66 | 16180 | 758.4375
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16181 | 758.4844
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16172 | 758.0625
66 | 16170 | 757.9688
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16171 | 758.0156
66 | 16182 | 758.5312
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16182 | 758.5312
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16172 | 758.0625
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16172 | 758.0625
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16181 | 758.4844
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16181 | 758.4844
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16181 | 758.4844
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16177 | 758.2969
66 | 16172 | 758.0625
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16170 | 757.9688
66 | 16181 | 758.4844
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16172 | 758.0625
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16172 | 758.0625
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16179 | 758.3906
66 | 16172 | 758.0625
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16172 | 758.0625
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16174 | 758.1562
66 | 16164 | 757.6875
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16180 | 758.4375
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16171 | 758.0156
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16173 | 758.1094
66 | 16169 | 757.9219
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16181 | 758.4844
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16180 | 758.4375
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16172 | 758.0625
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16180 | 758.4375
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16175 | 758.2031
66 | 16165 | 757.7344
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16179 | 758.3906
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16176 | 758.25
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16178 | 758.3438
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16176 | 758.25
66 | 16165 | 757.7344
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16179 | 758.3906
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16179 | 758.3906
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16167 | 757.8281
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16163 | 757.6406
66 | 16174 | 758.1562
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16171 | 758.0156
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16175 | 758.2031
66 | 16179 | 758.3906
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16173 | 758.1094
66 | 16178 | 758.3438
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16170 | 757.9688
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16164 | 757.6875
66 | 16173 | 758.1094
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16178 | 758.3438
66 | 16167 | 757.8281
66 | 16172 | 758.0625
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16177 | 758.2969
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16165 | 757.7344
66 | 16170 | 757.9688
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16176 | 758.25
66 | 16178 | 758.3438
66 | 16169 | 757.9219
66 | 16171 | 758.0156
66 | 16178 | 758.3438
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16173 | 758.1094
66 | 16176 | 758.25
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16172 | 758.0625
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16170 | 757.9688
66 | 16176 | 758.25
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16171 | 758.0156
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16174 | 758.1562
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16168 | 757.875
66 | 16173 | 758.1094
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16171 | 758.0156
66 | 16175 | 758.2031
66 | 16172 | 758.0625
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16177 | 758.2969
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16177 | 758.2969
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16178 | 758.3438
66 | 16177 | 758.2969
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16176 | 758.25
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16175 | 758.2031
66 | 16178 | 758.3438
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16171 | 758.0156
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16171 | 758.0156
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16171 | 758.0156
66 | 16166 | 757.7812
66 | 16171 | 758.0156
66 | 16178 | 758.3438
66 | 16176 | 758.25
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16178 | 758.3438
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16174 | 758.1562
66 | 16177 | 758.2969
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16170 | 757.9688
66 | 16177 | 758.2969
66 | 16176 | 758.25
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16172 | 758.0625
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16170 | 757.9688
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16177 | 758.2969
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16171 | 758.0156
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16170 | 757.9688
66 | 16177 | 758.2969
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16174 | 758.1562
66 | 16172 | 758.0625
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16172 | 758.0625
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16169 | 757.9219
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16176 | 758.25
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16170 | 757.9688
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16174 | 758.1562
66 | 16176 | 758.25
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16171 | 758.0156
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16170 | 757.9688
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16171 | 758.0156
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16171 | 758.0156
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16175 | 758.2031
66 | 16176 | 758.25
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16172 | 758.0625
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16175 | 758.2031
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16172 | 758.0625
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16172 | 758.0625
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16169 | 757.9219
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16171 | 758.0156
66 | 16174 | 758.1562
66 | 16175 | 758.2031
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16169 | 757.9219
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16171 | 758.0156
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16167 | 757.8281
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16172 | 758.0625
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16170 | 757.9688
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16167 | 757.8281
66 | 16173 | 758.1094
66 | 16175 | 758.2031
66 | 16174 | 758.1562
66 | 16171 | 758.0156
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16170 | 757.9688
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16170 | 757.9688
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16169 | 757.9219
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16167 | 757.8281
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16169 | 757.9219
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16169 | 757.9219
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16170 | 757.9688
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16172 | 758.0625
66 | 16176 | 758.25
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16169 | 757.9219
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16174 | 758.1562
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16167 | 757.8281
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16171 | 758.0156
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16169 | 757.9219
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16170 | 757.9688
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16168 | 757.875
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16169 | 757.9219
66 | 16171 | 758.0156
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16167 | 757.8281
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16174 | 758.1562
66 | 16174 | 758.1562
66 | 16173 | 758.1094
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16172 | 758.0625
66 | 16172 | 758.0625
66 | 16173 | 758.1094
66 | 16173 | 758.1094
66 | 16172 | 758.0625
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16170 | 757.9688
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16171 | 758.0156
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16172 | 758.0625
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16172 | 758.0625
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16171 | 758.0156
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16171 | 758.0156
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16160 | 757.5
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16170 | 757.9688
66 | 16170 | 757.9688
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16169 | 757.9219
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16162 | 757.5938
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16170 | 757.9688
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16162 | 757.5938
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16156 | 757.3125
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16166 | 757.7812
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16160 | 757.5
66 | 16155 | 757.2656
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16155 | 757.2656
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16169 | 757.9219
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16158 | 757.4062
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16162 | 757.5938
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16156 | 757.3125
66 | 16158 | 757.4062
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16162 | 757.5938
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16158 | 757.4062
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16164 | 757.6875
66 | 16156 | 757.3125
66 | 16161 | 757.5469
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16169 | 757.9219
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16169 | 757.9219
66 | 16167 | 757.8281
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16169 | 757.9219
66 | 16168 | 757.875
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16165 | 757.7344
66 | 16168 | 757.875
66 | 16167 | 757.8281
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16158 | 757.4062
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16166 | 757.7812
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16166 | 757.7812
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16168 | 757.875
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16168 | 757.875
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16164 | 757.6875
66 | 16167 | 757.8281
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16160 | 757.5
66 | 16156 | 757.3125
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16156 | 757.3125
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16157 | 757.3594
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16156 | 757.3125
66 | 16156 | 757.3125
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16165 | 757.7344
66 | 16167 | 757.8281
66 | 16164 | 757.6875
66 | 16155 | 757.2656
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16154 | 757.2188
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16167 | 757.8281
66 | 16167 | 757.8281
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16160 | 757.5
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16168 | 757.875
66 | 16162 | 757.5938
66 | 16156 | 757.3125
66 | 16158 | 757.4062
66 | 16166 | 757.7812
66 | 16167 | 757.8281
66 | 16157 | 757.3594
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16162 | 757.5938
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16166 | 757.7812
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16154 | 757.2188
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16160 | 757.5
66 | 16153 | 757.1719
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16164 | 757.6875
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16154 | 757.2188
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16153 | 757.1719
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16158 | 757.4062
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16161 | 757.5469
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16167 | 757.8281
66 | 16162 | 757.5938
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16161 | 757.5469
66 | 16154 | 757.2188
66 | 16158 | 757.4062
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16166 | 757.7812
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16155 | 757.2656
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16164 | 757.6875
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16166 | 757.7812
66 | 16166 | 757.7812
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16165 | 757.7344
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16165 | 757.7344
66 | 16165 | 757.7344
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16166 | 757.7812
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16159 | 757.4531
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16159 | 757.4531
66 | 16154 | 757.2188
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16164 | 757.6875
66 | 16156 | 757.3125
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16166 | 757.7812
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16165 | 757.7344
66 | 16166 | 757.7812
66 | 16152 | 757.125
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16154 | 757.2188
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16166 | 757.7812
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16156 | 757.3125
66 | 16165 | 757.7344
66 | 16164 | 757.6875
66 | 16154 | 757.2188
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16155 | 757.2656
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16163 | 757.6406
66 | 16165 | 757.7344
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16151 | 757.0781
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16164 | 757.6875
66 | 16165 | 757.7344
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16154 | 757.2188
66 | 16154 | 757.2188
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16154 | 757.2188
66 | 16152 | 757.125
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16157 | 757.3594
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16164 | 757.6875
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16150 | 757.0312
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16157 | 757.3594
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16159 | 757.4531
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16157 | 757.3594
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16164 | 757.6875
66 | 16155 | 757.2656
66 | 16158 | 757.4062
66 | 16164 | 757.6875
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16161 | 757.5469
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16151 | 757.0781
66 | 16162 | 757.5938
66 | 16159 | 757.4531
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16163 | 757.6406
66 | 16164 | 757.6875
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16164 | 757.6875
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16162 | 757.5938
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16164 | 757.6875
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16163 | 757.6406
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16164 | 757.6875
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16163 | 757.6406
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16158 | 757.4062
66 | 16163 | 757.6406
66 | 16154 | 757.2188
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16151 | 757.0781
66 | 16163 | 757.6406
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16162 | 757.5938
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16154 | 757.2188
66 | 16159 | 757.4531
66 | 16163 | 757.6406
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16156 | 757.3125
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16164 | 757.6875
66 | 16161 | 757.5469
66 | 16153 | 757.1719
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16152 | 757.125
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16162 | 757.5938
66 | 16155 | 757.2656
66 | 16153 | 757.1719
66 | 16162 | 757.5938
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16162 | 757.5938
66 | 16163 | 757.6406
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16160 | 757.5
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16160 | 757.5
66 | 16151 | 757.0781
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16148 | 756.9375
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16151 | 757.0781
66 | 16158 | 757.4062
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16151 | 757.0781
66 | 16156 | 757.3125
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16162 | 757.5938
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16147 | 756.8906
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16161 | 757.5469
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16161 | 757.5469
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16153 | 757.1719
66 | 16152 | 757.125
66 | 16159 | 757.4531
66 | 16155 | 757.2656
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16146 | 756.8438
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16161 | 757.5469
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16162 | 757.5938
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16151 | 757.0781
66 | 16161 | 757.5469
66 | 16162 | 757.5938
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16149 | 756.9844
66 | 16161 | 757.5469
66 | 16159 | 757.4531
66 | 16149 | 756.9844
66 | 16161 | 757.5469
66 | 16161 | 757.5469
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16156 | 757.3125
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16162 | 757.5938
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16161 | 757.5469
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16152 | 757.125
66 | 16160 | 757.5
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16153 | 757.1719
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16154 | 757.2188
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16155 | 757.2656
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16157 | 757.3594
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16160 | 757.5
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16160 | 757.5
66 | 16148 | 756.9375
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16162 | 757.5938
66 | 16148 | 756.9375
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16160 | 757.5
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16160 | 757.5
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16146 | 756.8438
66 | 16157 | 757.3594
66 | 16152 | 757.125
66 | 16148 | 756.9375
66 | 16160 | 757.5
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16161 | 757.5469
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16154 | 757.2188
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16156 | 757.3125
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16149 | 756.9844
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16161 | 757.5469
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16160 | 757.5
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16147 | 756.8906
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16154 | 757.2188
66 | 16152 | 757.125
66 | 16161 | 757.5469
66 | 16152 | 757.125
66 | 16156 | 757.3125
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16160 | 757.5
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16152 | 757.125
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16161 | 757.5469
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16160 | 757.5
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16146 | 756.8438
66 | 16158 | 757.4062
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16160 | 757.5
66 | 16156 | 757.3125
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16156 | 757.3125
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16151 | 757.0781
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16147 | 756.8906
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16148 | 756.9375
66 | 16160 | 757.5
66 | 16156 | 757.3125
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16158 | 757.4062
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16158 | 757.4062
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16158 | 757.4062
66 | 16156 | 757.3125
66 | 16149 | 756.9844
66 | 16159 | 757.4531
66 | 16157 | 757.3594
66 | 16147 | 756.8906
66 | 16159 | 757.4531
66 | 16156 | 757.3125
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16158 | 757.4062
66 | 16148 | 756.9375
66 | 16152 | 757.125
66 | 16156 | 757.3125
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16156 | 757.3125
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16154 | 757.2188
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16146 | 756.8438
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16154 | 757.2188
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16150 | 757.0312
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16157 | 757.3594
66 | 16159 | 757.4531
66 | 16148 | 756.9375
66 | 16158 | 757.4062
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16156 | 757.3125
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16144 | 756.75
66 | 16156 | 757.3125
66 | 16150 | 757.0312
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16159 | 757.4531
66 | 16147 | 756.8906
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16146 | 756.8438
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16156 | 757.3125
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16157 | 757.3594
66 | 16151 | 757.0781
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16158 | 757.4062
66 | 16155 | 757.2656
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16153 | 757.1719
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16146 | 756.8438
66 | 16157 | 757.3594
66 | 16156 | 757.3125
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16156 | 757.3125
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16146 | 756.8438
66 | 16157 | 757.3594
66 | 16157 | 757.3594
66 | 16147 | 756.8906
66 | 16157 | 757.3594
66 | 16154 | 757.2188
66 | 16147 | 756.8906
66 | 16158 | 757.4062
66 | 16153 | 757.1719
66 | 16146 | 756.8438
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16151 | 757.0781
66 | 16156 | 757.3125
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16158 | 757.4062
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16156 | 757.3125
66 | 16155 | 757.2656
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16146 | 756.8438
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16154 | 757.2188
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16146 | 756.8438
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16157 | 757.3594
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16144 | 756.75
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16144 | 756.75
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16157 | 757.3594
66 | 16146 | 756.8438
66 | 16151 | 757.0781
66 | 16156 | 757.3125
66 | 16143 | 756.7031
66 | 16156 | 757.3125
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16154 | 757.2188
66 | 16151 | 757.0781
66 | 16147 | 756.8906
66 | 16154 | 757.2188
66 | 16146 | 756.8438
66 | 16152 | 757.125
66 | 16156 | 757.3125
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16156 | 757.3125
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16144 | 756.75
66 | 16154 | 757.2188
66 | 16155 | 757.2656
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16157 | 757.3594
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16147 | 756.8906
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16156 | 757.3125
66 | 16156 | 757.3125
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16146 | 756.8438
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16142 | 756.6562
66 | 16149 | 756.9844
66 | 16153 | 757.1719
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16156 | 757.3125
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16152 | 757.125
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16147 | 756.8906
66 | 16155 | 757.2656
66 | 16149 | 756.9844
66 | 16144 | 756.75
66 | 16154 | 757.2188
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16152 | 757.125
66 | 16147 | 756.8906
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16144 | 756.75
66 | 16155 | 757.2656
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16157 | 757.3594
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16142 | 756.6562
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16154 | 757.2188
66 | 16155 | 757.2656
66 | 16143 | 756.7031
66 | 16154 | 757.2188
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16154 | 757.2188
66 | 16155 | 757.2656
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16148 | 756.9375
66 | 16140 | 756.5625
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16156 | 757.3125
66 | 16153 | 757.1719
66 | 16146 | 756.8438
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16152 | 757.125
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16152 | 757.125
66 | 16157 | 757.3594
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16152 | 757.125
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16154 | 757.2188
66 | 16152 | 757.125
66 | 16144 | 756.75
66 | 16155 | 757.2656
66 | 16150 | 757.0312
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16156 | 757.3125
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16154 | 757.2188
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16151 | 757.0781
66 | 16143 | 756.7031
66 | 16155 | 757.2656
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16153 | 757.1719
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16154 | 757.2188
66 | 16156 | 757.3125
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16156 | 757.3125
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16144 | 756.75
66 | 16152 | 757.125
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16156 | 757.3125
66 | 16148 | 756.9375
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16140 | 756.5625
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16148 | 756.9375
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16151 | 757.0781
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16155 | 757.2656
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16148 | 756.9375
66 | 16142 | 756.6562
66 | 16153 | 757.1719
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16146 | 756.8438
66 | 16152 | 757.125
66 | 16156 | 757.3125
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16154 | 757.2188
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16141 | 756.6094
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16141 | 756.6094
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16153 | 757.1719
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16154 | 757.2188
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16141 | 756.6094
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16154 | 757.2188
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16142 | 756.6562
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16152 | 757.125
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16154 | 757.2188
66 | 16150 | 757.0312
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16153 | 757.1719
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16152 | 757.125
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16154 | 757.2188
66 | 16149 | 756.9844
66 | 16143 | 756.7031
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16148 | 756.9375
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16155 | 757.2656
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16142 | 756.6562
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16141 | 756.6094
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16141 | 756.6094
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16155 | 757.2656
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16155 | 757.2656
66 | 16148 | 756.9375
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16150 | 757.0312
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16141 | 756.6094
66 | 16147 | 756.8906
66 | 16154 | 757.2188
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16151 | 757.0781
66 | 16151 | 757.0781
66 | 16140 | 756.5625
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16154 | 757.2188
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16152 | 757.125
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16149 | 756.9844
66 | 16142 | 756.6562
66 | 16153 | 757.1719
66 | 16154 | 757.2188
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16150 | 757.0312
66 | 16141 | 756.6094
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16151 | 757.0781
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16153 | 757.1719
66 | 16151 | 757.0781
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16152 | 757.125
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16152 | 757.125
66 | 16154 | 757.2188
66 | 16141 | 756.6094
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16153 | 757.1719
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16153 | 757.1719
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16141 | 756.6094
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16139 | 756.5156
66 | 16150 | 757.0312
66 | 16153 | 757.1719
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16152 | 757.125
66 | 16147 | 756.8906
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16151 | 757.0781
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16150 | 757.0312
66 | 16149 | 756.9844
66 | 16141 | 756.6094
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16152 | 757.125
66 | 16153 | 757.1719
66 | 16139 | 756.5156
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16147 | 756.8906
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16152 | 757.125
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16138 | 756.4688
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16152 | 757.125
66 | 16150 | 757.0312
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16151 | 757.0781
66 | 16153 | 757.1719
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16149 | 756.9844
66 | 16138 | 756.4688
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16138 | 756.4688
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16151 | 757.0781
66 | 16151 | 757.0781
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16142 | 756.6562
66 | 16137 | 756.4219
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16138 | 756.4688
66 | 16145 | 756.7969
66 | 16153 | 757.1719
66 | 16147 | 756.8906
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16149 | 756.9844
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16143 | 756.7031
66 | 16137 | 756.4219
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16137 | 756.4219
66 | 16141 | 756.6094
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16151 | 757.0781
66 | 16146 | 756.8438
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16149 | 756.9844
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16151 | 757.0781
66 | 16148 | 756.9375
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16152 | 757.125
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16152 | 757.125
66 | 16148 | 756.9375
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16151 | 757.0781
66 | 16152 | 757.125
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16151 | 757.0781
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16146 | 756.8438
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16147 | 756.8906
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16150 | 757.0312
66 | 16149 | 756.9844
66 | 16138 | 756.4688
66 | 16142 | 756.6562
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16151 | 757.0781
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16143 | 756.7031
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16136 | 756.375
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16136 | 756.375
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16147 | 756.8906
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16146 | 756.8438
66 | 16137 | 756.4219
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16136 | 756.375
66 | 16142 | 756.6562
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16143 | 756.7031
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16149 | 756.9844
66 | 16151 | 757.0781
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16137 | 756.4219
66 | 16145 | 756.7969
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16149 | 756.9844
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16149 | 756.9844
66 | 16152 | 757.125
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16144 | 756.75
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16150 | 757.0312
66 | 16148 | 756.9375
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16150 | 757.0312
66 | 16151 | 757.0781
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16149 | 756.9844
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16151 | 757.0781
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16144 | 756.75
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16150 | 757.0312
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16142 | 756.6562
66 | 16136 | 756.375
66 | 16141 | 756.6094
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16142 | 756.6562
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16136 | 756.375
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16150 | 757.0312
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16149 | 756.9844
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16143 | 756.7031
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16140 | 756.5625
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16136 | 756.375
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16149 | 756.9844
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16141 | 756.6094
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16141 | 756.6094
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16147 | 756.8906
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16138 | 756.4688
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16149 | 756.9844
66 | 16147 | 756.8906
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16145 | 756.7969
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16146 | 756.8438
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16150 | 757.0312
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16148 | 756.9375
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16149 | 756.9844
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16148 | 756.9375
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16138 | 756.4688
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16147 | 756.8906
66 | 16148 | 756.9375
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16147 | 756.8906
66 | 16147 | 756.8906
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16148 | 756.9375
66 | 16147 | 756.8906
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16147 | 756.8906
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16138 | 756.4688
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16144 | 756.75
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16146 | 756.8438
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16138 | 756.4688
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16138 | 756.4688
66 | 16143 | 756.7031
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16138 | 756.4688
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16137 | 756.4219
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16140 | 756.5625
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16129 | 756.0469
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16134 | 756.2812
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16146 | 756.8438
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16145 | 756.7969
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16142 | 756.6562
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16141 | 756.6094
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16141 | 756.6094
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16143 | 756.7031
66 | 16144 | 756.75
66 | 16143 | 756.7031
66 | 16143 | 756.7031
66 | 16142 | 756.6562
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16137 | 756.4219
66 | 16129 | 756.0469
66 | 16138 | 756.4688
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16140 | 756.5625
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16137 | 756.4219
66 | 16132 | 756.1875
66 | 16137 | 756.4219
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16133 | 756.2344
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16139 | 756.5156
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16139 | 756.5156
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16138 | 756.4688
66 | 16138 | 756.4688
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16131 | 756.1406
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16138 | 756.4688
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16139 | 756.5156
66 | 16133 | 756.2344
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16138 | 756.4688
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16137 | 756.4219
66 | 16135 | 756.3281
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16135 | 756.3281
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16138 | 756.4688
66 | 16137 | 756.4219
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16136 | 756.375
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16137 | 756.4219
66 | 16136 | 756.375
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16136 | 756.375
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16137 | 756.4219
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16137 | 756.4219
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16136 | 756.375
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16135 | 756.3281
66 | 16120 | 755.625
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16135 | 756.3281
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16136 | 756.375
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16135 | 756.3281
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16137 | 756.4219
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16135 | 756.3281
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16136 | 756.375
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16136 | 756.375
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16135 | 756.3281
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16134 | 756.2812
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16133 | 756.2344
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16136 | 756.375
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16132 | 756.1875
66 | 16125 | 755.8594
66 | 16135 | 756.3281
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16133 | 756.2344
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16133 | 756.2344
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16133 | 756.2344
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16133 | 756.2344
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16135 | 756.3281
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16134 | 756.2812
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16134 | 756.2812
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16132 | 756.1875
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16134 | 756.2812
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16132 | 756.1875
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16134 | 756.2812
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16132 | 756.1875
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16118 | 755.5312
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16133 | 756.2344
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16132 | 756.1875
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16133 | 756.2344
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16134 | 756.2812
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16131 | 756.1406
66 | 16118 | 755.5312
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16123 | 755.7656
66 | 16133 | 756.2344
66 | 16123 | 755.7656
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16128 | 756
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16119 | 755.5781
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16131 | 756.1406
66 | 16124 | 755.8125
66 | 16130 | 756.0938
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16131 | 756.1406
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16115 | 755.3906
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16130 | 756.0938
66 | 16117 | 755.4844
66 | 16131 | 756.1406
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16115 | 755.3906
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16131 | 756.1406
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16129 | 756.0469
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16132 | 756.1875
66 | 16121 | 755.6719
66 | 16131 | 756.1406
66 | 16129 | 756.0469
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16123 | 755.7656
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16131 | 756.1406
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16131 | 756.1406
66 | 16126 | 755.9062
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16125 | 755.8594
66 | 16115 | 755.3906
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16116 | 755.4375
66 | 16129 | 756.0469
66 | 16124 | 755.8125
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16130 | 756.0938
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16114 | 755.3438
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16126 | 755.9062
66 | 16115 | 755.3906
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16115 | 755.3906
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16115 | 755.3906
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16130 | 756.0938
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16114 | 755.3438
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16114 | 755.3438
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16130 | 756.0938
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16116 | 755.4375
66 | 16129 | 756.0469
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16115 | 755.3906
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16129 | 756.0469
66 | 16125 | 755.8594
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16128 | 756
66 | 16128 | 756
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16117 | 755.4844
66 | 16125 | 755.8594
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16128 | 756
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16129 | 756.0469
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16128 | 756
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16128 | 756
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16121 | 755.6719
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16129 | 756.0469
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16128 | 756
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16121 | 755.6719
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16123 | 755.7656
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16121 | 755.6719
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16127 | 755.9531
66 | 16128 | 756
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16127 | 755.9531
66 | 16129 | 756.0469
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16127 | 755.9531
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16127 | 755.9531
66 | 16126 | 755.9062
66 | 16128 | 756
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16128 | 756
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16127 | 755.9531
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16126 | 755.9062
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16126 | 755.9062
66 | 16125 | 755.8594
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16124 | 755.8125
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16125 | 755.8594
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16125 | 755.8594
66 | 16120 | 755.625
66 | 16124 | 755.8125
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16125 | 755.8594
66 | 16115 | 755.3906
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16124 | 755.8125
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16124 | 755.8125
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16124 | 755.8125
66 | 16123 | 755.7656
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16113 | 755.2969
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16123 | 755.7656
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16113 | 755.2969
66 | 16122 | 755.7188
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16122 | 755.7188
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16123 | 755.7656
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16122 | 755.7188
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16123 | 755.7656
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16113 | 755.2969
66 | 16110 | 755.1562
66 | 16120 | 755.625
66 | 16122 | 755.7188
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16123 | 755.7656
66 | 16122 | 755.7188
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16121 | 755.6719
66 | 16122 | 755.7188
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16122 | 755.7188
66 | 16123 | 755.7656
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16121 | 755.6719
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16122 | 755.7188
66 | 16115 | 755.3906
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16121 | 755.6719
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16121 | 755.6719
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16120 | 755.625
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16121 | 755.6719
66 | 16113 | 755.2969
66 | 16110 | 755.1562
66 | 16121 | 755.6719
66 | 16120 | 755.625
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16120 | 755.625
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16122 | 755.7188
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16122 | 755.7188
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16122 | 755.7188
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16121 | 755.6719
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16121 | 755.6719
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16120 | 755.625
66 | 16121 | 755.6719
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16122 | 755.7188
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16120 | 755.625
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16121 | 755.6719
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16120 | 755.625
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16120 | 755.625
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16122 | 755.7188
66 | 16107 | 755.0156
66 | 16120 | 755.625
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16120 | 755.625
66 | 16110 | 755.1562
66 | 16120 | 755.625
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16121 | 755.6719
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16120 | 755.625
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16120 | 755.625
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16120 | 755.625
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16121 | 755.6719
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16120 | 755.625
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16120 | 755.625
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16119 | 755.5781
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16118 | 755.5312
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16102 | 754.7812
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16117 | 755.4844
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16119 | 755.5781
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16119 | 755.5781
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16119 | 755.5781
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16111 | 755.2031
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16119 | 755.5781
66 | 16107 | 755.0156
66 | 16118 | 755.5312
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16106 | 754.9688
66 | 16119 | 755.5781
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16102 | 754.7812
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16103 | 754.8281
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16118 | 755.5312
66 | 16115 | 755.3906
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16104 | 754.875
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16110 | 755.1562
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16104 | 754.875
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16119 | 755.5781
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16103 | 754.8281
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16116 | 755.4375
66 | 16113 | 755.2969
66 | 16105 | 754.9219
66 | 16113 | 755.2969
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16117 | 755.4844
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16118 | 755.5312
66 | 16107 | 755.0156
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16113 | 755.2969
66 | 16118 | 755.5312
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16104 | 754.875
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16110 | 755.1562
66 | 16117 | 755.4844
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16117 | 755.4844
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16117 | 755.4844
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16116 | 755.4375
66 | 16118 | 755.5312
66 | 16113 | 755.2969
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16118 | 755.5312
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16112 | 755.25
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16107 | 755.0156
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16117 | 755.4844
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16116 | 755.4375
66 | 16116 | 755.4375
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16115 | 755.3906
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16115 | 755.3906
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16113 | 755.2969
66 | 16116 | 755.4375
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16115 | 755.3906
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16112 | 755.25
66 | 16115 | 755.3906
66 | 16116 | 755.4375
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16116 | 755.4375
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16114 | 755.3438
66 | 16114 | 755.3438
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16114 | 755.3438
66 | 16111 | 755.2031
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16114 | 755.3438
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16111 | 755.2031
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16111 | 755.2031
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16113 | 755.2969
66 | 16113 | 755.2969
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16112 | 755.25
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16113 | 755.2969
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16111 | 755.2031
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16111 | 755.2031
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16113 | 755.2969
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16112 | 755.25
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16112 | 755.25
66 | 16110 | 755.1562
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16111 | 755.2031
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16110 | 755.1562
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16110 | 755.1562
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16111 | 755.2031
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16112 | 755.25
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16111 | 755.2031
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16111 | 755.2031
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16111 | 755.2031
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16112 | 755.25
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16111 | 755.2031
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16110 | 755.1562
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16110 | 755.1562
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16102 | 754.7812
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16109 | 755.1094
66 | 16103 | 754.8281
66 | 16110 | 755.1562
66 | 16097 | 754.5469
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16110 | 755.1562
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16109 | 755.1094
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16110 | 755.1562
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16101 | 754.7344
66 | 16108 | 755.0625
66 | 16101 | 754.7344
66 | 16109 | 755.1094
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16109 | 755.1094
66 | 16100 | 754.6875
66 | 16107 | 755.0156
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16109 | 755.1094
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16109 | 755.1094
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16093 | 754.3594
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16097 | 754.5469
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16109 | 755.1094
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16109 | 755.1094
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16108 | 755.0625
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16093 | 754.3594
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16093 | 754.3594
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16092 | 754.3125
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16100 | 754.6875
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16099 | 754.6406
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16092 | 754.3125
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16092 | 754.3125
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16107 | 755.0156
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16108 | 755.0625
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16099 | 754.6406
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16107 | 755.0156
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16092 | 754.3125
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16092 | 754.3125
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16107 | 755.0156
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16093 | 754.3594
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16096 | 754.5
66 | 16107 | 755.0156
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16093 | 754.3594
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16105 | 754.9219
66 | 16104 | 754.875
66 | 16092 | 754.3125
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16093 | 754.3594
66 | 16105 | 754.9219
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16106 | 754.9688
66 | 16107 | 755.0156
66 | 16095 | 754.4531
66 | 16106 | 754.9688
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16092 | 754.3125
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16107 | 755.0156
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16097 | 754.5469
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16107 | 755.0156
66 | 16096 | 754.5
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16108 | 755.0625
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16102 | 754.7812
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16102 | 754.7812
66 | 16107 | 755.0156
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16096 | 754.5
66 | 16101 | 754.7344
66 | 16105 | 754.9219
66 | 16098 | 754.5938
66 | 16094 | 754.4062
66 | 16105 | 754.9219
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16106 | 754.9688
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16099 | 754.6406
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16092 | 754.3125
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16106 | 754.9688
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16106 | 754.9688
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16092 | 754.3125
66 | 16096 | 754.5
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16092 | 754.3125
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16106 | 754.9688
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16092 | 754.3125
66 | 16096 | 754.5
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16099 | 754.6406
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16090 | 754.2188
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16097 | 754.5469
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16104 | 754.875
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16098 | 754.5938
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16099 | 754.6406
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16098 | 754.5938
66 | 16091 | 754.2656
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16089 | 754.1719
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16100 | 754.6875
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16094 | 754.4062
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16099 | 754.6406
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16097 | 754.5469
66 | 16092 | 754.3125
66 | 16097 | 754.5469
66 | 16103 | 754.8281
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16096 | 754.5
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16105 | 754.9219
66 | 16096 | 754.5
66 | 16092 | 754.3125
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16098 | 754.5938
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16095 | 754.4531
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16090 | 754.2188
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16090 | 754.2188
66 | 16098 | 754.5938
66 | 16103 | 754.8281
66 | 16103 | 754.8281
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16104 | 754.875
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16089 | 754.1719
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16095 | 754.4531
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16092 | 754.3125
66 | 16096 | 754.5
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16097 | 754.5469
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16094 | 754.4062
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16104 | 754.875
66 | 16103 | 754.8281
66 | 16096 | 754.5
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16094 | 754.4062
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16104 | 754.875
66 | 16096 | 754.5
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16097 | 754.5469
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16097 | 754.5469
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16103 | 754.8281
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16094 | 754.4062
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16097 | 754.5469
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16097 | 754.5469
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16094 | 754.4062
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16096 | 754.5
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16094 | 754.4062
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16097 | 754.5469
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16102 | 754.7812
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16102 | 754.7812
66 | 16102 | 754.7812
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16101 | 754.7344
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16101 | 754.7344
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16100 | 754.6875
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16100 | 754.6875
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16099 | 754.6406
66 | 16098 | 754.5938
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16098 | 754.5938
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16091 | 754.2656
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16097 | 754.5469
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16086 | 754.0312
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16092 | 754.3125
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16097 | 754.5469
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16096 | 754.5
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16094 | 754.4062
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16092 | 754.3125
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16095 | 754.4531
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16095 | 754.4531
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16094 | 754.4062
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16092 | 754.3125
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16094 | 754.4062
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16093 | 754.3594
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16092 | 754.3125
66 | 16094 | 754.4062
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16093 | 754.3594
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16093 | 754.3594
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16093 | 754.3594
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16092 | 754.3125
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16091 | 754.2656
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16091 | 754.2656
66 | 16092 | 754.3125
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16090 | 754.2188
66 | 16092 | 754.3125
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16091 | 754.2656
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16091 | 754.2656
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16090 | 754.2188
66 | 16093 | 754.3594
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16091 | 754.2656
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16091 | 754.2656
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16090 | 754.2188
66 | 16091 | 754.2656
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16089 | 754.1719
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16089 | 754.1719
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16092 | 754.3125
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16091 | 754.2656
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16090 | 754.2188
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16090 | 754.2188
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16089 | 754.1719
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16091 | 754.2656
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16090 | 754.2188
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16091 | 754.2656
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16090 | 754.2188
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16080 | 753.75
66 | 16087 | 754.0781
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16089 | 754.1719
66 | 16079 | 753.7031
66 | 16086 | 754.0312
66 | 16088 | 754.125
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16088 | 754.125
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16088 | 754.125
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16090 | 754.2188
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16088 | 754.125
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16089 | 754.1719
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16089 | 754.1719
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16088 | 754.125
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16089 | 754.1719
66 | 16078 | 753.6562
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16088 | 754.125
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16087 | 754.0781
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16083 | 753.8906
66 | 16086 | 754.0312
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16088 | 754.125
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16081 | 753.7969
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16077 | 753.6094
66 | 16088 | 754.125
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16072 | 753.375
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16076 | 753.5625
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16075 | 753.5156
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16074 | 753.4688
66 | 16087 | 754.0781
66 | 16075 | 753.5156
66 | 16087 | 754.0781
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16075 | 753.5156
66 | 16084 | 753.9375
66 | 16075 | 753.5156
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16083 | 753.8906
66 | 16085 | 753.9844
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16087 | 754.0781
66 | 16083 | 753.8906
66 | 16087 | 754.0781
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16086 | 754.0312
66 | 16073 | 753.4219
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16086 | 754.0312
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16086 | 754.0312
66 | 16079 | 753.7031
66 | 16086 | 754.0312
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16085 | 753.9844
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16086 | 754.0312
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16071 | 753.3281
66 | 16083 | 753.8906
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16078 | 753.6562
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16075 | 753.5156
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16075 | 753.5156
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16085 | 753.9844
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16071 | 753.3281
66 | 16082 | 753.8438
66 | 16071 | 753.3281
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16078 | 753.6562
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16077 | 753.6094
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16078 | 753.6562
66 | 16080 | 753.75
66 | 16073 | 753.4219
66 | 16081 | 753.7969
66 | 16071 | 753.3281
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16076 | 753.5625
66 | 16078 | 753.6562
66 | 16080 | 753.75
66 | 16077 | 753.6094
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16078 | 753.6562
66 | 16078 | 753.6562
66 | 16079 | 753.7031
66 | 16077 | 753.6094
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16074 | 753.4688
66 | 16080 | 753.75
66 | 16070 | 753.2812
66 | 16082 | 753.8438
66 | 16069 | 753.2344
66 | 16081 | 753.7969
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16077 | 753.6094
66 | 16080 | 753.75
66 | 16074 | 753.4688
66 | 16080 | 753.75
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16071 | 753.3281
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16080 | 753.75
66 | 16073 | 753.4219
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16085 | 753.9844
66 | 16081 | 753.7969
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16081 | 753.7969
66 | 16077 | 753.6094
66 | 16080 | 753.75
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16071 | 753.3281
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16083 | 753.8906
66 | 16075 | 753.5156
66 | 16084 | 753.9375
66 | 16071 | 753.3281
66 | 16081 | 753.7969
66 | 16074 | 753.4688
66 | 16082 | 753.8438
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16074 | 753.4688
66 | 16083 | 753.8906
66 | 16072 | 753.375
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16074 | 753.4688
66 | 16085 | 753.9844
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16084 | 753.9375
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16079 | 753.7031
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16077 | 753.6094
66 | 16082 | 753.8438
66 | 16082 | 753.8438
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16072 | 753.375
66 | 16081 | 753.7969
66 | 16071 | 753.3281
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16072 | 753.375
66 | 16082 | 753.8438
66 | 16071 | 753.3281
66 | 16083 | 753.8906
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16076 | 753.5625
66 | 16082 | 753.8438
66 | 16071 | 753.3281
66 | 16082 | 753.8438
66 | 16069 | 753.2344
66 | 16081 | 753.7969
66 | 16078 | 753.6562
66 | 16078 | 753.6562
66 | 16082 | 753.8438
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16083 | 753.8906
66 | 16081 | 753.7969
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16072 | 753.375
66 | 16081 | 753.7969
66 | 16071 | 753.3281
66 | 16081 | 753.7969
66 | 16076 | 753.5625
66 | 16077 | 753.6094
66 | 16080 | 753.75
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16079 | 753.7031
66 | 16084 | 753.9375
66 | 16082 | 753.8438
66 | 16080 | 753.75
66 | 16079 | 753.7031
66 | 16078 | 753.6562
66 | 16080 | 753.75
66 | 16073 | 753.4219
66 | 16082 | 753.8438
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16074 | 753.4688
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16082 | 753.8438
66 | 16083 | 753.8906
66 | 16078 | 753.6562
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16077 | 753.6094
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16074 | 753.4688
66 | 16083 | 753.8906
66 | 16070 | 753.2812
66 | 16084 | 753.9375
66 | 16077 | 753.6094
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16080 | 753.75
66 | 16078 | 753.6562
66 | 16081 | 753.7969
66 | 16071 | 753.3281
66 | 16083 | 753.8906
66 | 16071 | 753.3281
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16074 | 753.4688
66 | 16084 | 753.9375
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16081 | 753.7969
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16071 | 753.3281
66 | 16083 | 753.8906
66 | 16076 | 753.5625
66 | 16073 | 753.4219
66 | 16081 | 753.7969
66 | 16070 | 753.2812
66 | 16079 | 753.7031
66 | 16075 | 753.5156
66 | 16082 | 753.8438
66 | 16081 | 753.7969
66 | 16077 | 753.6094
66 | 16083 | 753.8906
66 | 16073 | 753.4219
66 | 16085 | 753.9844
66 | 16073 | 753.4219
66 | 16084 | 753.9375
66 | 16080 | 753.75
66 | 16081 | 753.7969
66 | 16079 | 753.7031
66 | 16072 | 753.375
66 | 16083 | 753.8906
66 | 16071 | 753.3281
66 | 16080 | 753.75
66 | 16078 | 753.6562
66 | 16079 | 753.7031
66 | 16081 | 753.7969
66 | 16073 | 753.4219
66 | 16084 | 753.9375

## ADC autorange
Contents!A1
| mV range | scales | low range switch | high range switch
0 | 6144 | 0.1875 | 2250
1 | 4096 | 0.125 | 1500 | 3000
2 | 2048 | 0.0625 | 750 | 1500
3 | 1024 | 0.03215 | 385.8 | 771.6
4 | 512 | 0.015625 | 187.5 | 375
5 | 256 | 0.007813 |  | 187.512
|  |  | switch to lower range | switch to the highest range

## Sensor configuration set
Contents!A1
This table defines what configuration should be used on the start of the measurement with given parameter.
[14 empty] | Byte | Value | Message type
[14 empty] | 4 | 3 | Start measurement trigger
Sensor configuration set command | [13 empty] | 5 | uint8_t | Sample period in 0.1s units (1 is 100ms)
0b0000 0000 | No configuration, all analog inputs are generic, always powered, reading mV values | [12 empty] | 6 | uint8_t | Sensor configuration set
0bXXXX XXX1 | A3.2 used for TGS24444 (Amonia), selecting this will set 250ms period and ignore period setting | [12 empty] | EOP | EOP | EOP
0bXXX0 001X | Electrochemical sensors used, LMP91000 will be set up on the beginning of measurement with parameters: | [10 empty] | hex
|  | for | RLOAD | TIA_GAIN | BIAS | INT_Z | TIACN | REFCN | MODECN |  |  | TIACN | REFCN | MODECN
| A1.0 |  | 10 | 120k | 0V | 20%VDD | 0b11000 | 0b00000000 | 0b011 |  |  | 18 | 00 | 03
| A1.1 |  | 10 | 120k | 0V | 20%VDD | 0b11000 | 0b00000000 | 0b011 |  |  | 18 | 00 | 03
| A1.2 |  | 10 | 120k | 0V | 20%VDD | 0b11000 | 0b00000000 | 0b011 |  |  | 18 | 00 | 03
| A1.3 |  | 33 | 120k | 0V | 20%VDD | 0b11001 | 0b00000000 | 0b011 |  |  | 19 | 00 | 03
| A2.0 |  | 33 | 35k | 0V | 67%VDD | 0b10101 | 0b01000000 | 0b011 |  |  | 15 | 40 | 03
| A2.1 |  | 10 | 35k | 0V | 20%VDD | 0b10100 | 0b00000000 | 0b011 |  |  | 14 | 00 | 03
| A2.2 | O2-A1 | 50 | 14k | 0V | 20%VDD | 0b10010 | 0b00000000 | 0b011 |  |  | 12 | 00 | 03
| A2.3 |  | 10 | 7k | 0V | 20%VDD | 0b01100 | 0b00000000 | 0b011 |  |  | 0C | 00 | 03
| A3.1 |  | 33 | 3.5k | 300mV | 50%VDD | 0b01001 | 0b00100111 | 0b011 |  |  | 09 | 27 | 03
| A3.2 |  | 33 | 3.5k | 0V | 20%VDD | 0b01001 | 0b00000000 | 0b011 |  |  | 09 | 00 | 03
If LMP91000 is not detected on port, error will be reported, port will be still used and measured.
0bXXX0 011X | Electrochemical sensors used, LMP91000 is set up on all ports according to patemeters sent in datagram index 7:TIA,8:REF,9:MODE
[10 empty] | ref divider | ref | swing
|  | max uA | Load R | Leads | Bias mV |  | RLOAD | TIA_GAIN | BIAS | INT_Z |  |  | note
| O2-A1 | 240 | 47..100 | 2 | 0 |  | 50 | 14k | 0 | 20% |  |  | test polarity if less O2 is lower voltage
| TGS5042 | 24 | 0 | 2 | 0 |  | 10 | 120k | 0 | 20% | Vref 1.25V or VDD
| NO-AE | 400..800 | 10..47 | 3 | 300 |  | 10 | 3.5k | 12% | 50% | VDD |  | better from 1.25v ref
| NO2-AE | -34 | 33 | 3 | 0 |  | 33 | 35k | 0 | 67% | VDD | 3.35..1.19
| HCN-A1 | 8.5 | 10..33 | 3 | 0 |  | 10 | 120k | 0 | low |  |  | lower gain due to cross-sens?
| HCL-A1 | 13 | 10..33 | 3 | 0 |  | 10 | 120k |  | low |  |  | lower gain due to cross-sens?
|  | RLOAD | TIA_GAIN | BIAS | INT_Z | TIACN | REFCN | MODECN |  |  | TIACN | REFCN | MODECN |  | O2-A1 DC [mV] | actuated
| A2.2 | 50 | 14k | 0V | 20%VDD | 0b10010 | 0b00000000 | 0b001 |  |  | 12 | 00 | 01 |  | 3.25 | 5
[5 empty] | 67%VDD | [6 empty] | 40 |  |  | 3.25 | 16
|  |  | 120k | [7 empty] | 18 | 40 |  |  | 3.25 | 4.5
|  |  | 14k | [7 empty] | 18 | 0 |  |  | 3.25 | 3.4
| RLOAD | TIA_GAIN | BIAS | INT_Z | DC output and breathed out excitement
O2-A1 | 100 | 350k | +10% | 67% | 3mV not excited
| 100 | 350k | -10% | 67% | 3mV not excited
| 100 | 350k | +0% | 50% | 3mV not excited
| 100 | 35k | +0% | 50% | 9.3mV not excited

## El-chem fake sensor
Contents!A1
| TIACN | 0b00010111 | 35k |  |  |  |  | 100ohm |  |  |  | https://www.zimmerpeacocktech.com/knowledge-base/faq/lmp91000-trouble-shooting/
| REFCN | changing
| MODECN | 0b00000011 | 3-lead amperometric
| RefSource | Int
|  |  | mV | mV
|  |  | A2.0 | A2.2
REFCN | REFCN | Bias% | board1 35k | board2 35k
0x11 |  | 1 | 998.1 | 997.4
0x14 |  | 6 | 1006.7 | 1006.15
0x17 | 0b0001 0111 | 12 | 1017 | 1016.55
0x18 | 0b0001 1000 | 14 | 1020.45 | 1020
0x19 | 0b0001 1001 | 16 | 1023.55 | 1023.9
0x1A | 0b0001 1010 | 18 | 1027.3 | 1027
0x1B | 0b0001 1011 | 20 | 1030.7 | 1030.4
0x1D | 0b0001 1101 | 24 | 1030.7 | 1030.37
| diff 20-12 | 8 | 13.7 | 1.7125 | mV/%
|  |  |  | 171.25 | 100%
|  |  | board1 350k | board2 350k
0x14 |  | 6 | 1099.4 | 1099.9
0x19 |  | 16 | 1271.25 | 1273.7
0x1B |  | 20 | 1339 | 1342.3
INT_Z | 67%
TIA | 35k
[5 empty] | VDD | 5 | V
[5 empty] | Vref divided 67% | 3.35 | V
[5 empty] | 20% bias | 0.67 | V |  | VDD | 5 | V
[5 empty] | R RE,CE to WE | 1000000 | ohm |  | Vref divided 20% | 1 | V
[5 empty] | I_WE | 2.68e-06 | A |  | 20% bias | 0.2 | V
[6 empty] | 2.68 | uA |  | R RE,CE to WE | 1000000 | ohm
[5 empty] | R TIA gain | 35000 |  |  | I_WE | 8e-07 | A
[5 empty] | Out volt | 3.4438 | V |  |  | 0.8 | uA
0x1B | 0b0001 1011 | 20 | 3379 | 3380 |  | 3443.8 | mV |  | R TIA gain | 350000
[9 empty] | Out volt | 1.28 | V
[10 empty] | 1280 | mV

## LabVIEW data log
Contents!A1
Loggging data definition
Data are found in VI path under folder DATA
Column decriptions
Time | Unit the data was received from | Sample number | Sensor type1 | Data1 | Sensor type2 | Data2 | Sensor type3 | Data3 | Sensor type4 | Data4 | Sensor type5 | Data5 | Sensor type6 | Data6 | Sensor type7 | Data7 | Sensor type8 | Data8 | .. | …

## Error Codes
Contents!A1
|  | Byte | Value
ERROR CODES |  | 4 | 3 | ERROR
0 | No error | 5 | uint_8 | ERROR CODE
1 | failed to initalize Ethernet controller
2 | failed to obtaiin IP from DHCP server
3 | SD card communication fail. continuing with no logging to SD card | DO NOT ADD ERROR CODES TO THIS LIST, EDIT SOURCE FILE HERE:
4 | failed to initalize I2C bus | D:\PROJECTS\01_GIT_REPOS\FRA_Sensor_platform\3-FW_SW\2-SW_log_PC\1-LabVIEW_app
5 | failed to initalize IO Expander
6 | failed to initalize ADC
7 | error in setting time. time is from the past. time not set. continuing
8 | IO expander output does not match the setting - power setting not valid
9 | failed to read IO expander status - power settings not valid
10 | test error reporting - NO ACTUAL ERROR
11 | ADC1 readout failed - timeout
12 | ADC2 readout failed - timeout
13 | ADC3 readout failed - timeout

## LW Error Codes
Contents!A1
Labview error codes
500 | Power setting not accepted. Invalid input

## Time budget
Contents!A1
| bytes | time [us]
Sending UDP | 1 | 284 |  | SD logging
Sending UDP | 8 | 300
Writing to SD | 68 | 9000 | 20MHz SPI
I2C ADC
I2C speed | 400 | kHz
time to get one bit | 0.0025 | ms
2 bytes / one reading | 0.04 | ms
9 sensors | 0.36 | ms

## TO DO
Contents!A1
Issue | Status
Inputs for setting up IP address - digital - JTAG pins | cancelled - set in DHCP server
Analog input for sensor detection | cancelled - would have to be switched to each sensor port
Baseline test for connected sensors | cancelled
change SHCP timeout from 1s back to 60s | done
SD status flag to UDP | done
SD full card resolution
BOARD V1 FEEDBACK>
IC9 heat pad unmask | done
10k is 1206 - fix in BOM to 0602 | done, was never there
R46 replace with ferite bead | done
cp2101 edit footprint, corner out | done
remove SD connector | not done, more work
rotate microSD connector 180 deg. | done
blocking cap to IO expander? | done
Error message not received in some cases | done - due to interrupt, packetloop must run also on sending
R1. R2 10k? Test on 2nd board | not changed, to be safe
C10 10u/25V place | done
C17 place 10u 10v normal | done
pca9555 check packageing | done
change pca9555 so also the other one can be used | done
C16 change to 100n, avoid loading fw problems | done | 10u tested, works but 100n better for CP2102
SMPS inductors to SRP5030CA-1R5M | done
Connector power | done | https://www.mouser.com/ProductDetail/CUI/PJ-059A?qs=sGAEpiMZZMtnOp%252BbbqA009lE0K0K%252BPZGgL2j%2Fo1XCBQtBpzQrCB9dQ%3D%3D
5V jumper to ethernet port | done
LED changed to header | done
cross check BOMsafter export | done
add resettable fuse on the input
IO expander, add 0R to A2 input to VCC and GND. This will enable using PCA9539 too | done
add I2C address selecting resistors in case some sensor has conflicting address | I2C address of first ADC adjusted
FW>
add error code array instead of one value | done
debug error code sending from multiple units | done | reset was not performed properly, software reset was not wquivalent to HW reset
power report not received from 24 | done | timeout needed to be longer
Alive units display in Labview | done
init all units with 5V ON | done
Button hold trigger | done
Button long hold reset | done
Light sensor integration | done
Humidity and temperature sensor integration | done
set time on each measurement start, or on each received dtg | done
Name array for Labview charts legend | done
Autorange Analog inputs | done
Labview, second trigger opens another file, does not close old one. - resolve | done
Light temperature and humidity, change to unsigned saving | done
Add sampling period to SD file or time | done
adjust range on ammonia sensor - is full range
System>
Power critical Units from separate power supplies, in case they short they will not take more units with them - inform
Research FW loading tool
LMP91000 not found, create error message
Sensor board>
Change comparator power to 5V not calm | cancelled
heat relief on LMP91000
remove C 100n on A3.2 on all boards - for ammonia | done in BOM
add RC to FET gate - avoid switching on 1ms pulse | done in new design
LMV611 fix footprint to SOT-32 or BOM to SC70 | deleted from design
re-export data for sensor board TGS2xxx, 1nF on low pass to FET | done

## System setup
Contents!A1
1 | set up DHCP on router
2 | connect Unit with ethernet cable
3 | turn on power to unit
4 | watch DHCP client list in router configuration
5 | Set up IP address reservation for the sensor unit
6 | power cycle or reset unit
7 | open labview app
8 | Unit should appear in acknowledged clients
9
10
11
12
13
14
to be derived from code block diagram

## Sensors Example datasheet val.
Contents!A1
|  | PN | Property measured | Precision | Voltage | Power | Ouptut | Size | Connection method | Interface | AA | Note
| 1 | iSweek HVAC - ZG09 | CO2 0-10k ppm | .+-50ppm, 3% reading | 5 | 190mA, avg 34mA | UART, I2C, PWM, DAC | 32x20x14 | 2x Pin header 4 + 5 contacts, 2.54 pitch | I2C (5v or 3v3?), cables or custom board | get exact part numebr
| 2 | iSweek NO-AE | Nitride Oxide 0+5k ppm |  | 0.3 V bias |  | x uA Reference, Worker, counter | 20x20x16 | GasPCB1, Pins 1.5mm dia | RWC board | get application note
| 3 | Spec Sensors 110-501 | NO2 0-20ppm | 20ppb, 3% reading |  | 50uW | Reference, Worker, counter | 20x20x3 | 3 pin header 2+2+2 contacts 2mm pitch | RWC board |  | https://eu.mouser.com/ProductDetail/3M-Electronic-Solutions-Division/950502-6102-AR?qs=mk7%252BW39U1H95%252BMM5da7bdg%3D%3D
| 4 | iSweek HCL-A1 | HCl 0-100ppm |  | not required bias |  | 0-12 uA Reference, Worker, counter | 20x20x16 | GasPCB1, Pins 1.5mm dia | RWC board
| 5 | iSweek HCN-A1 | HCN 0-100ppm |  | not required bias |  | 0-10 uA Reference, Worker, counter | 20x20x16 | GasPCB1, Pins 1.5mm dia | RWC board
| 6 | iSweek O2-M2 | O2 |  |  |  | 0-120 uA +Ve,-Ve | 20x20x16 | GasPCB1, Pins 1.5mm dia
| 7 | Figaro TGS 2444 | Ammonia |  | 5V switched, timed | mean 56mW, 300mA instant | resistor div, + transistor | 9x9x7.8 | FigaroPCB1, Pins to solder 0.55 dia |  | find alternative
| 8 | Figaro TGS 2602 | VOCs |  | 5V | 280 mW | resistor div | 9x9x7.8 | FigaroPCB1, Pins to solder 0.55 dia
| 9 | Figaro TGS 5042 | CO 0-10k ppm |  |  |  | 0-24uA, ext opamp needed | 49x10x10 | To solder only
| 10 | Figaro TGS 6812 | H, Methane 0-25k ppm |  | 3V +-0.1V |  | resistor whatstone bridge compensator detector | 12x12x9 | Pins to solder 0.8 dia
| 11 | Figaro TGS 2611 | [6 empty] | FigaroPCB, Pins to solder 0.55 dia
| 12 | 4P50 CiTipeL |  |  |  |  | Reference, Worker, Counter? |  | 4P50PCB,Pins 1.5mm
13
14
15

## IssueTracking
Contents!A1
add I2C address selecting resistors in case some sensor has conflicting address
IO expander, add 0R to A2 input to VCC and GND. This will enable using PCA9539 too | NXP datasheet: The PCA9539; PCA9539R is identical to the PCA9555 except for the removal of the internal I/O pull-up resistor which greatly reduces power consumption when the I/Os are held LOW, replacement of A2 with RESET and a different address range
TO DO'!A1

## Button and LED UI table
Contents!A1
LED state | Period [s] | Duty cycle [%]
OFF |  |  | initalizing unit after power-up
OFF |  |  | main loop not responding
Blinking | 1 | 20 | Waiting for DHCP (max 60 seconds)
Blinking | 1.5 | 50 | Button press registered, pressed 1 to 4s, measurement will start on button release
Blinking | 0.2 | 50 | Button press registered, pressed 4 - 15s, Unit will restart on release or after 15s from press
Measuring | measurement period | see note | LED off for about 40ms of measurement period - wink every meas. period