## Hárok1
|  |  |  | 8000 ppm Methane | 20000 PPM Methane |  | 2000 ppm CO2 | 6000 ppm CO2 | 10000 ppm CO2 |  | 110 ppm amonia | 280 ppm amonia
AI | Sensor | Gas | Baseline mV air | mV |  | Air |  |  |  | air
A2.0 | TGS5042 | CO | 984 | 1000 | 1000 | 1000 | 1000
A2.1 | 4P50 | Methane | 2037 | 2000 | 1950 | 2038 | 2038
A2.2 | TGS2611 | Methane | 485 | 3798 | 4239 | 350 | 380
A2.3 | MG-812 | CO2 | 260 | 340 | 340 | 382 | 248.8 | 216 | 201.7
A3.1 | O2-AE | Oxygen | 1950 | 1940 | 1960 | 1925 | 1925
A3.2 | TGS2444 | Amonia | 510 | 250 | 225 | 200 | 160 |  |  | 700 | 2800 | 3835
RH |  |  | 43% | -2 | -2.5
C |  |  | 28d C | 29 | 29
TGS2611 Rs |  |  | 47477.3196 | 1614.06 | 915.5697
R in 5000ppm Methane guess | 1880 |  | 0.039598 | 1.1648 | 2.0534
|  |  | 1.2935 | 8036.3791 | 34759.0606
[6 empty] | CO2 mV diff | -133.2 | -166 | -180.3
[6 empty] | Co2 ppm | 2029.9027 | 6191.565 | 10068.2597
[9 empty] | Rs amonia | 31328.5714 | 4007.1429 | 1549.2829
| NO-AE | NO-AE | RE and CE switched. |  |  |  |  | sampling 1s
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 35k | 33R | VDD | 50% | POS | 0.06 | 3-lead amperometric
|  | NO-AE | mV | at sample
|  | air | 2573 | to 260
|  | 1000ppm NO | 4155 | 270-370
|  | gas off, hose kept on sensor |  | from 370
|  | expected at 1000ppm - datasheet | 6e-05 | A
|  | expected rise with gain 35k | 2.1 | V
|  | measured rise | 1.582 | V
|  | measured sensor current | 4.52e-05 | A
|  | measured sensoitivity | 4.52e-08 | A/ppmNO
|  | NO-AE | RE and CE switched. |  |  |  |  | sampling 1s
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 35k | 33R | VDD | 20% | POS | 6% | 3-lead amperometric
|  | NO-AE | mV
|  | air | 1062
|  | 1000ppm NO | 2699
gas off, hose kept on sensor
|  | expected at 1000ppm - datasheet | 6e-05 | A
|  | expected rise with gain 35k | 2.1 | V
|  | measured rise | 1.637 | V
|  | measured sensor current | 4.6771e-05 | A
|  | measured sensoitivity | 4.6771e-08 | A/ppmNO
|  | NO-AE | RE and CE switched. |  |  |  |  | sampling 1s
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 14k | 33R | VDD | 20% | POS | 0.06 | 3-lead amperometric
|  | NO-AE | mV
|  | air | 1024 | to 37
|  | 1000ppm NO | 1674 | 37-120
|  | gas off, hose off the sensor |  | from120
|  | expected at 1000ppm - datasheet | 6e-05 | A
|  | expected rise with gain 14k | 0.84 | V
|  | measured rise | 0.65 | V
|  | measured sensor current | 4.6429e-05 | A
|  | measured sensitivity | 4.6429e-08 | A/ppmNO
THIS IS FINAL SETTING! Verified on 2 boards
|  | max measureable current till saturation | 0.00028571 | A
|  | corresponds to ppm NO - max measurable | 6153.8462 | ppm CO
|  | re-measured with final settings | [5 empty] | 1s sampling
Sensor inserted correctly
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 14k | 33R | VDD | 20% | POS | 0.06 | 3-lead amperometric
|  | air | 1016
|  | 1000ppm NO | 1659
gas off, hose off the sensor
|  | expected at 1000ppm - datasheet | 6e-05 | A
|  | expected rise with gain 14k | 0.84 | V
|  | measured rise | 0.643 | V
|  | measured sensor current | 4.5929e-05 | A
|  | measured sensitivity | 4.5929e-08 | A/ppmNO
| NO2-AE | NO2-AE | RE and CE switched.
|  | middle sensitivity - datasheet | -120 | nA/ppm NO2 | -70 to -170 nA/ppm
|  |  | -1.2e-07 | A/ppm NO2
|  | exposed by 200ppm, expected current | -2.4e-05 | A
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 35000 | 33R | VDD | 20% | NEG | 0 | 3-lead amperometric
|  | expected output at 200ppm NO2 | -0.84 | V |  |  | 1s sampling
sample
|  | Output on air in mV | 1039 | to 40
|  | 200ppm NO2 outuput in mV | 710 | 40-560
|  | gas off |  | from 560
|  | measured rise (V) | -0.329 | V
|  | measured sensor current | -9.4e-06 | A
|  | measured sensoitivity on 200ppm | -4.7e-08 | A/ppmNO2
|  | measured sensoitivity on 200ppm | -47 | nA/ppmNO2
|  | NO2-AE | RE and CE switched.
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 35000 | 33R | VDD | 67% | NEG | 0 | 3-lead amperometric
1s sampling
sample
|  | Output on air in mV | 3385 | to 55
|  | 200ppm NO2 outuput in mV | 3065 | 55-155
|  | gas off |  | from 155
|  | measured rise (V) | -0.32 | V
|  | measured sensor current | -9.1429e-06 | A
|  | measured sensitivity on 200ppm | -4.5714e-08 | A/ppmNO2
|  | measured sensitivity on 200ppm | -45.7143 | nA/ppmNO2
THIS IS FINAL SETTING! Verified on 2 boards
|  | max measureable current till saturation | -9.5714e-05 | A
|  | corresponds to ppm NO - max measurable | 2093.75 | ppm CO
| HCL-A1 | HCL-A1 | RE and CE insterted correctly
|  | middle sensitivity - datasheet | 105 | nA/ppm HCL | 80 to 130 nA/ppm
|  |  | 1.05e-07 | A/ppm HCL
|  | exposed by 25ppm, expected current | 2.625e-06 | A
settings
|  | Gain | Load R | Reference | Internal zero from 2.5V ref | Polarity bias | bias | mode
|  | 35000 | 33R | VDD | 20% | NEG | 0 | 3-lead amperometric
294.4
|  | expected output at 25ppm HCL | 0.091875 | V | 391.78
|  | air | 294.4 | mV
|  | exposed by 25ppm | 391.78 | mV
|  | measured rise (V) | 0.09738 | V
|  | measured sensor current | 2.7823e-06 | A
|  | measured sensitivity on 25ppm | 1.1129e-07 | A/ppmHCL
|  | measured sensitivity on 25ppm | 111.2914 | nA/ppmHCL
|  | settings | RE and CE insterted correctly
|  | Gain | Load R | Reference | Internal zero from 2.5V ref | Polarity bias | bias | mode
|  | 14000 | 10R | VDD | 50% | NEG | 0 | 3-lead amperometric
|  | expected output | 0.03675 | V
|  | air | 1590 | mV
|  | exposed by 25ppm | 1629 | mV
|  | measured rise (V) | 0.039 | V
|  | measured sensor current | 2.7857e-06 | A
|  | measured sensitivity on 25ppm | 1.1143e-07 | A/ppmHCL
|  | measured sensitivity on 25ppm | 111.4286 | nA/ppmHCL
|  | settings | RE and CE insterted correctly
|  | Gain | Load R | Reference | Internal zero from 3.3V ref | Polarity bias | bias | mode
|  | 35000 | 10R | VDD | 67% | NEG | 0 | 3-lead amperometric
|  | expected output | 0.091875 | V
|  | air | 2220 | mV
|  | exposed by 25ppm | 2292 | mV
|  | measured rise (V) | 0.072 | V
|  | measured sensor current | 2.0571e-06 | A
|  | measured sensitivity on 25ppm | 8.2286e-08 | A/ppmHCL
|  | measured sensitivity on 25ppm | 82.2857 | nA/ppmHCL
|  | settings | RE and CE insterted correctly
|  | Gain | Load R | Reference | Internal zero from 5V VDD | Polarity bias | bias | mode
|  | 120000 | 33R | VDD | 20% | NEG | 0 | 3-lead amperometric
|  | expected output | 0.315 | V
|  | air | 1070 | mV
|  | exposed by 25ppm | 1416 | mV
|  | measured rise (V) | 0.346 | V
|  | measured sensor current | 2.8833e-06 | A
|  | measured sensitivity on 25ppm | 1.1533e-07 | A/ppmHCL
|  | measured sensitivity on 25ppm | 115.3333 | nA/ppmHCL
THIS IS FINAL SETTING! Verified on 1 board
|  | max measureable current till saturation | 3.3333e-05 | A
|  | corresponds to ppm HCL - max measurable | 289.0173 | ppm HCL
| HCN-A1 | HCN-A1 | RE and CE switched. |  |  |  | 1s sampling
|  | middle sensitivity - datasheet | 70 | nA/ppm HCN | 55 to 85 nA/ppm
|  |  | 7e-08 | A/ppm HCN
|  | exposed by 50ppm, expected current | 3.5e-06 | A
settings
|  | Gain | Load R | Reference | Internal zero | Polarity bias | bias | mode
|  | 120000 | 33R | VDD | 20% | NEG | 0 | 3-lead amperometric
|  | expected output at 50ppm HCN | 0.42 | V
sample
|  | air mV, slow stabilization, would be lower | 1309 | to 50-340
|  | 50ppm HCN outuput in mV | 1861 | from 340
|  | measured rise (V) | 0.552 | V
|  | measured sensor current | 4.6e-06 | A
|  | measured sensitivity on 50ppm | 9.2e-08 | A/ppmHCN
|  | measured sensitivity on 50ppm | 92 | nA/ppmHCN
THIS IS FINAL SETTING! Verified on 1 board
|  | max measureable current till saturation | 3.3333e-05 | A
|  | corresponds to ppm NO - max measurable | 362.3188 | ppm HCN
repeated measurement with long stabilization on air (overnight)
|  | air | 1116 | mV
|  | 50ppm HCN outuput in mV | 1681 | mV | sample 300-800
|  | measured rise (V) | 0.565 | V
|  | measured sensor current | 4.7083e-06 | A
|  | measured sensitivity on 50ppm | 9.4167e-08 | A/ppmHCN
|  | measured sensitivity on 50ppm | 94.1667 | nA/ppmHCN
ca. 30 min
exposed to 25ppm / 10.5min